## Slide 1: Treatment facility optimization
Andrea De Franco
MedAustron
International Conference on Medical Accelerator and Particle Therapy
CNA Seville, 4th - 6th Sep 2019

## Slide 2: Treatment Time Reduction
Intra-fractional movement
Patient comfort
Patient throughput
M. Stock, OMA school Vienna, 2019

## Slide 3: MedAustron
Research Room:
Treatment/Clinical Research:
Irradiation Rooms
Ion Sources
LINAC
Synchrotron
Proton treatment from Dec 2016
Carbon treatment from July 2019

## Slide 4
Injector Hall - LEBT
3 Ion Sources and particle type switching area

## Slide 5
LINEAR ACCELERATOR - LI
Solenoid, RFQ, IH tank, acceleration to 7 MeV/n

## Slide 6
SYNCHROTRON - MR
Proton 60-250 (800) MeV, carbon 120-400 MeV/n

## Slide 7: HEBT
HIGH ENERGY TRANSFER LINE - HEBT
Beam to 3 treatment & 1 research rooms

## Slide 8: Workflow & Statistics
(H. Hentschel)

## Slide 9: Irradiation
New spill requested when:
No more particles stored in synchrotron
New energy required
Sync. B-Field
Statistics* ( Dec 16 – May 19)
Avg. Slices per portal: 41.43
Avg. Spills per portal: 44.65
Avg. Spills per Slice: 1.072
*Patient selection biased by machine performance!

## Slide 10: Irradiation time red. Options
Flux
Flux [p/s]
Time [s]
Flux [p/s]
Time [s]
T between
spills
Flux [p/s]
Time [s]
nr. of spills

## Slide 11: Irradiation time red. Options
Flux
Flux [p/s]
Time [s]
Flux [p/s]
Time [s]
T between
spills
Flux [p/s]
Time [s]
nr. of spills

## Slide 12: Time between spills - 1
Flux [p/s]
Time [s]
T between
spills
Unused beam
Sync. B-Field
Dynamic timings for accelerator phases control
Implemented at MedAustron in 2018
Project Leader: C. Schmitzer
Deputy: A. De Franco
Up to -40% treatment time

## Slide 13: Resonant Slow Extraction
Wrong way to use resonances:
Tacoma Narrows Bridge example
Amplitude
Momentum or Tune
Condition of Resonance
Unstable beam
Stable beam
Condition for resonance
Notes: Prior to extraction: Acceleration

## Slide 14: Resonant Slow Extraction
    Slowly increase momentum spread to move beam into resonance Qx=1.6666
histogram dp/p
Protons measured Schottky monitor
(A. Goetz bachelor thesis)
Hor. phase space of one particle during the extraction
Electrostatic septum wire
(F. Kühteubl Master thesis)
Notes: Prior to extraction: Acceleration

## Slide 15: Time between spills - 2
Combined
~2 sec per spill
Up to -30% treatment time
Non-adiabatic capture
Dual phase Betatron Core settings
Parallel beam preparation (Longitudinal and Transverse)
Avoid dipole ‘washing’
(B-Field active reg.)
Sync. B-Field

## Slide 16: Irradiation time red. Options
Flux
Flux [p/s]
Time [s]
Flux [p/s]
Time [s]
T between
spills
Flux [p/s]
Time [s]
nr. of spills

## Slide 17: Higher Intensities
Flux [p/s]
Time [s]
Safety (Standards & Risk Management)
Functional
MinP
Machine failure
Intensity ripples (<ms)
MinP:= ~lowest spot weight in the slice(~106p) < -- > minimum beam dwell time (~1ms)
We use slow extraction to stretch the pulse. Doing the opposite is ‘easy’, but…
Need to:
Model beam dynamics during failures
Reduce termination times
Reduce intensity ripples
Speed up scanning system
…Voxel based flux
Expected implementation Q1 2020
Project Leader: A. De Franco
> x2 Flux
(x4 stored intensity in synchrotron)
Up to -35% treatment time

## Slide 18: Scanning Speed

## Slide 19: Ripples suppression
Spill time structure (50kHz)
En:62.4 MeV - Frev: 1.33 MHz.
∆Fch:15kHz – 1500V
With RF channeling
Change Spot Sizes, Range !!!
 Re-commissioning

## Slide 20: Irradiation time red. Options
Flux
Flux [p/s]
Time [s]
Flux [p/s]
Time [s]
T between
spills
Flux [p/s]
Time [s]
nr. of spills

## Slide 21: Nr. Spills Reduction
New spill requested when:
No more particles stored in synchrotron
New energy required
Sync. B-Field
Statistics* ( Dec 16 – May 19)
Avg. Slices per portal: 41.43
Avg. Spills per portal: 44.65
Avg. Spills per Slice: 1.072
*Patient selection biased by machine performance!

## Slide 22: Nr. Spills Reduction
Need to:
Model beam dynamics during failures
Reduce termination times
Might have to:
Upgrade Ion source
Use higher harmonics capture
New spill requested when:
No more particles stored in synchrotron
New energy required
Ripple filters
Higher stored intensities in the synchrotron
Spread Bragg peak to reduce the number of energy layers necessary
Change the energy of the beam stored in the synchrotron
Multi energy extraction

## Slide 23: Multi Energy Extraction
Sync. B-Field
Flux [p/s]
Time [s]
En1
En2
En4
En5

## Slide 24: Resonant Slow Extraction
    Slowly increase momentum spread to move beam into resonance Qx=1.6666
histogram dp/p
Protons measured Schottky monitor
(A. Goetz bachelor thesis)
Hor. phase space of one particle during the extraction
Electrostatic septum wire
(F. Kühteubl Master thesis)
Notes: Prior to extraction: Acceleration

## Slide 25: Resonant Slow Extraction
    Slowly increase momentum spread to move beam into resonance Qx=1.6666
histogram dp/p
Protons measured Schottky monitor
(A. Goetz bachelor thesis)
Hor. phase space of one particle during the extraction
Electrostatic septum wire
(F. Kühteubl Master thesis)
De-bunched! Not easy to re-accelerate
Need bunched extraction!!!
Notes: Prior to extraction: Acceleration

## Slide 26: RF-KO extraction
Amplitude
P, Q
Amplitude
P, Q
(A. Goetz bachelor thesis)
(A. Goetz bachelor thesis)

## Slide 27: RF-KO extraction
Measured intensity time profile of extracted beam with RF-KO
Preliminary optics
re-design fro RF-KO
Betatron Core
RF-KO
(F. Kühteubl Master thesis)

## Slide 28
THANK YOU
Andrea De Franco
andrea.de.franco@medaustron.at
This project has received funding from the European Union’s Horizon 2020 research and innovation programme under the Marie Skłodowska-Curie grant agreement No 675265.

## Slide 29
Proton treatment from Dec 2016
Carbon treatment from July 2019

## Slide 30
ION SOURCES S1 – S3
Generation of or beams

## Slide 31: Phase Jump – Enlarge dp/p
Momentum Distribution (Schottky mon.)
Blue: Injection (± 0.3 ‰)
Red: Flat Top (± 0.3 ‰)
Green: After Phase Jump (± 3 ‰)
Intensity (a.u.)
Momentum (‰)
± 3 ‰
± 0.3 ‰
Notes: Prior to extraction: Acceleration

## Slide 32: Ripples suppression