## Slide 1: Beamline measurement of a non-invasive BCM: its advantages and limits
Sudharsan Srinivasan

## Slide 2: COMET
Superconducting cyclotron
250MeV proton beam
Compact cyclotron
Relevant Beam Parameters

## Slide 3: Specific Diagnostic needs
Beam parameter of interest: Beam intensity (Most Important)
Two different modes of operation: Diagnostics for commissioning/Diagnostics for Standard operation
Diagnostics for Standard Operation [1]:
  To control and improve accelerator operation (precise)
  Diagnosis of unwanted errors and to trigger interlocks
  High demands for accuracy
  Minimum beam disturbing schemes
Repetitive and stable beam: CW system (Closed loop system)
Signals are treated typically in Frequency domain
Beam current is a non- or slowly varying parameter
High precision by averaging

## Slide 4: Review of the BCM
Standard Beam Current Measurements:
ionization chambers are the most commonly used detector type
caveat: intercept beam  decrease the beam’s quality by scattering
An alternative:
a sensitive RF-based current monitor
fully non-interceptive beam current measurement
simple and robust system (KISS), good for safety purpose/applications
Challenge:
For very low charge beams (bunch charge ~ 1.36e-17 C)
Trade-off between bandwidth and sensitivity
Potential benefit:
design a new beam position monitor system using the same technics
IC
NOTE: Slide provided by Pierre André Duperrex

## Slide 5
BPM types | Advantages | Bandwidth | PROSCAN
Linear-cut | highly linear and good position sensitivity | Broadband | No
Button | non-linear and good position sensitivity | Broadband | No
Stripline | directivity and better position sensitivity | Broadband | Yes (3)
Pill-box | extermely linear and highly position sensitive | Narrowband | Yes (2)
Reentrant | extermely linear and highly position sensitive | Narrowband | Yes (1)
Resistive WCM | linear and average position sensitivity | Broadband | No
Inductive WCM | linear and position sensitivity better than linear-cut | Broadband | No
BPM types | Bunch/Beam size | Signal strength & quality | Mechanical realization
Linear-cut | Longer | smaller & deformations | complex
Button | comparable | smaller & deformations | simple
Stripline | short | larger & minor deformations | complex
Pill-box | very short | larger ; minuscule deformations | simple
Reentrant | very short | larger; minuscule deformations | simple
Resistive WCM | short | large with thermal noise | complex
Inductive WCM | short | better than linear-cut | complex
[2,3,4]

## Slide 6: Review of the BCM
Dielectric-Filled Reentrant Cavity Resonator as a Low-Intensity Proton Beam Diagnostic, Instruments 2018, 2,24
TM010
Induced E field
Induced B field
For 1nA at 145.7MHz
Matched to 2nd harmonic of the pulse repetition rate=72.85MHz

## Slide 7: PROSCAN Beamline
BCM located at 22m from the cyclotron exit
Placed behind an Ioniozation Chamber Ionization Chamber used as reference for measurements

## Slide 8: Measurement Chain
2*50Ohm LNA with a BP filter inbetween (center freq: 145.7MHz ; span=10MHz)
Subsystems of VME MESTRA (PSI developed)
  Preamplifier
  Digitizer (ADC); ADC3110 : max. input power of 10dBm (2V peak-peak) 16bit resolution (0.0001V)
  Digital Down Converter (DDC)
  Current Calculations
VME MESTRA converts RF signals to digital signals proportional to absolute beam current
Friss Formula

## Slide 9: Measurement Chain: Power Budget
The typical power budget for the measurement system is:
Resonator at 1nA: -134dBm
Local amplifiers: +70dB -64dBm
Cellflex cable 150m: -3dB -67dBm
1st Amplifier: +25dB -42dBm
2nd Amplifier: +25dB -17dBm (0.089V peak to peak)
Noise Level

## Slide 10: Resonator Vs MMAC5
Uncertainty in noise arises from uncertainty of MMAC5 (15-20% measurement uncertainty )

## Slide 11
For lower beam energies, energy spread is higher beyond degrader. Hence, pulse length increases thus decreasing the 2nd harmonic component of the beam current.

## Slide 13: Influence of MMAC5

## Slide 14: Influence of MMAC5
For higher beam energy (200-230MeV), the influence of IC is minor
As for lower beam energies, IC influence is detrimental affecting signal sensitivity
IC has a huge measurement uncertainty (15-20%)
Beam measurements without IC in beamline recommended

## Slide 15
Summary
TM010 resonance
60pA at 230MeV; 500pA at 140MeV
Few 100ms integration time
No activation unlike in IC
Energy spread affects signal level
NO IC in Beamline

## Slide 16
G.Kube, CAS on Beam Diagnostics, ``Specific diagnostic needs for different machines``, 2008.
J.C. Denard, CAS on Beam Diagnostics, ``Beam Current Monitors``, 2008
P.Forck, P.Kowina; D. Liakin, CAS on Beam Diagnostics, ``Beam Position Monitors``, 2008
Q. Luo; B.G. Sun; Z.R. Zhou; Q.K. Jia, ``Design and Cold test of Reentrant Cavity for HIs``, Proceedings of PAC2011
References

## Slide 17: Thank You
Questions???
This project has received funding from the European Union’s Horizon 2020 research and innovation programme under the Marie Skłodowska-Curie grant agreement No 675265.

## Slide 18: Potential Benefits
TM110