## Slide 1: Gas jet diagnostics for hollow electron lens
Carsten Welsch & Hao Zhang
for the
BGC Collaboration
9th HL-LHC Collaboration meeting, Fermilab, USA, 15th October 2019
Notes: Good afternoon, Everyone. My name is Hao Zhang from University of Liverpool and Cockcroft Institute. I will talk about the Beam gas curtain experimental programme.

## Slide 2
Motivation
Principle of the supersonic gas jet beam profile monitor in beam induced fluorescence (BIF) mode.
Experimental program in CI with prototype setup
Ongoing v3 setup for HEL test stand
HL-LHC-UK phase II proposal and plans
Outline
Notes: The motivation for this project is to develop a non-invasive beam profile monitor for hollow Electron lens and HLLHC. Although it is covered before, we will briefly talk about the principle of this supersonic gas jet beam profile monitor based on beam induced fluorescence mode.
Then we will discuss the progress of a prototype monitor we have built last recently as well as the experimental program.

## Slide 3
Problem: HLLHC aims to increase the luminosity by 10 times than current LHC condition. There is no mature diagnostics for continuous, non-invasive 2D profile monitoring for such condition.
Hollow e-lens (HEL) requires a diagnostic for on-line monitoring of the overlap between proton and electron beams in high magnetic field region.
A solution was developed during HL-LHC-UK phase I where it was successfully demonstrated through a prototype setup at the Cockcroft Institute that a supersonic gas jet in fluorescence mode can measure the 2D transverse beam profile online with sufficiently short integration times.
Motivation
Notes: Need to edit

## Slide 4
*Credit: GSI, http://www-bd.gsi.de/dokuwiki/doku.php
Particle Beam propagate to the jet curtain perpendicularly
Supersonic gas jet curtain with 45 degree tilted
Beam induced fluorescence using gas jet
Key: Increase gas jet density
test working gases
PROS:
  Very simple source: incoherent, isotropic, monochromatic.
  Detection is simple: imaging
CONS:
  Low light yield, noise sources (SR, losses).
Notes: First of all, we will mention a technique been used in GSI and many other places called beam induced fluorescence profile monitor. As shown in the picture, particle beams, here the ion beam, pass through a vacuum pipe, the particles will interact with the residual gas molecules, cause the molecule to either ionized or into excite mode, or both. When the excited molecules or excited ions returned to the ground level, they will emit fluorescent light, where the wavelength of the fluorescence will depend on the energy difference between the excited state and ground state. The cross section for the fluorescence is usually very low, and a fracture of light collecting angle can be used. To make this method practical, gases will be introduced into the system to make a local pressure bump to increase the fluorescence rate. However, the drawback is the beam loss and unwanted high vacuum pressure. Two independent camera system will be installed if a two dimensional profile is required.
Here, in our design, instead of pouring gas into the whole chamber, we use a supersonic curtain gas jet with transverse dimension about 5*0.5 mm2. The benefit of it will be the directionality of the jet, which means it could be easily pumped away. The curtain could be tilted 45 degree, which means a 2D measurement at the same time. The interaction length depends on the thickness of the gas jet, which could be optimized, which means a minimum invasive method.
The photon detection system is similar to one of the BIF monitor.

## Slide 5: Prototype (v2) at the Cockcroft Institute
HIPACE 300L/S
Turbo-molecular pump
HIPACE 700L/S
Turbo-molecular pump
Nozzle Chamber
Skimmer Assembly
Gas Separator
Skimmer Chamber
Imaging System
Interaction Chamber
Moveable Gauge Chamber
Dump Chamber
Third Skimmer
Nozzle Chamber
Skimmer Chamber
Interaction Chamber
Moveable Gauge Chamber
Dump Chamber
Imaging System
1.6 m
Gas jet

## Slide 6: Photon counting method
1
Identify location of each dot. Each dot is a photon (either signal or noise)
Repeat this for all frames
Integrate all frames to form a photon count matrix. The total integration time will be exposure time * frame number.
2
Exposure time: 1 s, e-beam: 5 keV, 0.65 mA
Inlet pressure: 5 bar
30 mm nozzle
3.8 mm nozzle to skimmer distance (optimum)
3
Gas jet
Electron beam
x
y

## Slide 7: Working gas test: Nitrogen gas jet
Integration time : 400 s
FWHM = 1.8 mm
FWHM = 1.5 mm
Integration time : 4 s
Filter used for Nitrogen: 391 nm with 10 nm bandwidth Number of photons per second from the jet: ~ 17
Electron beam signal is visible in just 4 s via binning.
Gas jet
Electron beam
x
y
FWHM = 1.5 mm
Integration time: 400 s
Notes: Redo the plot

## Slide 8: Neon and Argon gas-jet
Integration time: 4,000 s (+ 4,000 s for background);
Good agreement between normalised profiles of neon and nitrogen;
Filter used for Neon: 585 nm with 10 nm bandwidth
Number of photons per second from jet: 1.6 per second
Integration time: 4,000 s (+ 4,000 s for background);
Filter used for Neon: 476 nm with 10 nm bandwidth
The background was subtracted from the gas jet image
Number of photons per second from jet: 1.3 per second.
FWHM = 1.5 mm
FWHM = 1.5 mm
Notes: Redo plot

## Slide 9: Density measurements with pin-hole gauge
A scan was taken using the pin-hole opening for nitrogen, neon and argon under similar experimental conditions. (30 mm nozzle at optimum nozzle-skimmer distance and third skimmer set vertically).
Density measured by :
0.7 mm
6.3 mm
Nitrogen
Neon
Argon
Third skimmer
D: 40 mm
0
0

## Slide 10: Summary: working gases in lab condition
 | Experimental Photon Number/s | Advantages | Disadvantages
N2 N2+ emitter 391.4 nm | 17 | Higher photon rate, shorter integration time. lower dark counts rate for the photocathode Higher quantum efficiency of photocathode | Ion emitter with 60 ns lifetime, potential distortion due to space charge and external EM field on the N2+ Need further vacuum approval
Neon Ne emitter 585.4 nm | 1.6 | Fluorescence due to neutral Ne (yellow line), not affect by space charge of the primary beam and external EM field LHC vacuum compatible | Lower photon rate, longer integration time Higher dark counts rate for the photocathode Lower quantum efficiency of photocathode Background light from lab e-gun
Argon Ar+ emitter 476.5 nm | 1.3 | Short decay time, less affected by space charge or external EM field Higher quantum efficiency of photocathode | Lower photon rate, longer integration time Background light from lab e-gun
Argon Ar+ emitter 420±75 nm | 4 |  | 

## Slide 11: Estimation for HEL condition
Based on a gas jet with N=2.5e10 cm-3, d=0.5 mm.
Cross section data for 7 TeV proton beam is extrapolated from references. Tests are required for LHC condition.
<ti>mcp is the expected time to detect one photon. A few hundred photons are required for the proton beam and 104 for the HEL electron beam.
Expected integration time
  Electron: sub second
  Proton: few or few tens of seconds

## Slide 12: BIF test during Run 2
Pb+ at injection
Proton test => Noises from SR
585 nm, gas ON, beam ON
“photon counting” over 1286.4 s
(3216 frames, 400 ms exp time)
585 nm, gas OFF, beam ON
“photon counting” over 210.8 s
(527 frames, 400 ms exp time)

## Slide 13: Challenge of BGC for HEL
Gas choice (distortion due to space charge)
Light pollution (SR, stray light) => Black coating, filtering
  NEG on Cu (R = 44.3%) vs. Polyteknik black coating (R = 0.12 % @ 585 nm)
Shorten the integration time => Increase jet density.
Notes: Multilayer sputttering

## Slide 14: Compact design for gas jet generation
400 mm
3rd skimmer
Versatile for space limitation
Increase gas jet density at the interaction chamber.

## Slide 15
V3 gas jet setup for HEL test stand
Nozzle Chamber
Skimmer Chamber
Interaction Chamber
Dump Chamber
Gate valve
Gate valve
Gate valve
Imaging system
Collector solenoid
Room for beam diagnostics (BGC)
Second drift solenoid + correctors

## Slide 16: Plans for v3 gas jet for HEL test stand
Procurement ongoing
By end of March 2020: v3 gas jet system delivered and commissioned at CI;
Connected to June 2020 milestone: ARIES e-gun and modulator measurements
Solenoid

## Slide 17: HL-LHC-UK Phase II
Gas jet beam profile monitor (BGC) is strongly supported by CERN for LHC as a general non-invasive profile measurement diagnostic and the baseline instrument for on-line monitoring of the overlap between proton and electron beams for HEL;
We will carry out R&D into minimum-invasive transverse profile monitoring and deliver two supersonic gas jet-based monitors based on beam-induced fluorescence
The specific objectives are
  Build and commission two gas jet monitors; deliver the instruments to CERN for integration in HL-LHC and contribute to commissioning and operation;
  Optimize optical imaging and acquisition system through simulation and experiments;
  Fully characterize different jet formation schemes and their impact on density at interaction point;
  Study options for gas jet wire scanner and achievable resolution.

## Slide 18: HL-LHC-UK Phase II
This work is a direct continuation of our previous studies in HLLHCUK – phase I.
Large parts of the monitors will be manufactured by UK industry through established contacts from phase I. Opportunities for further commercialization of the monitor will be exploited with the help of experts from D-Beam, STFC-CERN BIC alumni.
Phase II work will make this profile monitor a versatile solution for essentially any high energy, high intensity accelerator.
Research and installation plan
  Stage 1
    Install interaction chamber assembly in LHC include blackened chamber, gate valves, supports, optical calibration target during LS2. 01/04/20-24/12/20
  Stage 2
    M3.2.1: Commission and optimize gas-jet design based on Phase I final deliverable both at CI and at HEL test stand. 01/04/20-31/11/21
    Carry out investigations into the adaptation of the current scheme to a gas jet wire scanner. 01/01/21-31/11/21
    D3.2.1: Get-jet monitor engineering design: show that proposed design meets all specifications. 31/12/21
Interaction chamber
Image Courtesy: J. Glutting

## Slide 19: HL-LHC-UK Phase II
Research and installation plan - continued
  Stage 3
    Prototype monitor installed during LHC technical stop 01/12/22-28/02/22
    M3.2.2: Finalize design and production drawings. 01/03/22-30/06/22
    Review of gas-jet drawings to launch production completed 01/03/22-30/06/22
    Prototype monitor commissioning and study of 7 TeV proton fluorescence 01/03/22-31/08/22
    D3.2.2: Final design: report with final drawings, integration and commissioning plan 01/09/22-31/10/22
  Stage 4
    Manufacture interaction chambers and delivery to CERN for HEL setup 30/06/22-28/03/23
    D3.2.3: Delivery of interaction chambers for integration in HEL 31/03/23
    M3.2.3: Orders placed for major gas-jet unit 1 components:
      Manufacture and delivery of gas jet monitor unit 1 28/03/23-26/03/24
      Manufacture and delivery of gas jet monitor unit 2 26/03/24-17/12/24
    D3.2.4: Delivery of gas-jet monitor unit 1, pre-tested at CI, for integration in HEL. Participation in commissioning 28/03/23-31/03/2024
    D3.2.5: Delivery of gas-jet monitor unit 2 for integration at CERN, pre-tested at CI 31/12/2024
    M3.2.4: Production of gas-jets completed and shipped to CERN 02/01/25
Preliminary image of BGC integrated with LHC.
Image Courtesy: J. Glutting

## Slide 20: HEL integration with gas jet monitor
Main solenoid
Main solenoid
Gas jet monitor
Space for BGC
GAP for mounting
~ 200 mm

## Slide 21: Summary
In HL-LHC-UK phase I, we successfully demonstrated 2D minimum invasive transverse profile monitor based on beam-induced fluorescence, using supersonic gas jet;
Various options of gases were studied in terms of achievable gas jet density and integration time;
Key technologies required to produce a working instrument in the LHC have been identified and development is well under-way
V3 monitor is being progressed and will be tested in HEL test stand;
For HL-LHC-UK phase II, primary goal will be delivery of two working monitors, but R&D into gas jet wire, BGV and monitor simplification will continue.

## Slide 22: Thanks for your attention!
Work presented on behalf of the BGC collaboration:
J. Resta Lopez, N. Kumar, A. Salehilashkajani,
C.P. Welsch, H. Zhang (U Liverpool / Cockcroft Institute)
P. Forck, S. Udrea (GSI)
M. Ady, N. Chritin, T. Doddington, J. Glutting, R. Jones, R. Kersevan, S. Mazzoni,
A. Rossi, G. Schneider, R. Veness, (CERN BE-BI, TE-VSC, EN-MME)
9th HL-LHC Collaboration meeting, 15th October 2019