## Slide 1: MQXFAP1b Alignment and Field Quality
J. DiMarco, D. Cheng, B. Ghiorso, C. Hernikl, C. Myers, H. Pan, G. L. Sabbi, H. Song, P. Wanderer, X. Wang, H. Zhu
HiLumi Collaboration Meeting, Fermilab, October 16, 2019

## Slide 2: Magnetic Measurements during Assembly
MQXFAP1b Magnet geometry and fabrication:
Shorter magnetic length (4 m)
Three previously used first-generation coils
One new coil made with shorter length but incorporating many new features
  See coil P06 review information at https://indico.fnal.gov/event/19381/
  Mid-plane/pole shim for b6 correction still not implemented
Magnetic measurement system and plan
Measurements performed for coil pack and after complete assembly
  MQXFAP2: drive system was not complete for coil pack measurement, and splice box was not completed for “final” measurement
New power supply (more stable, better controls, capability to switch polarity)
Field harmonics as a function of longitudinal position
  Including a detailed scan of the ends
Magnetic axis vs. longitudinal position relative to external fiducials
Relative change in the field orientation vs. longitudinal position
  Including a scan step allowing to overlap the positions of the two probes mounted on the PCB

## Slide 3: Magnetic Measurement System at LBNL
Rotating probe includes two identical PCBs from FNAL with 108.74 mm effective length and 59.5 mm radius
Probe support tube
(inside magnet)
Rotating
probe
Drive
system
MQXFAP2
Fiducial
locations
PCB
Retroreflectors
X. Wang, J. DiMarco, D. Cheng, J. Herrera, S. Myers, B. Ghiorso, A. Pekedis et al.
Probe is inserted from lead end
Return end

## Slide 4: Longitudinal scan: Transfer Function
TF increases by 0.4% from coil pack to final assembly
  Effect of loading (iron pads are already present in coil pack measurement)
Scan length is sufficient to cover both ends – however this is a shorter magnet, may need some improvement for full length magnet

## Slide 5: Longitudinal Scan: b6
b6 is around -2 units in the straight section (about 5 units smaller than AP2)
Also much more stable along the length (about +/- 1 unit)
Change of about 1 unit (-3 to -2) from coil pack to final.
Fine (1/4) scan of magnet ends allows to accurately capture the peaks
Useful to compare with models and for longitudinal reference
Will be repeated in vertical test

## Slide 6: Comparison of MQXFAP1b and MQXFAP2
b6 is smaller and more stable along the length
Zoom in

## Slide 7: Longitudinal scan – other harmonics
Change of 2 units of b3 from coil pack to final. All other harmonics do not change significantly
Comparatively large a5

## Slide 8: Comparison of harmonics after assembly

## Slide 9: Reference axis and probe position survey
Survey performed by Chris Hernikl, Hongyan Zhu, Dan Ellis, and Federico Carrara (Survey & Alignment Group, Engineering Division)

## Slide 10: Magnetic Axis
Magnetic center relative to reference axis
The magnetic axis will be defined as the best fit, independently for x and y
Variations of magnetic center relative to best fit are within +/- 0.2 mm (straight section)
Requirement is +/- 0.5 mm
Additional benefit from integration length (0.5 m)

## Slide 11: Magnetic Field Angle
Magnetic field angle along the magnet length
Magnet angle can be adjusted for best fit
Variations of magnetic field angle relative to best fit are within +/- 1 mrad
Requirement is +/- 2 mrad
Ends are excluded
Additional benefit from integration length (0.5 m)

## Slide 12: Magnetic Measurements System at BNL

## Slide 13: Magnetic Measurements during Vertical Test
MQXFAP1b quench history, warm and cold magnetic field measurement time
Stair-step measurement

## Slide 14: Warm Measurements during Vertical Test
Comparison between warm measurements during assembly at LBNL and prior to vertical test at BNL are generally in good agreement
No indications of changes during shipping, small differences are mainly due to different probe length
New identical 440 mm probes from FNAL are under fabrication for future tests

## Slide 15: Cold Measurements during Vertical Test
Decrease on the transfer function from injection to nominal by ~ 9 % due to iron saturation, consistent with calculations and results from other model magnets and prototypes
Transfer function

## Slide 16: Cold Measurements during Vertical Test
Similar longitudinal variations at different temperature/current, with some offset (0-2 units)

## Slide 17: Cold Measurements during Vertical Test
Upper bound: s + u + 3s
Lower bound: s - u - 3s

## Slide 18: MQXFAP1b Calculated Harmonics
MQXFAP1 and MQXFAP1b included both 1st and 2nd generation coils
One coil first generation (Quadrant 1 for the computations)
Three coils second generation
Calculated field errors at nominal current (16.47 kA):

## Slide 19: Cold Measurements during Vertical Test
Relative shift of magnetic vs probe axis up to ~12 mm was observed
Calculated - assuming measured dipole stems from probe offset in pure quad field
Significant changes in value and pattern from warm to cold measurements
About a factor of two change from low to high current
Addressed with additional alignment features for the warm bore between AP2 and AP1b
X
Y
Z
I
Cold AP1b
Cold AP1b
Cold AP1b
Cold AP1b

## Slide 20
Additional slides

## Slide 21: Field Quality Reference Table

## Slide 22: Reference Frame for Magnetic Measurements
C102
C106
C105
C104
Longitudinal axis points from the return to the lead end
z = 0 is defined as the magnetic center based on the integrated quadrupole strength
Consider defining z=0 based on a specific geometric location/drawing
Also need to introduce in the quadrant/coil diagram a geometric feature (e.g. main leads) to clarify angular reference

## Slide 23: Quadrupole reference frame and powering
Positive or negative quadrupole depending on powering
Positive or negative depending on rotation (switches every 90 degrees)
Positive or negative quadrupole depending on direction of travel
Two magnets will be oriented opposite to each other in the cold mass
Next steps:
Confirm quadrupole sign for given powering orientation at LBNL and BNL
Some issue at LBNL comparing AP2 with AP1b:opposite quadrupole sign for nominally same powering
But, AP2 measurement was performed with no splice box (coils connected manually)

## Slide 24: MQXFAP1b Calculated Harmonics
MQXFAP1 and MQXFAP1b included both 1st and 2nd generation coils
One coil first generation in Quadrant 3
Three coils second generation
Calculated field errors at nominal current (16.47 kA):
2G
2G
2G
1G

## Slide 25: Cold Measurements during Vertical Test
Upper bound: s + u + 3s
Lower bound: s - u - 3s