## Slide 1
2nd part of this lecture covers:
Transverse profile techniques
Emittance determination at transfer lines
Diagnostics for bunch shape determination
Beam Instrumentation & Diagnostics Part 2
CAS Introduction to Accelerator Physics
Vysoké Tatry, 19thof September 2019
Peter Forck
Gesellschaft für Schwerionenforschnung (GSI)
p.forck@gsi.de

## Slide 2
Measurement of Beam Profile
The beam width can be changed by focusing via quadruples.
Transverse matching between ascending accelerators is done by focusing.
→ Profiles have to be controlled at many locations.
Synchrotrons: Lattice functions  (s) and D(s) are fixed  width  and emittance  are:
Transfer lines: Lattice functions are ‘smoothly’ defined due to variable input emittance.
A great variety of devices are used:
Optical techniques: Scintillating screens (all beams),
synchrotron light monitors (e−), optical transition radiation (e−, high energetic p),
ionization profile monitors (protons)
 Electronics techniques: Secondary electron emission SEM grids, wire scanners (all)
Typical beam sizes:
e−-beam: typically Ø 0.01 to 3 mm, protons: typically Ø 1 to 30 mm
Notes: bla

## Slide 3
Scintillation Screen
Scintillation: Particle’s energy loss in matter causes emission of light
the most direct way of profile observation as used from the early days on!
Pneumatic feed-through
with Ø70 mm screen:
Flange
200 mm
& window
Screen
70 mm
beam
Pneumatic
drive
CCD
1 m

## Slide 4
Example of Screen based Beam Profile Measurement
Observation with
a CCD, CMOS or video camera
Advantage of screens:
Direct 2-dim measurement
High spatial resolution
Cheap realization
 widely used at transfer lines
Disadvantage of screens:
Intercepting device
Some material might brittle
Low dynamic range
Might be destroyed
by the beam
b/w CCD:
artificial false-color
Example: GSI LINAC, 4 MeV/u, low current, YAG:Ce screen
LINAC
LINAC
Scintillation Screen (beam stopped)

## Slide 5
Alumina: Al2O3
CsI:Tl
Chromox: Al2O3:Cr
P43
YAG:Ce
Quartz
Quartz:Ce
ZrO2:Mg
Very different light yield i.e. photons per ion‘s energy loss
Different wavelength of emitted light
Light output from various Scintillating Screens
Example: Color CCD camera: Images at different particle intensities determined for U at 300 MeV/u

## Slide 6
Material Properties for Scintillating Screens
Some materials and their basic properties:
Properties of a good scintillator:
Large light output at optical wavelength
 standard CCD camera can be used
Large dynamic range  usable for different currents
Short decay time  observation of variations
Radiation hardness  long lifetime
Good mechanical properties  typ. size up to Ø 10 cm
(Phosphor Pxx grains of Ø  10 μm on glass or metal).
Name | Type | Material | Activ. | Max. λ | Decay
Chromox | Cera- mics | Al2O3 | Cr | 700 nm |  10 ms
Alumina |  | Al2O3 | Non | 380 nm |  10 ns
YAG:Ce | Crystal | Y3Al5O12 | Ce | 550 nm | 200 ns
P43 | Powder | Gd2O3S | Tb | 545 nm | 1 ms
P46 |  | Y3Al5O12 | Ce | 530 nm | 300 ns
P47 |  | Y3Si5O12 | Ce&Tb | 400 nm | 100 ns
Standard drive with P43 screen
Flange
200 mm
& window
Screen
70 mm
beam
Pneumatic
drive
CCD
1 m

## Slide 7
Outline:
Scintillation screens:
emission of light, universal usage, limited dynamic range
SEM-Grid: emission of electrons, workhorse, limited resolution
Wire scanner
Ionization Profile Monitor
Optical Transition Radiation
Synchrotron Light Monitors
Summary
Measurement of Beam Profile

## Slide 8
8
Secondary Electron Emission by Ion Impact
Energy loss of ions in metals close to a surface:
Closed collision with large energy transfer:  fast e- with Ekin >> 100 eV
Distant collision with low energy transfer :  slow e- with Ekin  10 eV
‘diffusion’ & scattering with other e-: scattering length Ls  1 - 10 nm
at surface  90 % probability for escape
Secondary electron yield and energy distribution comparable for all metals!
 Y = const. * dE/dx (Sternglass formula)
beam
Ls  10 nm
e-
e-
δ-ray
Electrons per ion
Different targets:
From E.J. Sternglass, Phys. Rev. 108, 1 (1957)
Notes: 8

## Slide 9
Secondary Electron Emission Grids = SEM-Grid
Beam surface interaction: e− emission → measurement of current.
SEM-Grid feed-through on CF200:
Example: 15 wire spaced by 1.5 mm:
Parameter | Typ. value
# wires per plane | 10 ...100
Active area | (5...20 cm) 2
Wire  | 25....100 m
Spacing | 0.3...2 mm
Material | e.g. W or Carbon
Max. beam power | 1 W/mm
5 cm
 1 m
SEM-Grid
wire
feedthru
beam

## Slide 10
Secondary Electron Emission Grids = SEM-Grid
Each wire is equipped with one I/U converter
different ranges settings by Ri
 very large dynamic range up to 106.
Example: 15 wire spaced by 1.5 mm:
5 cm
Beam surface interaction: e− emission → measurement of current.

## Slide 11
Example of Profile Measurement with SEM-Grids
Even for low energies, several SEM-Grid can be used due to the  80 % transmission
 frequently used instrument beam optimization: setting of quadrupoles, energy….
Example: C6+ beam of 11.4 MeV/u at different locations at GSI-LINAC
injection
extraction
SEM-Grids
synchrotron
horizontal
vertical
beam

## Slide 12
The Artist view of a SEM-Grid = Harp

## Slide 13
Outline:
Scintillation screens:
emission of light, universal usage, limited dynamic range
SEM-Grid: emission of electrons, workhorse, limited resolution
Wire scanner: emission of electrons, workhorse, scanning method
Ionization Profile Monitor
Optical Transition Radiation
Synchrotron Light Monitors
Summary
Measurement of Beam Profile

## Slide 14
Slow, linear Wire Scanner
Idea: One wire is scanned through the beam!
Wire diameter 100 µm < dwire < 10 µm
Slow, linear scanner are used for:
Low energy protons
High resolution measurements for e− beam
by de-convolution σ2beam=σ2meas−d2wire
 resolution down to 1 μm range can be reached
Detection of beam halo
5 cm

## Slide 15
The Artist view of a Beam Scraper or Scanner

## Slide 16
Fast, Flying Wire Scanner
In a synchrotron one wire is scanned though the beam as fast as possible.
Fast pendulum scanner for synchrotrons; sometimes it is called ’flying wire’:
From https://twiki.cern.ch/twiki/
bin/viewauth/BWSUpgrade/
injection
extraction
particle
detector
synchrotron

## Slide 17
Usage of Flying Wire Scanners
Material: carbon or SiC → low Z-material for low energy loss and high temperature.
Thickness: down to 10 μm → high resolution.
Detection: High energy secondary particles with a detector like a beam loss monitor
Proton impact on
scanner at CERN-PS Booster:
Secondary particles:
Proton beam → hadrons shower (π, n, p...)
Electron beam → Bremsstrahlung photons.
Rest mass:
m± = 140 MeV/c2
m0 = 135 MeV/c2
Kinematics of flying wire:
Velocity during passage typically 10 m/s = 36 km/h and
typical beam size  10 mm  time for traversing the beam t  1 ms
Challenges: Wire stability for fast movement with high acceleration
U. Raich et al., DIPAC 2005

## Slide 18
The Artist View of a Wire Scanner

## Slide 19
Comparison between SEM-Grid and slow Wire Scanners
Grid: Measurement at a single moment in time
Scanner: Fast variations can not be monitored
 for pulsed LINACs precise synchronization is needed
__________________________________________________________________________
Grid: Not adequate at synchrotrons for stored beam parameters
Scanner: At high energy synchrotrons flying wire scanners are nearly non-destructive
__________________________________________________________________________
Grid: Resolution of a grid is fixed by the wire distance (typically 1 mm)
Scanner: For slow scanners the resolution is about the wire thickness (down to 10 μm)
 used for e−-beams having small sizes (down to 10 μm)
__________________________________________________________________________
Grid: Needs one electronics channel per wire
→ expensive electronics and data acquisition
Scanner: Needs a precise movable feed-through → expensive mechanics.

## Slide 20
Outline:
Scintillation screens:
emission of light, universal usage, limited dynamic range
SEM-Grid: emission of electrons, workhorse, limited resolution
Wire scanner: emission of electrons, workhorse, scanning method
Ionization Profile Monitor:
secondary particle detection from interaction beam-residual gas
Optical Transition Radiation
Synchrotron Light Monitors
Summary
Measurement of Beam Profile

## Slide 21
Ionization Profile Monitor at GSI Synchrotron
Non-destructive device for proton synchrotron:
beam ionizes the residual gas by electronic stopping
gas ions or e- accelerated by E -field 1 kV/cm
spatial resolved single particle detection
Realization at GSI synchrotron:
One monitor per plane
Typical vacuum pressure:
Transfer line: N2 10−8...10−6 mbar  3108...31010cm-3
Synchrotron: H2 10−11...10−9 mbar  3105...3107cm-3
R
voltage divider
+ 6 kV
+ 5 kV
+ 4 kV
+ 2 kV
+ 3 kV
+ 1 kV
0 kV
MCP
anode
position readout
HV electrode
ion
e.g. H+
beam

## Slide 22
Ionization Profile Monitor Realization
The realization for the heavy ion storage ring ESR at GSI:
Horizontal camera
Horizontal IPM:
E-field box
MCP
IPM support
& UV lamp
Ø250 mm
beam
E-field separation disks
View port Ø150 mm
electrodes
Insertion 650 mm
175mm
Vertical IPM
Vertical camera
Realization at GSI synchrotron:
One monitor per plane

## Slide 23
Ionization Profile Monitor Realization
 IPM measure
Horizontal camera
Horizontal IPM:
E-field box
MCP
IPM support
& UV lamp
Ø250 mm
beam
E-field separation disks
View port Ø150 mm
electrodes
Insertion 650 mm
175mm
Vertical IPM
Vertical camera
Realization at GSI synchrotron:
One monitor per plane
The realization for the heavy ion storage ring ESR at GSI:

## Slide 24
The emittance is defined via the position deviation and angle in lab-frame
After acceleration the longitudinal velocity is increased  angle  is smaller
The angle is expressed in momenta: x’ = p / p|| the emittance is <xx’> = 0 :  = x  x’ = x  p / p||
under ideal conditions the emittance can be normalized to the momentum p || =   m  c
normalized emittance norm =    is preserved with the Lorentz factor  and velocity  =v/c
Example: Acceleration in GSI-synchrotron for C6+ from
6.7  600 MeV/u ( = 12  79 %) observed by IPM
theoretical width:
measured width:
‘Adiabatic’ Damping during Acceleration
slow
v
v||
v
before acceleration
fast
v
after acceleration
v
v||
IPM is well suited
for long time observations
without beam disturbance
 mainly used at proton synchrotrons.
acc.
injection
extraction
IPM
synchrotron with acceleration
 OTR

## Slide 25
Outline:
Scintillation screens:
emission of light. universal usage, limited dynamic range
SEM-Grid: emission of electrons, workhorse, limited resolution
Wire scanner: emission of electrons, workhorse, scanning method
Ionization Profile Monitor:
secondary particle detection from interaction beam-residual gas
Optical Transition Radiation:
crossing material boundary, for relativistic beams only
Synchrotron Light Monitors
Summary
Measurement of Beam Profile

## Slide 26
Optical Transition Radiation: Depictive Description
Optical Transition Radiation OTR for a single charge e:
Assuming a charge e approaches an ideal conducting boundary e.g. metal foil
image charge is created by electric field
dipole type field pattern
field distribution depends on velocity  and Lorentz factor  due to relativistic trans. field increase
penetration of charge through surface within t < 10 fs: sudden change of source distribution
emission of radiation with dipole characteristic
perfect
metal
vacuum
charge e
velocity 
E-field
pattern of
dipole type
image
charge -e
velocity -
perfect
metal
vacuum
charge e
velocity 
image
charge -e
velocity -
perfect
metal
vacuum
charge e
inside
metal
‘dipole
radiation’

 max 1/
sudden change charge distribution
rearrangement of sources  radiation
Other physical interpretation: Impedance mismatch at boundary leads to radiation


## Slide 27
Optical Transition Radiation OTR can be described in classical physics:
approximated formula
for normal incidence
& in-plane polarization:
Angular distribution of radiation in optical spectrum:
lope emission pattern depends on velocity or Lorentz factor 
peak at angle   1/
emitted energy i.e. amount of photons scales with W   2
broad wave length spectrum (i.e. no dependence on )
 suited for high energy electrons
Optical Transition Radiation: Depictive Description
W: radiated energy
: frequency of wave
perfect
metal
vacuum
‘dipole
radiation’

 max 1/

charge e
inside
metal
sudden change charge distribution
rearrangement of sources  radiation

## Slide 28
Technical Realization of Optical Transition Radiation OTR
OTR is emitted by charged particle passage through a material boundary.
Insertion of thin Al-foil under 45o
Observation of low light by CCD.
Photon distribution:
within a solid angle d and
Wavelength interval begin to end
Detection: Optical 400 nm <  < 800 nm
using image intensified CCD
Larger signal for relativistic beam  >> 1
Low divergence for  >> 1  large signal
well suited for e- beams
p-beam only for Ekin > 10 GeV   > 10

## Slide 29
OTR-Monitor: Technical Realization and Results
Example of realization at TERATRON:
Insertion of foil
e.g. 5 m Kapton coated with 0.1m Al
Advantage: thin foil  low heating & straggling
2-dim image visible
Courtesy V.E. Scarpine (FNAL) et al., BIW’06
 = 0.66 mm
 = 1.03 mm
rad-hard
camera
Beam pipe
Window
Filter wheel
Lens
Results at FNAL-TEVATRON synchrotron
with 150 GeV proton
Using fast camera: Turn-by-turn measurement

## Slide 30
Optical Transition Radiation compared to Scintillation Screen
Installation of OTR and scintillation screens on same drive:
Courtesy of U. Iriso et al., DIPAC’09
Example: ALBA LINAC 100 MeV
Results:
Much more light from YAG:Ce
for 100 MeV ( =200) electrons
light output IYAG  10 5 IOTR
Broader image from YAG:Ce
due to finite YAG:Ce thickness
OTR
YAG:Ce
OTR
YAG:Ce
projection
OTR projection

## Slide 31
Comparison between Scintillation Screens and OTR
OTR: electrodynamic process → beam intensity linear to # photons, high radiation hardness
Scint. Screen: complex atomic process → saturation possible, for some low radiation hardness
OTR: thin foil Al or Al on Mylar, down to 0.25 μm thickness
→ minimization of beam scattering (Al is low Z-material e.g. plastics like Mylar)
Scint. Screen: thickness  1 mm inorganic, fragile material, not always radiation hard
OTR: low number of photons → expensive image intensified CCD
Scint. Screen: large number of photons → simple CCD sufficient
OTR: complex angular photon distribution → resolution limited
Scint. Screen: isotropic photon distribution → simple interpretation
OTR: large γ needed → e−-beam with Ekin > 100 MeV, proton-beam with Ekin > 100 GeV
Scint. Screen: for all beams
Remark: OTR not suited for LINAC-FEL due to coherent light emission (not covered here)
but scintillation screens can be used.

## Slide 32
Outline:
Scintillation screens:
emission of light, universal usage, limited dynamic range
SEM-Grid: emission of electrons, workhorse, limited resolution
Wire scanner: emission of electrons, workhorse, scanning method
Ionization Profile Monitor:
secondary particle detection from interaction beam-residual gas
Optical Transition Radiation:
crossing optical boundary, for relativistic beams only
Synchrotron Light Monitors
photon detection of emitted synchrotron light in optical and X-ray range
Summary
Measurement of Beam Profile

## Slide 33
Synchrotron Light Monitor
An electron bent (i.e. accelerated) by a dipole magnet emit synchrotron light
see lecture of Lenny Rivkin
This light is emitted
into a cone of
opening 2/ in lab-frame.
Well suited for rel. e-
For protons:
Only for energies Ekin > 100 GeV
The light is focused to a intensified CCD.
Advantage:
Signal anyhow available!
orbit of electrons
orbit of electrons
dp/dt
dp/dt
radiation field
radiation field
power: P  4/  2
Rest frame of electron:
Laboratory frame:

## Slide 34
Realization of a Synchrotron Light Monitor
Extracting out of the beam’s plane by a (cooled) mirror
→ Focus to a slit + wavelength filter for optical wavelength
→ Image intensified CCD camera
Example: ESRF monitor from dipole with bending radius 22 m (blue or near UV)
injection
extraction
SRM
synchrotron
 beam
dipole
optical table
e- beam
Courtesy K. Scheidt et al., DIPAC 2005

## Slide 35
Result from a Synchrotron Light Monitor
Example: Synchrotron radiation facility APS accumulator ring and blue wavelength:
Advantage: Direct measurement of 2-dim distribution, good optics for visible light
Realization: Optics outside of vacuum pipe
Disadvantage: Resolution limited by the diffraction due to finite apertures in the optics.
B.X. Yang (ANL) et al. PAC’97
injection
extraction
SRM
synchrotron

## Slide 36
‘Adiabatic Damping’ for an Electron Beam
The beam emittance in influenced by:
Adiabatic damping
Longitudinal momentum contribution
via dispersion
total width
Quantum fluctuation due to light emission
Example: Booster at the light source ALBA acceleration from 0.1  3 GeV within 130 ms
Courtesy U. Iriso & M. Pont (ALBA) et al. IPAC 2011
vert. y [mm]
hor. x [mm]
injection
extraction
SRM
synchrotron
Profile measure by synchrotron light monitor:

## Slide 37
The Artist View of a Synchrotron Light Monitor

## Slide 38
Diffraction Limit of Synchrotron Light Monitor
Limitations:
Diffraction limits the resolution
due to Fraunhofer diffraction
Improvements:
Shorter wavelength:
Using X-rays and an aperture of Ø 1mm
→ ‘X-ray pin hole camera’,
achievable resolution   10 m
Interference technique:
At optical wavelength
using a double slit
→ interference fringes
achievable resolution   1 m.
with wavelength 

## Slide 39
X-ray Pin-Hole Camera
The diffraction limit is  shorter wavelength by X-rays.
Example: PETRA III
Courtesy K. Wittenburg, DESY
X-ray optics
CCD
Monochromator
electron beam
X-ray beam
 = 44 µm
 = 32µm
Example: PETRA III result:

## Slide 40
Double Slit Interference for Radiation Monitors
The blurring of interference pattern due to
finite size of the sources
spatial coherence parameter  delivers rms beam size
i.e. ‘de-convolution’ of blurred image!
 highest resolution, but complex method
Typical resolution for three methods:
Direct optical observation:   100 µm
Direct x-ray observation :   10 µm
Interference optical obser:   1 µm
Courtesy of V. Schlott PSI
Ideal double slit interference pattern:
x
P(x)
Blurring by finite source size:
x
P(x)

## Slide 41
Summary for Beam Profile Measurement
Different techniques are suited for different beam parameters:
e−-beam: typically Ø 0.01 to 3 mm, protons: typically Ø 1 to 30 mm
Intercepting  non-intercepting methods
Direct observation of electrodynamics processes:
Optical synchrotron radiation monitor: non-destructive, for e−-beams, complex, limited res.
X-ray synchrotron radiation monitor: non-destructive, for e−-beams, very complex
OTR screen: nearly non-destructive, large relativistic γ needed, e−-beams mainly
Detection of secondary photons, electrons or ions:
Scintillation screen: destructive, large signal, simple setup, all beams
Ionization profile monitor: non-destructive, expensive, limited resolution, for protons
Wire based electronic methods:
SEM-grid: partly destructive, large signal and dynamic range, limited resolution
Wire scanner: partly destructive, large signal and dynamics, high resolution, slow scan.

## Slide 42
Measurement of transverse Emittance
The emittance characterizes the whole beam quality, assuming linear
behavior as described by second order differential equation.
It is defined within the phase space as:
The measurement is based on determination of:
Either profile width σx and angular width σx′ at one location
Or profile width σx at different locations and linear transformations.
Synchrotron: lattice functions results in stability criterion
 beam width delivers emittance:
Different devices are used at transfer lines:
Lower energies Ekin < 100 MeV/u: slit-grid device, pepper-pot
(suited in case of non-linear forces).
All beams: Quadrupole variation, ’three grid’ method using linear transformations
(not well suited in the presence of non-linear forces)

## Slide 43
Trajectory and Characterization of many Particles
Single particle trajectories are forming a beam
They have a distribution of
start positions and angles
Characteristic quantity is
the beam envelope
Goal:
Transformation of envelope
behavior of whole ensemble
see lecture of Wolfgang Hillert
Courtesy K.Wille
Focus. quad.
Focus. quad.
drift
drift
envelope of all particles
single particles
trajectory

## Slide 44
Definition of Coordinates and basic Equations
The basic vector is 6 dimensional:
The transformation of a single particle from a location s0 to s1 is given by the Transfer Matrix R:
The transformation of a the envelope from a location s0 to s1 is given by the
Beam Matrix  :
6-dim Beam Matrix with decoupled hor., vert. and long. plane:
Horizontal
beam matrix:
Beam width for
the three
coordinates:
horizontal
vertical
longitudinal
hor.-long. coupling
 9 values

## Slide 45
The Emittance for Gaussian and non-Gaussian Beams
The beam distribution can be non-Gaussian, e.g. at:
 beams behind ion source
space charged dominated beams at LINAC & synchrotron
cooled beams in storage rings
General description of emittance
using terms of 2-dim distribution:
It describes the value for 1 standard derivation
Variances
Covariance
i.e. correlation
Care:
No common definition
of emittance concerning
the fraction f
For Gaussian beams only: εrms ↔ interpreted as area containing a fraction f of ions:
Emittance (f) | Fraction f
1  rms | 15 %
  rms | 39 %
2  rms | 63 %
4  rms | 86 %
8  rms | 98 %

## Slide 46
The Slit-Grid Measurement Device
Slit-Grid: Direct determination of position and angle distribution.
Used for protons with Ekin < 100 MeV/u  range R < 1 cm.
Slit: position P(x) with typical width: 0.1 to 0.5 mm
Distance: typ. 0.5 to 5 m (depending on beam energy 0. 1 ... 100 MeV)
SEM-Grid: angle distribution P(x′)
LINAC
RFQ
SEM-Grid
Slit
0.5...5 m

## Slide 47
Display of Measurement Results
The distribution is depicted as a function of
position [mm] & angle [mrad]
The distribution can be visualized by
Mountain plot
Contour plot
Calc. of 2nd moments <x2> , <x’2> & <xx’>
Emittance value εrms from
 Problems:
Finite binning results in limited resolution
Background  large influence on <x2>, <x’2> and <xx’>
Or fit of distribution with an ellipse
 Effective emittance only
Beam: Ar4+, 60 keV, 15 μA
at Spiral2 Phoenix ECR source.
P. Ausset, DIPAC 2009
Remark: Behind a ion source the beam might very non-Gaussian
due to plasma density and aberration at quadrupoles

## Slide 48
Outline:
Definition and some properties of transverse emittance
Slit-Grid device: scanning method
scanning slit  beam position & grid  angular distribution
Quadrupole strength variation and position measurement
emittance from several profile measurement and beam optical calculation
Measurement of transverse Emittance

## Slide 49
Emittance Measurement by Quadrupole Variation
From a profile determination, the emittance can be calculated via linear transformation,
if a well known and constant distribution (e.g. Gaussian) is assumed.
Measurement of beam width
x2max = σ11(1, k)
matrix R(k) describes the focusing.
With the drift matrix the transfer is
R(ki ) = Rdrift  Rfocus(ki )
Transformation of the beam matrix
(1,ki) = R(ki )  (0)  RT (ki )
Task: Calculation of (0)
at entrance s0 i.e. all three elements

## Slide 50
Measurement of transverse Emittance
Example: Square of the beam width at
ELETTRA 100 MeV e- Linac, YAG:Ce:
Using the ‘thin lens approximation’ i.e. the quadrupole has a focal length of f:
Measurement of the matrix-element (1,K) from (1,K) = R(K)∙(0)∙RT(K)
G. Penco (ELETTRA) et al., EPAC’08
Focusing strength K [m-1]
For completeness: The relevant formulas
The three matrix elements at the quadrupole:

## Slide 51
Summary for transverse Emittance Measurement
Emittance is the important quantity for comparison to theory.
It includes size (value of ε) and orientation in phase space (ij or α,  and γ)
three independent values
assuming no coupling between horizontal, vertical and longitudinal planes
Transfer line, low energy beams → direct measurement of x- and x′-distribution:
 Slit-grid: movable slit → x-profile, grid → x′-profile
Transfer line, all beams → profile measurement + linear transformation:
 Quadrupole variation: one location, different setting of a quadrupole
Assumptions:  well aligned beam, no steering
 no emittance blow-up due to space charge
Remark: non-linear transformation possible via tomographic reconstruction
Important remark: For a synchrotron with a stable beam storage,
width measurement is sufficient using

## Slide 52
Measurement of longitudinal parameter:
Bunch length measurement at
Synchrotron light sources
Linear light sources
Summary
Longitudinal  transverse correspondences:
position relative to rf  transverse center-of-mass
bunch structure in time  transverse profile
momentum or energy spread  transverse divergence
longitudinal emittance  transverse emittance.
Measurement of longitudinal Parameters

## Slide 53
Bunch Length Measurement for relativistic Electrons
Electron bunches are too short (t < 100 ps) to be covered by the bandwidth of
pick-ups (f < 3 GHz  trise > 100 ps) for structure determination.
 Time resolved observation of synchr. light with a streak camera: Resolution  1 ps.
injection
extraction
Streak camera
e- synchrotron
acc.
freq. frf
rf cavity
Scheme of a streak camera

## Slide 54
Technical Realization of a Streak Camera
Hardware of a streak camera
Time resolution down to 0.5 ps:
acceleration
focusing
deflection
 30 cm
Input optics
Streak tube
CCD
camera
 60 cm

## Slide 55
Technical Realization of a Streak Camera
Hardware of a streak camera
Time resolution down to 0.5 ps:
acceleration
focusing
deflection
 30 cm
Input optics
Streak tube
CCD
camera
 60 cm

## Slide 56
Results of Bunch Length Measurement by a Streak Camera
The streak camera delivers a fast scan in vertical direction (here 360 ps full scale)
and a slower scan in horizontal direction (24 μs).
Example: Bunch length at the synchrotron light source SOLEIL for Urf = 2 MV
for slow direction 24 μs and scaling for fast scan 360 ps: measure t = 35 ps.
Slow Scan:
Fast Scan:
bunch length
2t = 70 ps
Short bunches are desired by
the users
Example: Bunch length t
as a function of stored current
(i.e. space charge de-focusing)
at SOLEIL
Courtesy of M. Labat et al., DIPAC’07
injection
extraction
Streak camera
e- synchrotron
acc.
freq. frf
rf cavity

## Slide 57
The Artist View of a Streak Camera
 conclusion

## Slide 58
Bunch Length Measurement by electro-optical Method
For Free Electron Lasers  bunch length below 1 ps is achieved
Below the resolution of streak camera
Short laser pulses with t  10 fs and electro-optical modulator
Electro optical modulator: birefringent, rotation angle depends on external electric field
Relativistic electron bunches: transverse field E, lab = γE, rest carries the time information
Scanning of delay between bunch and laser  time profile after several pulses.
From S.P.Jamison et al., EPAC 2006
LINAC
LINAC
EO monitor
Undulator
Bunch compressor

## Slide 59
Bunch Length Measurement by electro-optical Method
For Free Electron Lasers  bunch length below 1 ps is achieved
Short laser pulse  broad frequency spectrum (property of Fourier transformation)
Optical stretcher: Separation of colors by different path length  single-shot observation
Courtesy S.P.Jamison et al., EPAC 2006
LINAC
LINAC
EO monitor
Undulator
Bunch compressor

## Slide 60
Realization of EOS Scanning
Setup of a scanning EOS method.
Delay line
X. Yan et al, Phys. Rev. Lett. 85, 3404 (2000)
Using 12fs pulses from
Ti:Al2O3 laser at 800nm and
ZnTe crystal 0.5mm thick
with a e- - beam 46MeV of 200pC

## Slide 61
Hardware of a compact EOS Scanning Setup
B. Steffen et al, DIPAC 2009
B. Steffen et al., Phys. Rev. AB 12, 032802 (2009)
crystal
beam
Example: Bunch length at FLASH
100 fs bunch duration = 30 µm length
2 = 110 fs

## Slide 62
Summary of longitudinal Measurements
Devices for bunch length at light sources:
Streak cameras:
 Time resolved monitoring of synchrotron radiation
→ for relativistic e−-beams, tbunch < 1 ns
reason: too short bunches for rf electronics.
Laser scanning:
 Electro-optical modulation of short laser pulse
→ very high time resolution down to some fs

## Slide 63
Conclusion for Beam Diagnostics Course
Diagnostics is the ’sensory organ’ for the beam.
It required for operation and development of accelerators
Several categories of demands leads to different installations:
Quick, non-destructive measurements leading to a single number or simple plots
Complex instrumentation used for hard malfunction and accelerator development
Automated measurement and control of beam parameters i.e. feedback
The goal and a clear interpretation of the results is a important design criterion.
General comments:
Quite different technologies are used, based on various physics processes
Accelerator development goes parallel to diagnostics development
Thank you for your attention!

## Slide 64
General Reading on Beam Instrumentation
H. Schmickler (Ed.) Beam Instrumentation, Proc. CERN Accelerator School, Tuusula 2018 in prep.
D. Brandt (Ed.), Beam Diagnostics for Accelerators, Proc. CERN Accelerator School, Dourdan,
CERN-2009-005, 2009;
Proceedings of several CERN Acc. Schools (introduction & advanced level, special topics).
V. Smaluk, Particle Beam Diagnostics for Accelerators: Instruments and Methods,
VDM Verlag Dr. Müller, Saarbrücken 2009.
P. Strehl, Beam Instrumentation and Diagnostics, Springer-Verlag, Berlin 2006.
M.G. Minty and F. Zimmermann, Measurement and Control of Charged Particle Beams,
Springer-Verlag, Berlin 2003.
S-I. Kurokawa, S.Y. Lee, E. Perevedentev, S. Turner (Eds.), Proceeding of the School on Beam
Measurement, Proceedings Montreux, World Scientific Singapore (1999).
P. Forck, Lecture Notes on Beam Instrumentation and Diagnostics, JUAS School, JUAS Indico web-site.
Contributions to conferences, in particular to International Beam Instrumentation Conference IBIC.

## Slide 65
Backup slides

## Slide 66
Example: Variation of vertical injection angle by magnetic steerer
Beam: C6+ at 6.7 MeV/u acc. to 600 MeV/u, up to 6109 ions per fill
with multi-turn injection, IPM integration 0.5 ms i.e.  100 turns
Emittance Enlargement by Injection Mis-steering
Vertical profile at injection:
Horizontal profile at injection:
Horizontal profile after acc.:
Vertical profile after acc.:
before
acc.
after acc.
Emittance conservation requires
precise injection matching
Wrong angle
of injected beam:
injection into outer
phase space  large
-amplitude i.e. large beam
might result in
‘hollow’ beam
filling of acceptance
i.e. loss of particles
Hadron beams: larger
emittance after acceleration
injection
extraction
IPM
synchrotron
vertical
steerer
injection:angle
mismatch
Schematic simulation:
Courtesy M. Syphers
misplace injection
filamentation
larger emittance
x
x’

## Slide 67
100 ns delayed emission
no OTR as expected (classical process)
emission by scint. screen due to lifetime
 correct profile image!
Coherent Optical Transition Radiation
Observation of coherent OTR for compressed bunches at LINAC based light sources
Reason: Coherent emission if bunch length  wavelength (tbunch=2 fs  lbunch =600 nm)
or bunch fluctuations  wavelength
Contrary of M. Yan et al., DIPAC’11 & S. Wesch, DIPAC’11
prompt emission for OTR and scint. screen
 coherent and in-coherent OTR
OTR screen
scint. screen
Parameter reach
for most LINAC-based FELs!
Beam parameter: FLASH, 700 MeV,
0.5 nC, with bunch compression
beam
photons
OTR screen
in-coherent emission
beam
photons
OTR screen
coherent emission