## Slide 1: November test beam results
Within the MAMMA collaboration
Arizona, Athens(U, NTU, Demokritos), Brookhaven,
CERN, Harvard, Istanbul, Naples, CEA Saclay, Seattle, USTC Hefei, South Carolina, St. Petersburg, Shandong, Stony Brook, Thessaloniki
Shuoxing Wu / CEA Saclay, USTC Hefei
D. Attié, T. Alexopoulos, M. Boyer, P. Colas, J. Derré, F. Diblen,
E. Ferrer Ribas, T. Geralis, A. Giganon, I. Giomataris, F. Jeanneau, K. Karakostas, S. Kirch, E. Ntomari, M. Titov ,W.Wang

## Slide 2: Outline:
Spark current-voltage monitoring.
Spark rate.
Efficiency and spatial resolution.

## Slide 3: Test beam set up:
Telescope:
3 X-Y detectors(10 x 10 cm2) manufactured at Saclay
Aim: Test different resistive films detectors manufactured by Rui De Oliveira at CERN and compare behaviour to non-resistive detectors
Electronics: GASSIPLEX
DAQ: realised by Demokritos
Gas: 95%Ar + 3% CF4 + 2% isobutane
120 Gev π+
Y
Resistive
Non-Resistive
X
Y
X
X
Beam
1 mm
0.25 mm
1 mm
Detectors in test
Y
Tested detectors:
Standard bulk detectors; Resistive coating detectors; Segmented mesh detector
SPS-H6

## Slide 4: Summary of tested detectors:
Type | Name | Properties:
Standard bulk: | SLHC3 | 1mm pitch
 | SLHC2 | 2mm pitch
Resistive coating: | R3&R4 | 2mm,2MΩ/ ,kapton+insulator
 | R5 | 2mm,250MΩ/ ,resistive paste
 | R6 | 1mm,400KΩ/ , resistive strip
 | R7 | 0.5mm, tens of KΩ/ , resistive pad
Segmented mesh | S1 | 1mm pitch,8 segmentations

## Slide 5
Different sparking behaviors of standard and resistive detector:
Standard SLHC2 (2mm) (@10KHz):
R6(1mm,400kΩ/ resistive strip) (@10KHz):
SLHC2: HV=400 V (Gain ~3000): current when sparking < 0.4 mA
voltage drop< 5%
R6: HV=390 V (Gain ~3000): current when sparking < 0.08 mA
voltage drop<0.5%

## Slide 6
Sparking behavior of resistive detector R3:
Standard SLHC2 (2mm) (@10KHz):
Resistive R3(2mm,2MΩ/ ) (@10KHz):
SLHC2: HV=400 V (Gain ~3000): current when sparking < 0.4 mA
voltage drop< 5%
R3: HV=410 V (Gain ~3000): current when sparking < 0.2 mA
voltage drop<2%

## Slide 7: Sparking behavior of S1:
Seven,eight sparking
six sparking
five sparking
four sparking
three sparking
two sparking
one sparking
S1:
Standard bulk: SLHC2

## Slide 8: Charge vs. channel number:
1mm
1mm
1mm
1mm
250mm
250mm
R3
R6
Pedestal shift in standard MM (telescope) due to sparks.
No pedestal shift in resistive detectors (R3&R6).

## Slide 9
Chart: Spark rate in all detectors:
| 390 | 400 | 410
R7 | 0.0005595 | 0.00059175 | 0.00063225
R5 | 0.00014 | 0.00016578 | 0.00014585 | 0.00018294 | 0.00043609
Standard | 1.05e-05 | 1.7344e-05 | 4.375e-05 | 7.2115e-05 | 0.00014224
S1 | 1.7571e-05 | 2.1136e-05 | 5.122e-05 | 5.3625e-05
R3 | 5.5777e-06 | 5.9615e-06 | 7.5051e-06 | 9.6852e-06 | 1.1731e-05 | 1.0336e-05 | 1.3466e-05 | 1.1894e-05
R6 | 7.9286e-06 | 7.8333e-06 | 6.4583e-06 | 6.4286e-06 | 7.8191e-06 | 1.0588e-05
Spark rate vs mesh voltage and beam intensity:
Chart: 
| 5.2 | 11.5 | 30 | 40
R3 spark rate | 2.0911e-06 | 3.4357e-06 | 7.3026e-06 | 8.7539e-06
R3(2mm, 2MΩ/ ):

## Slide 10
Detector | type | Spark rate | Spark current /mA | Voltage drop
SLHC2 | standard | 7*E-5 | 0.4 | 5%
R3 | 2 M Ω/ resistive kapton | 9.6*E-6 | 0.2 | 2%
R6 | 400K Ω/ resistive strip | 6.4*E-6 | 0.08 | 0.5%
R5 | 250 M Ω/ resistive paste | 1.6*E-4 | 0.1 | 1.5%
R7 | tens of K Ω/ resistive pad | 5.9*E-4 | 0.35 | 4.5%
Detector performance at same gas gain (~3000):

## Slide 11
R6 (1mm, 400KΩ/ ) efficiency :
>98%
drop<3%
Chart: Beam Intensity/KHz*cm-2
| 5.2 | 11 | 30 | 40
 | 0.9925 | 0.9911 | 0.982 | 0.968
Chart: Gain
| 1300 | 2100 | 3300 | 5100 | 10000 | 27000
 | 0.9742 | 0.984 | 0.9895 | 0.9903 | 0.9917 | 0.9924
@beam intensity of 11kHz/cm2:
@voltage of 400V:
Chart: Mesh Voltage/V
| 370 | 380 | 390 | 400 | 410 | 420
 | 0.9742 | 0.984 | 0.9895 | 0.9903 | 0.9917 | 0.9924
Chart: Gain
| 1300 | 2100 | 3300 | 5100 | 10000 | 27000
R6 | 0.9742 | 0.984 | 0.9895 | 0.9903 | 0.9917 | 0.9924
SLHC2 | 0.9759 | 0.9802 | 0.9744 | 0.966 | 0.9651

## Slide 12
Resolution:
beam
MM
MM under test
MM
MM
1mm
0.25 mm
1mm
•Residuals of MM cluster position and extrapolated track from MM telescope:
•Convolution of:
- Intrinsic MM resolution
- Track resolution (extrapolated)
- Multiple scattering
R6(1mm pitch, 400kΩ/ )
δ=134mm
δmm= 91±2.2mm
~98mm

## Slide 13
R6(1mm 400kΩ/ resistive strip) resolution vs mesh voltage and beam intensity:
Chart: Mesh voltage/V
| 370 | 380 | 390 | 400 | 410 | 420
 | 123.9355 | 91.3893 | 97.1597 | 110.9054 | 117.4947 | 136.4551
Mesh voltage:
Beam intensity:
Chart: Beam Intensity/KHz*cm-2
| 5.2 | 11 | 30 | 40
 | 105.5083 | 111.7048 | 119.9584 | 108.8956

## Slide 14
Cluster size (mean of distribution) :
Chart: Voltage/V
| 380 | 390 | 400 | 410 | 420 | 430 | 440
R3 | 5.968 | 6.996 | 7.96 | 8.966 | 9.138 | 9.656 | 9.728
R6 | 1.824 | 2.101 | 2.413 | 2.851 | 3.291 | 3.598
SLHC2 | 2.5 | 2.568 | 2.668 | 2.712
R3 : 2mm , 2 M Ω/ Resistive kapton +insulator
R6 : 1mm , 400 k Ω/ Resistive strip
SLHC2: 2mm standard bulk

## Slide 15
Resistive coating is a successful method to reduce the micromegas spark rate and limit the change in mesh voltage and current.
Good spatial resolution<100mm can be reached with a resistive strip coating detector of 1mm pitch.
High efficiency (>98%) can be achieved with resistive strip coating micromegas detector, and efficiency drops less than 3% when increasing the beam intensity from 5KHz to 40KHz . R6 can operate at high gain up to 30000 without loosing efficiency.
R&D and studies will continue inside the MAMMA collaboration
Conclusion and outlook:

## Slide 16
Back up slides

## Slide 17: Number of missing strips in cluster:
R6( 1mm, 400 KΩ/ ):
R3( 2mm, 2 mΩ/ ):
17

## Slide 18: Probability to have missing strips in cluster and number of missing strips:
18

## Slide 19: Sparks in R7(few tens of kΩ/ Resistive pads):
All the ‘sparks’ in R7:
Three ‘spark’ type:
‘Spark’ amplitude distribution:

## Slide 20: Sparks in R5(250 MΩ/ Resistive kapton):
All the ‘sparks’ in R5:
‘Spark’ amplitude distribution:
Four ‘spark’ type:

## Slide 21: Sparks in SLHC2(standard bulk):
All the ‘sparks’ in SLHC2:
One ‘spark’ type:
‘Sparks’ amplitude distribution:

## Slide 22: Sparks in R3(2 MΩ/ resistive kapton):
All ‘sparks’ in R3:
‘sparks’ amplitude distribution:
One ‘spark’ type:

## Slide 23: Sparks in R6(400 kΩ/ resistive strip):
All ‘sparks’ in R6:
‘sparks’ amplitude distribution:
One ‘spark’ type:

## Slide 24: Spark amplitude vs voltage and beam intensity:
R5 vs voltage:
R7 vs voltage:
SLHC2 vs voltage:
R3 vs intensity:
R5 vs voltage:
R3 vs voltage: