## Slide 1: VACUUM ACCEPTANCE TESTS FOR PARTICLE ACCELERATOR COMPONENTS AT CERN

## Slide 2: OUTLINE
What is a vacuum acceptance test?
Why is it needed?
How is it performed? which physical quantities and information are we interested in?
How to use the results?
Few examples;
What happens in case of non conformity?
Some statistics.

## Slide 3: What is a vacuum acceptance test of a component of a particle accelerator?
In order to check if the component outgassing and eventual contamination can alter the vacuum conditions of the accelerator, specific tests are performed to evaluate its gas load and its residual gases footprint.
Kicker
Scrapers
Beam Position Monitors

## Slide 4: Why is it needed?
In order to guarantee a certain beam lifetime, or certain machine operation conditions, a maximum acceptable pressure is defined for each machine. Therefore, vacuum components’ outgassing and residual gas analysis have to be compliant with these acceptance levels.
One of the main mandates of TE-VSC is to provide beam operation with the required vacuum in our accelerators.
Questions box:
Is the existing pumping layout going to cope with the additional gas load?
Is the component going to degrade the beam vacuum?(leaks, hydrocarbons..)
For new layouts: do I have to modify the vacuum layout accordingly? Or foresee a specific procedure for critical components?
Is the component going to limit the machine operation?
REMINDER!! Outgassing rates spans over orders of magnitude, while the effective pumping speed is conductance limited.

## Slide 5: Beam and residual gas interactions
CnHn+2
e -
e -
Types of interactions:
Nuclear interactions  high energy particles ( )  particle loss, radiation to equipment and background to detectors;
Electromagnetic interactions  elastic scattering  emittance growth;
   inelastic scattering  energy spread;
Ionisation of residual gases  Ion induced desorption;
 e- cloud;
Synchrotron radiation  photon induced desorption;
 e- cloud;
H20
H2
CO
CO2
CH4
N2
Beam instabilites and particles loss

## Slide 6: Beam and residual gas interactions
Particles loss leads to a decrease of bunch population:
where
The probability of a certain interaction depends on the mass of the molecule (Z) and energy of the particles.
LHC vacuum system, O.Grobner
Ex. LHC

## Slide 7: Particle accelerator vacuum requirements
Beam lifetime is crucial, but is only a part of the full picture:
EQUIPMENT REQUIREMENTS:
  UHV insulating conditions for high voltage component (septa, kickers);
  MACHINE OPERATION:
  Beam downtime: operation should be resumed ASAP!

## Slide 8: Operating pressure levels for CERN accelerator complex
Pressure spans over orders of magnitudes!

## Slide 9: Operating pressure levels for CERN accelerator complex
LHC arcs <10-8 mbar
<10-12 mbar
LEIR
LHC LSS<10-10 mbar
UNBAKED SYSTEMS (SEPTA..):
TMP, ION PUMPS;
SUBLIMATORS;
BAKED SYSTEMS:
ION PUMPS;
Non Evaporable Getter (TiZrV) coating
CRYO SYSTEMS:
CRYOPUMPING;
ELENA ring<4*10-12 mbar
Linac 4 <2*10-6 mbar*
PSB <5*10-8 mbar*
PS<2*10-8 mbar*
SPS LSS <10-7 mbar*
* 24h pumpdown

## Slide 10: Pressure requirements for LHC and ELENA
Area | Pressure requirements | Effective pumping speed (indicative)
Arcs | ≤10-8 mbar | ≥100 l.s-1
Experiments | ≤10-10 mbar | 
LHC
ELENA ACCEPTANCE THRESHOLDS
Ensures the limitation of momentum and emittance blow up induced by the interaction of a low energy (100 keV) antiproton beam with residual gases.
ELENA
Area | Pressure requirements | Effective pumping speed (indicative)
Ring | ≤4 x10-12 mbar | Depend upon position (NEG sticking probability)
Transfer lines | ≤10-10 mbar | 
LHC ACCEPTANCE THRESHOLDS
Ensures 100 h of circulating beam lifetime and minimize the background to the experiments.
In both examples, the pressure requirements are driven by beam lifetime requirements.
Notes: 100 KeV

## Slide 11: Pressure requirements for the LHC Injectors
Accelerator | Area | Operational pressure requirement (24h pumping) [mbar] | Average effective pumping speed (indicative) [l s-1] | Outgassing rate limit at 24 h [mbar.l.s-1]
PS complex | LINACS, ISOLDE AND TRANSFER LINES CLOSE TO PSB AND PS | ≤2.10-6 | 100 | 5∙10-5 (*)
 | PSB, HIE-ISOLDE, REX AND TRANSFER LINES CLOSE TO THE PS RING | ≤5.10-8 | 100 | 5.10-6 (*)
 | PS ring | ≤2.10-8 | 70 | 1.5.10-6 (*)
SPS | Arcs | ≤10-6 | 10 | 10-5
 | LSS (kickers, septa, RF cavities) | ≤10-7 | 100 | 10-5 (*)
 | TI2&TI8: From SPS to TED | ≤10-5 | 1-2 | 2.10-5
 | TI2&TI8: From TED to LHC | ≤5.10-7 | 5 | 2.5.10-6
  Ensure beam operation after 24h pumpdown;
  Ensure proper functioning of High Voltage and RF devices;
  Ensure ion operation at low energy (e.g. LEIR, PS);
SPS ARCs PUMPDOWN
We must comply with beam requirements minimising the beam downtime

## Slide 12: What about contaminations?
Criteria for vacuum acceptance tests, P. Chiggiato, J.A.F. Somoza, G. Bregliozzi, EDMS 1752123

## Slide 13: Contamination acceptance levels for unbaked components
The unbaked component is considered accepted if the ratio between the water peak and the peaks up to mass 44 is higher then 100 and if the same ratio is higher then 1000 for peaks above mass 44.
*simulated spectra
cPS, SPS, transfer lines
H2O
H2
N2/CO
N+
CO2
C+

## Slide 14: Contamination acceptance levels for baked components
For baked components the H2 peak becomes the reference with a more stringent acceptance criteria (i.e. 104 ratio for masses > 44).
*simulated spectra
e.g. LHC, AD, ELENA
H2O
H2
N2/CO
N+
CO2
C+

## Slide 15: What about leak tightness?
External Leak
Internal leak
Helium leak rate acceptance threshold is 1*10-10 mbar *l/s;
A leak is an indication of possible material defects (precipitates, cracks..), flaws in the fabrication process or in the mechanical assembly.
No deviation from this value are accepted.
Internal leak rate acceptance threshold is < 5*10-9 mbar*l/s both for LHC (NEG saturation), Injectors (<20% of the total pressure).
It is an indication of flaws in the conceptual design (trapped volumes).
Leak can be evaluated by accumulation;

## Slide 16: Finally, a vacuum component is accepted for installation if…
It is leak tight;
Its outgassing is compliant;
Its residual gases spectrum is compliant;
If no virtual leaks or acceptable virtual leaks are detected.

## Slide 17: How is a vacuum acceptance test performed?
THROUGHPUT METHOD:
1
2
3
4
5
1
2
3
4
5
The pressure rise method (accumulation) and the coupled method (througput + accumulation) are often used as well.

## Slide 18: How is an acceptance test performed?
UNBAKED COMPONENTS
BAKED COMPONENTS
Pumpdown;
Leak detection;
Outgassing estimation at 24h;
RGA scan after 4h of filament conditioning.
Pumpdown;
Leak detection;
Bake – out (dedicated recipes for specific components);
Outgassing evaluation at RT after 48h from the end of the bake-out;
RGA scan;
Evaluation of possible internal leaks by accumulation.

## Slide 19: Which physical quantities and information are we interested in? Pumpdown curve
Outgassing rate:
  Unbaked components:
  Baked components:
For unbaked components the slope of the pumpdown gives a crucial information concerning the pumpdown time: are we desorption limited or diffusion limited (metallic surface Vs polymers)?

## Slide 20: Which physical quantities and information are we interested in? Residual Gas Analisys
In this example:
  Leak?
  Hydrocarbon contamination?
The residual gas scan up to 100 a.m.u./e- gives an insight of the residual gases composition: once compared with the system scan, it is possible to attribute to the component its outgassing footprint.

## Slide 21: Which physical quantities and information are we interested in? Internal leak
Leak / internal leak evaluation by accumulation:
With a calibrated RGA, an estimation of the leak rate can be obtained by analysing mass 40, 20 pressure rises.

## Slide 22: Which physical quantities and information are we interested in? NEG coated components
Getter surfaces are characterized by the sticking probability:
The sticking probability cannot be directly measured, it is obtained by combining experimental pressure measurements and MolFlow simulations.
Two methods are used:
  Fisher-Momsen dome, where the probability that a molecule entering the deposited component is pumped at the surface (capture probability) is evaluated.
  Measurement in transmission, where the probability that a molecule is not pumped and it is transmitted through a component is evaluated.

## Slide 23: Which physical quantities and information are we interested in? NEG coated components
SAMPLE
Transmission Probability of NEG coated components
The measured pressure difference must be correlated through a Molflow simulation to the actual sticking probability of the coated surface.

## Slide 24: Which physical quantities and information are we interested in? NEG coated components
Courtesy of J. Hellstrom

## Slide 25: How do we use these results?
Compare them with the acceptance levels:
  Depending on the machine;
  Depending on the final location of the component (local effective pumping speed);
  The installation of the component is finally accepted, tolerated or rejected.
For new machine or new layouts:
  The measured outgassing data can be used as input for vacuum simulations;
  The vacuum design can be adapted accordingly;

## Slide 26: How do we use these results?
Compare them with the acceptance levels:
  Depending on the machine;
  Depending on the final location of the component (local effective pumping speed);
  The installation of the component is finally accepted, tolerated or rejected.
For new machine or new layouts:
  The measured outgassing data can be used as input for vacuum simulations;
  The vacuum design can be adapted accordingly;

## Slide 27: Few examples: PS & SPS FWS
Non conformities detected coming from the fork.
Courtesy of A.Michet and N. Thaus

## Slide 28: Few examples: PS & SPS FWS
Working toward the conformity:
Tanks stripped down and cleaned again;
Mild bake-outs;
N2 hot injections;

## Slide 29: Few examples: SPS MKDV
The MKDVs are kickers magnet meant to vertically deflect the beam toward the dump. They are composed of a stacked lamination of ferrites (8C11, NiZn alloy) and metallic ground plates (AlMg3F26) over 3m length.
300 l/s Ion pumps are directly connected to the tank to cope with the high outgassing rate.

## Slide 30: Few examples: SPS MKDV
Chart: MKDV2 tank#2 TMP pumpdown curve
Acceptance test is performed with ion pumps as well:
Outgassing rate at 24h with turbomolecular pumping exceeds the SPS acceptance levels (≈ 9*10-5 mbar l/s);
The additional ion pumps, reduce the pressure to an acceptable level for operation (≈ 7*10-8 mbar at 48h of ion pumping).
MKDV are and example of tolerated components.
Courtesy of A.Harrison and K. Owens

## Slide 31: Few examples: TCDI, TI2-TI8 collimators
12 collimators will be installed in TI2 and TI8 to protect LHC beam injection from eventual mis-stirring of the beam during extraction from the SPS.
Each collimator is based on 2 graphite (3D-CC + isostatic) jaws, positioned at 4.5 - 5σ from the beam.
All collimators will be baked in the machine to reduce their impact to the vacuum lines:
  TI2 and TI8 transfert lines are caracterized by a low conductance and lumped, spaced ion pumping;
  Sectors 1204 and 1805 are interfacing LHC NEG coated sectors.

## Slide 32: Few examples: TCDI, TI2-TI8 collimators
Courtesy of G. Cattenoz
For complex components, subassemblies are tested in order to anticipate possible non conformities before the final assembly.
3D CC blocks
Jaws
TCDI – Final assembly

## Slide 33: Few examples: TCDI, TI2-TI8 collimators
Courtesy of G. Cattenoz
Total Outgassing rate measured after a bake out at 250 C for 48h. RGA scans measured after 48h at room temperature:
    Outgassing rates ≈ 6*10-8 mbar l/s;  conform!
    RGA scans all within acceptance limit.
Graphite conformity reached thanks to a vacuum firing at 950 ◦C for 6 h.

## Slide 34: Non conformities: why?
Material choice (polymers, synthered materials, high vapour pressure materials);
Missing Thermal treatments (vacuum firing, air baking, prebaking for polymers);
Poor mechanical vacuum design (trapped volumes, wrong welding procedure, brazing alloys);
Non conformity in the design (lubricants under vacuum, glues..)
Non conformity in the handling procedures;
Non conformity in the cleaning procedures;
OUTGASSING RATE
CONTAMINATIONS

## Slide 35: What happens in case of non conformity?

## Slide 36: What happens in case of non conformity?
  Best way to deal with a non conformity is to avoid it!
  Early involvement of vacuum experts in the design phase of the components is crucial to avoid common design mistakes that can lead to non conformities at the final stage of the production cycle of the component.

## Slide 37: A bit of statistics..
Courtesy of I. Wevers
Chart: 
| CPS | SPS | AD | ELENA | LHC | ISOLDE | GENERAL
 | 0.28333 | 0.125 | 0.016667 | 0.05 | 0.475 | 0.025 | 0.025
More than 120 components tested last year!
Chart: Solution by machine
| Non conform without solution | CPS | SPS | AD | ELENA | LHC | ISOLDE
 | 0.53704 | 0.16667 | 0.14815 | 0 | 0.018519 | 0.12963 | 0

## Slide 38: A bit of statistics..
Chart: Machine
| CPS | SPS | AD | ELENA | LHC
 | 0.23326 | 18.0341 | 0.042 | 0.4437 | 0.4334
Outgassing rate acceptance limit
Courtesy of I. Wevers

## Slide 39: Conclusions
A brief (non –exhaustive!) overview of vacuum acceptance tests;
Crucial role of BVO in testing and assessing the compliance of vacuum components to be installed in the accelerator chain;
Thanks to all BVO colleagues involved into these activities!