## Slide 1: Germanium Charge-Coupled Devices
Christopher Leitz (on behalf of the germanium imagers team)
16 December 2022
christopher.leitz@ll.mit.edu
DISTRIBUTION STATEMENT A. Approved for public release. Distribution is unlimited.
This material is based upon work supported by the National Aeronautics and Space Administration under Air Force Contract No. FA8702-15-D-0001. Any opinions, findings, conclusions or recommendations expressed in this material are those of the author(s) and do not necessarily reflect the views of the National Aeronautics and Space Administration .
© 2022 Massachusetts Institute of Technology.
Delivered to the U.S. Government with Unlimited Rights, as defined in DFARS Part 252.227-7013 or 7014 (Feb 2014). Notwithstanding any copyright notice, U.S. Government rights in this work are defined by DFARS 252.227-7013 or DFARS 252.227-7014 as detailed above. Use of this work other than as specifically authorized by the U.S. Government may violate any copyrights that exist in this work.
Acknowledgements:
We gratefully acknowledge support for this work from NASA's Astrophysics Division to MIT Lincoln Laboratory through Astrophysics Research and Analysis grant 80HQTR19T0048.

## Slide 2: Unique Focal Planes in Support of National Security and Science
National Ignition Facility (LLNL)
Nuclear Stockpile Stewardship (LANL)
AMOS (AF)
Space Surveillance
Astronomy & Near Earth Object Detection
Directed Energy
AFRL SOR
Space-Based Astronomy

## Slide 3: Germanium Detectors Opportunity
MITLL Microelectronics Laboratory
X-Ray Sensitivity Comparison
Vision: extend the advantages of CCDs (format, noise…) into new material
Germanium wafers processed in same tools used to build silicon detectors for flight missions
Elemental high-Z detector material with broadband sensitivity

## Slide 4: Germanium CCD Development Timeline
2013
2017
2015
2020+
Test Devices
CTI = Charge-Transfer Inefficiency
C.W. Leitz, SPIE Proc. 11454 (2020) C.W. Leitz, SPIE Proc. 11118 (2019); C.W. Leitz, SPIE Proc. 10709 (2018); C.W. Leitz, J. Inst. 12(5) CO5014 (2017).
2nd-Gen Imagers
Area of current focus
1st-Gen Imagers
Notes: Previously cleared under RR#126466

## Slide 5: Silicon and Germanium CCD Comparison
400 nm
Pixel
Pixel
Serial Register
Serial Register
Output Amplifier
Output Amplifier
Gates (TiN)
Gates (poly-Si)
Routing Metal
Routing Metal
20 µm
24 µm
Germanium CCDs draw upon long heritage of silicon CCD designs and processes
Silicon CCD
Germanium CCD
Gate dielectric (GeO2/Al2O3)
Gate dielectric (SiO2/Si3N4)

## Slide 6: Outline
Background and motivation
Status of technology development
  2nd-generation imager design
  First batch of wafers (“pilot wafers”) & yield enhancement
  Latest batch of wafers
Plans for FY23 and beyond

## Slide 7: Second-Generation Imagers
512 × 512 frame-transfer CCD, 24 µm pixels
1 Mpixel frame-transfer CCD, 8.1 µm pixels
Wide range of new devices, reflecting lessons from first-generation imagers
  512 × 512 frame-transfer CCD with 24 µm pixels, two MOSFET outputs
  2562, 1282, 642 full-frame CCDs with 8.1 µm pixels, one MOSFET & one JFET output
  Key yield diagnostic testable at metal-1

## Slide 8: Elimination of Parasitic MOSFET in Output Region
Channel Stop
Contact / via
Metal-1
Metal-2
Legend
5 µm
Output
gate
Reset gate
5 µm
1st-Generation Design
2nd-Generation Design
Rerouted metal-1 features to eliminate parasitic transistors to further improve device performance

## Slide 9: New Al2O3 Deposition System
Control
(Metal-1 on SiO2)
Metal-1 on Al2O3
(existing system)
Metal-1 on Al2O3
(new system)
Defect Maps
New, more capable Al2O3 deposition system used on four pilot wafers
  Particles from previous Al2O3 deposition system caused patterning defects which short phases of device
  New system with vastly lower particle counts installed and qualified in late 2020
Region of die measured

## Slide 10: Shorts-Opens Test Summary on Pilot Wafers
Shorts/opens testing showed poor yield on pilot wafers despite use of new Al2O3 deposition system
Analysis of pilot wafers began with shorts/opens testing
  A single short among the 15 different tests results in device failure
No significant yield improvement over previous device runs
  Multiple shorts on all 2562 devices tested (at left)
  Average short spacing a few mm
New investigation into yield launched
  Discovered that new Al2O3 deposition tool, with much lower particle counts, lifted the veil on a previously hidden failure mechanism
Sample Yield Map for 256 × 256 Full-Frame CCD
All greens needed in a column to ensure that device functions

## Slide 11: “Starry Night” Defects
Root cause traced to nodules present on wafer surface after gate dielectric growth, composed only of Ge & O
Nodules micro-mask metal gate etch, leaving ring of metal around the center nodule & leading to shorts
Post-Oxidation
Post-Metal Dep & Etch
Diagnosis
Ge + O
Metal
Mitigation
25 µm
25 µm
Previous Process
New Process
New oxidation condition sharply reduces density of starry night defects
Electrical tests on monitor wafers positively correlated starry night defect density to shorts-opens yield

## Slide 12: Shorts-Opens Test Summary on Second Batch of Wafers
Sample Yield Map for 256 × 256 Full-Frame CCD
All greens needed in a column to ensure that device functions
Average short spacing ~ 1 m
  Approximately 200x improvement in average short spacing
Enables us to routinely yield larger arrays
  Example here: 7/14 2562 arrays are clean
Known issue at metal patterning on these wafers
  Wafer flatness issue mentioned in previous slide
  Yield would have been even higher if not for that
Significant yield improvement in this second batch of wafers & no shortage of ideas for additional improvement!

## Slide 13: First (Preliminary) Test Results on Newest Devices
256 × 256-pixel array being clocked through MOSFET output at -60°C
Dark
Flood-Illuminated
Encouraging early result; still a lot of work to do to optimize clocking biases

## Slide 14: Conclusions and Future Work
We have realized germanium CCDs that show promise for scientific detectors
Our current efforts are focused on improving performance and yield
  New output amplifier designs to lower read noise and dark current
  Improved processes to reduce density of gate-to-gate shorts which limit device format
Future work:
  Perform device screening on newly completed wafers
  Transition device testing to cryogenic environment
    Collaboration with MIT Kavli Institute for Astrophysics & Space Research
  Begin back-illuminated detector development (pending NASA award)
  Incorporate lessons learned on newly completed wafers to partially-processed wafers
    Seven wafers staged ~6 weeks from completion