## Slide 1: FSI qualification for Internal monitoring Alignment review 2019
Vivien RUDE
2019-08-26

## Slide 2
Outline
Internal monitoring
Crab-Cavities project
Dipole Test project
(Test for HL-LHC quadrupole)
Crab-Cavities in SPS accelerator
Dipole-test in SM18

## Slide 3
Internal monitoring
Objective : determine the position of inner component inside a vacuum vessel (Cryostat, Cryomodule)
Environment :
Temperature : 1.9 K (Cryogenics conditions)  - 271°C
Vacuum : 10-6 mBar
Radiation : 1 MGy / year
Accuracy :
0.1 mm w.r.t. vacuum vessel
Crab-Cavities
Dipole

## Slide 4
FSI : Frequency Scanning Interferometry
Accuracy :
0.5 µm per meter
Absolute distance measuring interferometric technique

## Slide 5
Crab-cavities project
Crab-cavities
Cryomodule
Feedthrough

## Slide 6
Configuration
8 FSI distances by Cavity for 7 unknowns
beam
Cavity
Cryomodule
Thermal shielding
Vacuum : 10-6 mbar
Temperature : 4 K
Vacuum : Ambient
Temperature : Ambient

## Slide 7
Validation at Warm condition
FSI
(absolute distances)
BCAM
(Angle measurements)
Laser Tracker
(Angles and distances measurements)

## Slide 8
Intercomparison between FSI / BCAM / Laser Tracker (Before cooling down)
FSI, BCAM systems precision better than 50µm (1σ), crosschecked with AT401 laser tracker measurements

## Slide 9
Cooling Down in SPS : From 293 K to 4K
Accuracy : < 50 µm (1)

## Slide 10
Intercomparison at Cold condition (FSI – BCAM)

## Slide 11
Long term stability
Since April 2018, the monitoring of the Crab-cavities worked correctly.
Parameter | Precision (1σ)
Tx (radial) | +/- 25 μm
Ty (longitudinal) | +/- 45 μm
Tz (vertical) | +/- 10 μm
Rx (pitch) | +/- 30 μrad
Ry (roll) | +/- 150 μrad
Rz (yaw) | +/- 70 μrad
Scale | +/- 60 ppm
Redundancy :
8 observation for 7 unknowns  1 degree of freedom
If some optical sight views are not possible  Hypothesis (for example : Fixed Ry or Tx)

## Slide 12
Dipole test

## Slide 13
Objective
Vacuum : 10-6 mbar
Temperature : 1.9 K
In order to validate the internal monitoring solution for the Low-Beta quadrupoles, a dipole has been used as test setup.
Section 1
Section 2
Section 3

## Slide 14
Configuration
Section 1
Section 2
Section 3
Long object (15 m)  Cryostat and cold mass may not be considered as rigid body  Measurement on 3 sections
12 distances measurements (FSI system)
3 sections
Deformation up to 0.25 mm has been observed
on the cryostat and the cold mass
during transport and during cool down

## Slide 15
Summary of first tests (2017 - 2018)
Cryo-condensation on the laser reflectors is a main showstopper for optical distance measurements in “dusty” cryostat
Cryo-condensation

## Slide 16
Insulated support

## Slide 17
Test summary (end 2018)
Cold mass
Cryostat
Thermal
shielding
Vacuum : 10-6 mbar
Vacuum : Ambient
Temperature : Ambient
Acquisition system : FSI
Commercial : Etalon
Multi-Targets : CERN
Collimator
Parallel lens
Divergent collimator
FSI Head
Feedthrough (adjustable)
Feedthrough (fixed)
Window
Targets
Newport reflector (3 mirrors)
Glass sphere without coating
Glass sphere with coating
Temperature : 4 K
Temperature : 180 K
Target Support
Insulated target
Nothing

## Slide 18
Chart: Date
3 Cooling-down
Schedule

## Slide 19
Section 1
COLD (4K)
Legend
WARM (293 K)
ETALON (FSI)
Window (crossing)
Newport prism
Laser tacker
Newport prism / Glass sphere
MT-FSI
Window (crossing)
Glass sphere
MT-FSI
Window (crossing)
Newport prism
ETALON (FSI)
Window (crossing)
Glass sphere with coating
MT-FSI
Window (crossing)
Glass sphere with coating
MT-FSI
Feedthrough
Glass sphere with coating
+/- 0.1 mm
Section 1

## Slide 20
Section 2
WARM (293 K)
COLD (4K)
Legend
+/- 0.1 mm
+/- 0.1 mm
+/- 0.1 mm
Section 2
ETALON (FSI)
Window (crossing)
Newport prism
Laser tacker
Newport prism / Glass sphere
MT-FSI
Window (crossing)
Glass sphere
MT-FSI
Window (crossing)
Newport prism
ETALON (FSI)
Window (crossing)
Glass sphere with coating
MT-FSI
Window (crossing)
Glass sphere with coating
MT-FSI
Feedthrough
Glass sphere with coating

## Slide 21
Section 3
COLD (4K)
Legend
WARM (293 K)
+/- 0.1 mm
+/- 0.1 mm
Section 3
ETALON (FSI)
Window (crossing)
Newport prism
Laser tacker
Newport prism / Glass sphere
MT-FSI
Window (crossing)
Glass sphere
MT-FSI
Window (crossing)
Newport prism
ETALON (FSI)
Window (crossing)
Glass sphere with coating
MT-FSI
Window (crossing)
Glass sphere with coating
MT-FSI
Feedthrough
Glass sphere with coating

## Slide 22
Conclusion
Acquisition system : FSI
Commercial : Etalon
Multi-Targets : CERN
Collimator
Parallel lens
Divergent collimator
FSI Head
Feedthrough (adjustable)
Feedthrough (fixed)
Window
Targets
Newport reflector (3 mirrors)
Glass sphere without coating
Glass sphere with coating
Target Support
Insulated target
Nothing
Cold mass
Cryostat
Multi-Targets
Acquisition system
Feedthrough
Divergent
collimator
Glass sphere
WITH coating
Insulated target

## Slide 23
Parameter | Precision (1σ)
Tx (radial) | +/- 25 μm
Ty (longitudinal) | +/- 45 μm
Tz (vertical) | +/- 10 μm
Rx (pitch) | +/- 30 μrad
Ry (roll) | +/- 150 μrad
Rz (yaw) | +/- 70 μrad
Scale | +/- 60 ppm
Parameter | Precision (1σ)
Tx (radial) | +/- 50 μm
Ty (longitudinal) | +/- 55 μm
Tz (vertical) | +/- 15 μm
Rx (pitch) | +/- 5 μrad
Ry (roll) | +/- 1500 μrad
Rz (yaw) | +/- 5 μrad
Scale | +/- 10 ppm
Internal monitoring
X (mm)
Y (mm)
Z (mm)
X (mm)
Y (mm)
Z (mm)
Crab-cavity
Triplet

## Slide 24
Thank you for your attention
On behalf of :
Mateusz Sosin
Hélène Mainaud Durand
Thibault Dijoud
Mathieu Duquenne
Anna Zemanek
Kacper Widuck
Jan Gabka

## Slide 25: SPARE

## Slide 26
Section 1
COLD (4K)
Legend
WARM (293 K)
ETALON (commercial)
Window (crossing)
Newport prism
Laser tacker
Newport prism / Glass sphere
Multi-targets
Window (crossing)
Glass sphere
Multi-targets
Window (crossing)
Newport prism
ETALON (commercial)
Window (crossing)
Glass sphere with coating
Multi-targets
Window (crossing)
Glass sphere with coating
Multi-targets
Feedthrough
Glass sphere with coating
+/- 0.1 mm
Section 1

## Slide 27
Section 2
WARM (293 K)
COLD (4K)
Legend
ETALON (commercial)
Window (crossing)
Newport prism
Laser tacker
Newport prism / Glass sphere
Multi-targets
Window (crossing)
Glass sphere
Multi-targets
Window (crossing)
Newport prism
ETALON (commercial)
Window (crossing)
Glass sphere with coating
Multi-targets
Window (crossing)
Glass sphere with coating
Multi-targets
Feedthrough
Glass sphere with coating
+/- 0.1 mm
+/- 0.1 mm
+/- 0.1 mm
Section 2

## Slide 28
Section 3
COLD (4K)
Legend
WARM (293 K)
ETALON (commercial)
Window (crossing)
Newport prism
Laser tacker
Newport prism / Glass sphere
Multi-targets
Window (crossing)
Glass sphere
Multi-targets
Window (crossing)
Newport prism
ETALON (commercial)
Window (crossing)
Glass sphere with coating
Multi-targets
Window (crossing)
Glass sphere with coating
Multi-targets
Feedthrough
Glass sphere with coating
+/- 0.1 mm
+/- 0.1 mm
Section 3

## Slide 29
Thermal contraction : From 293 K to 4 K
CONTRACTION | Cav 1 |  | Cav 2 | 
 | IN (mm) | OUT (mm) | IN (mm) | OUT (mm)
Simulation | 1.097 | 0.678 | 1.097 | 0.678
SM18 | 1.321 | 0.843 | 1.295 | 0.832
SPS | 1.310 | 0.834 | 1.320 | 0.835

## Slide 30
Cryostat
X
Z
COLD mass
Calculation

## Slide 31: Phase 4 : Results (Warm condition)
X (mm)
Y (mm)
Z (mm)
Section 1
A / B/ C / D
Section 2
I / J / K / L
Section 3
E / F / G / H

## Slide 32: Phase 4 : Results (Warm condition)
WARM 1
WARM 2
WARM 3
WARM 1
WARM 2
WARM 3
WARM 1
WARM 2
WARM 3

## Slide 33: Phase 4 : Results (Cold condition)
X (mm)
Y (mm)
Z (mm)
Section 1
A / B/ C / D
Section 2
I / J / K / L
Section 3
E / F / G / H

## Slide 34: Phase 4 : Results (Cold condition)
+/- 0.1 mm
+/- 0.1 mm
+/- 0.1 mm