## Slide 1: LHC synchrotron light imaging system “BSRT” Run I-II and plans for Run III
LHC Beam Size Measurement Review
G. Trad on the behalf of BE-BI-PM
01/10/2019

## Slide 2
CONTENT
System Introduction
Operational Experience in Run I/II
  Milestones and Challenges
  Optical System
  Cross-calibration
  Reliability
  Achieved Performance
Plans for Run III
Summary
2
G. Trad – LHC Beam Size Measurement Review

## Slide 3
Introduction
3
BSRT MONITOR
G. Trad – LHC Beam Size Measurement Review
Beam Emitted SR
SR Extraction
Optical System
Detector
Post-Process
Synchrotron radiation imaging system
Real time control system
System protection and cope with configuration changes
Accurate source description.
(Especially when energy dependent)
In-vacuum system dictated by machine aperture.
Critical beam interaction;
SR exposure.
SR focus for image formation.
Working point, magnification, & resolution:
(aberrations and diffraction)
Profiles from images projections.
Fits .
Users displays
2D distribution of photons.
Continuous, high rep rate acquisition
Sensitive

## Slide 4
4
BSRT MONITOR
G. Trad – LHC Beam Size Measurement Review
Beam Emitted SR
SR Extraction
Optical System
Detector
Post-Process
Synchrotron radiation imaging system

## Slide 5
5
SR source
G. Trad – LHC Beam Size Measurement Review
an Undulator (5T superconductive, 2 periods of 28cm)
  Peak wavelength ~600 nm at 450 GeV
  Out of visible range ~0.8-1TeV
  Peak wavelength ~2nm at 7TeV
a D3 type dipole (~9.5m, 6Km bending radius)
  Edge radiation enters visible ~1 TeV
  Core radiation dominates visible >2-2.5 TeV
  Broadband radiation and DOF of few meters (~3m)
Visible SR in the LHC is emitted from:
The LHC source is complex, with the energy change:
  Light intensity increase by 3 orders of magnitude
  Wavelength working point shifts from visible to NUV (intermediate energies = mix)
  Longitudinal position of the source shifts by ~3 m
  The source at high energy is longitudinally extended (few meters)

## Slide 6
6
BSRT MONITOR
G. Trad – LHC Beam Size Measurement Review
Beam Emitted SR
SR Extraction
Optical System
Detector
Post-Process
Synchrotron radiation imaging system

## Slide 7
7
G. Trad – LHC Beam Size Measurement Review
Light extraction:
Dictated by the minimum aperture to respect (distance beam mirror edge at injection).
Requested clearance (22mm: ~20 sigma) obtained at ~30 m after emission.
Problems in Run I:
Mirror heating due to beam coupling impedance resulted in mirror coating damage and mechanical structure failure.
SR extraction

## Slide 8
8
G. Trad – LHC Beam Size Measurement Review
OLD Extraction Mirror
NEW Extraction Mirror
Reduced beam coupling impedance by design
No heating issues observed in Run II
Extraction mirror support redesigned for Run II:
SR extraction

## Slide 9
9
SR extraction
G. Trad – LHC Beam Size Measurement Review
Experience in Run II
Observed a fade black spot ~12 mm away from mirror edge
Loss of reflectivity in the damaged area of ~ 15 % (white light)
No deposit was found on the viewports surface.
Damage is localized on a really sharp region (only were SR hits)
It appears as an etching of the material itself.
Damage seems a replica of the light distribution from the D3 for NUV
Damage coincide with undulator SR at 6.5 TeV in the soft X-ray region.

## Slide 10
10
BSRT MONITOR
G. Trad – LHC Beam Size Measurement Review
Beam Emitted SR
SR Extraction
Optical System
Detector
Post-Process
Synchrotron radiation imaging system

## Slide 11
11
Optical System
G. Trad – LHC Beam Size Measurement Review
Run I
Run II
Optical table mounted on rails to:
ease manipulation of the optics (minimizing risk of misalignment during interventions)
allow installation of complex systems (interferometry & coronagraph) in parallel to the imaging system.

## Slide 12
12
Optical System
G. Trad – LHC Beam Size Measurement Review
Focal length:
L1 ~5 m
L2 ~0.3 m
Total MAG
~0.3
Due to the peculiar characteristics of the SR, the parameters of the optical system (magnification and resolution of the Keplerian telescope) can only be determined using the real beam as source.
This implies that the only way to calibrate the BSRT is by cross calibrating it with another system providing the absolute beam emittance, in our case the WS.

## Slide 13
13
Optical System
G. Trad – LHC Beam Size Measurement Review
Experience in Run II:
Simplified system with refractive optics (2 achromatic lenses) with focusing obtained by moving L2.
Shifted operation at high energy to NUV to gain resolution (reduced diffraction, expected gain 20%, effective 10%). (Mention: External help for β increase in IR 4 from ABP.)
Issues:
Translation stages moving L2 from focus position at injection (source undulator) to Top energy (source D3) losing steps
-> losing focusing with time.
    Solved by offsetting the 2 lenses (Vis/UV) to avoid TS movement.
Optical filters (250 nm +/-10nm) degradation to the SR deposited power interaction
-> Resolution degradation with time.
    Solved by moving filters to a location where the SR spot is defocused.
Experience in Run I:
Optical system based on reflective optics (focusing mirrors):
Very sensitive to alignment
Degraded resolution wrt nominal due to SR cuts in the optical path (8 mirrors trombone delay line to compensate source shift).

## Slide 14
14
BSRT MONITOR
G. Trad – LHC Beam Size Measurement Review
Beam Emitted SR
SR Extraction
Optical System
Detector
Post-Process
Synchrotron radiation imaging system

## Slide 15
15
Detector
G. Trad – LHC Beam Size Measurement Review
Experience in Run I:
Analog cameras (20ms exposure) fiber coupled to image intensifiers (Gen I) + In house frame grabber (25 Hz).
Slow, limited dynamic range, noisy images.
Experience in Run II:
Switched to digital cameras optically coupled to image intensifiers (Gen II).
Potentially 300 frame/sec, high dynamic range, high SNR.
Issues:
Intensifier lifetime:
Photocathode worn out
-> I.I non uniform response
Mitigated with the implementation of the “spot painting”
(x,y camera TS)
Image intensifiers are used in the LHC mainly for gating capabilities (bunch by bunch)
Rather than for light amplification

## Slide 16
16
BSRT MONITOR
G. Trad – LHC Beam Size Measurement Review
Beam Emitted SR
SR Extraction
Optical System
Detector
Post-Process
Synchrotron radiation imaging system
Real time control system

## Slide 17
17
Control system
G. Trad – LHC Beam Size Measurement Review
Control server in FESA 3 operational working reliably since Run II.
Bunch by bunch measurements of full ring in less than 1 minute.
Provides “profiles” bunch by bunch and estimated beam size.
Manages the configuration changes from injection to flat top and the online feedbacks (MCP gain, filters, steering and focusing positions).
Single point failures of the communication due to radiation SEU are being addressed.

## Slide 18
18
BSRT performance
G. Trad – LHC Beam Size Measurement Review
BSRT cross calibration
Characterize the optical system (coupled to the SR source): Compute MAG, LSF
True beam size from reference instrument: Wirescanner
MAG calculated by performing closed orbit bumps at the source location.
Bump amplitude read from BPMs
LSF recovered as correction in quadrature of the scaled BSRT sigma (from Gaussian fit)
Run I
Run II
MAG and LSF computed by
Independent of BPMs
Requires a range of emittances in the LHC (typically from 1- 5um)

## Slide 19
19
BSRT performance
G. Trad – LHC Beam Size Measurement Review
BSRT cross calibration
Lens Scan [mm]
Lens Scan [mm]
At each position of the lens scan:
(MAG,LSF) computed
Focus defined as position with minimum LSF
Typical output of a calibration in Run II

## Slide 20
20
BSRT performance
G. Trad – LHC Beam Size Measurement Review
BSRT cross calibration
Calibration process (quadrature correction for resolution deconvolution) relies on:
Hyp1: LSF is Gaussian (by construction, it isn’t: trajectory effect in D3).
Hyp2: Beam is Gaussian (not always true, specially for small beam size)
Experimentally validated that such hypothesis do not deteriorate over the extreme range of emittances (1 um at Top energy)
324th LMC – 18/10/2016
G. TRAD
20
1.08
0.92
0.97
1.04
6.5 TeV

## Slide 21
21
BSRT performance
G. Trad – LHC Beam Size Measurement Review
Recalibration reproducibility B2
In terms of emittance:
Applying several calibrations starting from same raw data (sigma in pixels)
Emittance ratio
Emittance ratio
Emittance [μm]
Emittance [μm]
Any effect on changed optical functions in collision (as not measured) is folded into this “reproducibility”

## Slide 22
22
BSRT performance
G. Trad – LHC Beam Size Measurement Review
BSRT calibration during Ramp
Philosophy:
-> Keep the optics focusing and setup in FT configuration from 2TeV onward
-> Light source is shifting from undulator to dipole (interplay edge and core)
-> Work in “not on focus” condition: worst LSF.
-> Hope that the ratio LSF/σ stays relatively low.
-> Trust the WS single shot value as absolute true beam size at the scan energy
-> Hope that beta beating is not a limitation

## Slide 23
23
BSRT performance
G. Trad – LHC Beam Size Measurement Review
BSRT calibration during Ramp
The WS noise level improvement was promising to try a BSRT Calibration during RAMP. (from 2 TeV upward)
Automatic WS to cross calibrate along the cycle;

## Slide 24
24
BSRT performance
G. Trad – LHC Beam Size Measurement Review
BSRT calibration during Ramp
Optics used :
<2 TeV injection optics
>2 TeV collision optics
Consistent calibration factors could be used from E >3 TeV
Optics knowledge and precision of WS dominate the error contribution to the calibration in the ramp.
Accuracy turned out much worse than at FT > 20%

## Slide 25
25
BSRT performance
G. Trad – LHC Beam Size Measurement Review
Miscellaneous
Beta Beating caused by HO collisions never accounted for:
B1H = 3.5%,B1V = 2.5% B2H = -1%, B2V = -5.3%
Courtesy T.Pieloni
Optical system and magnification linked to slope and offset of

## Slide 26
26
BSRT performance
G. Trad – LHC Beam Size Measurement Review
Miscellaneous
Beta Beating caused by HO collisions never accounted for:
B1H = 3.5%,B1V = 2.5% B2H = -1%, B2V = -5.3%
Coupling effects are assumed small (can’t be observed) when 2D projection is ~round.

## Slide 27
27
BSRT performance
G. Trad – LHC Beam Size Measurement Review
Miscellaneous
Beta Beating caused by HO collisions never accounted for:
B1H = 3.5%,B1V = 2.5% B2H = -1%, B2V = -5.3%
Coupling effects are assumed small (can’t be observed) when 2D projection is ~round.
Dispersion is assumed null;
For small emittances ( 1μm)
Emittance calculation not accounting for dispersion (pure model) results into significant changes
Beam 1
LHC 2018 nominal cycle

## Slide 28
28
BSRT performance
G. Trad – LHC Beam Size Measurement Review
Miscellaneous
Beta Beating caused by HO collisions never accounted for:
B1H = 3.5%,B1V = 2.5% B2H = -1%, B2V = -5.3%
Coupling effects are assumed small (can’t be observed) when 2D projection is ~round.
Dispersion is assumed null;
Observation in 2017 (thanks to the new ADT activity monitor) showed a correlation between beam size from BSRT (sum of 200 2D images over 200 LHC turns) and the bunch oscillation.
=>Could be a limitation to the reproducibility of the calibration fills.

## Slide 29
Performance
29
G. Trad – 9th LHC Operations Evian Workshop
SR MONITORS
S. Papadopoulo
2018

## Slide 30
Performance
30
G. Trad – 9th LHC Operations Evian Workshop
SR MONITORS
M. Hostettler
2017

## Slide 31
Performance
31
G. Trad – 9th LHC Operations Evian Workshop
SR MONITORS
M. Hostettler
2016

## Slide 32
BSRT
32
G. Trad – 9th LHC Operations Evian Workshop
SR MONITORS
Ratio of Emittances from WS vs BSRT
When everything went smooth!
repeatability over 2 months period.

## Slide 33
33
G. Trad – 9th LHC Operations Evian Workshop
SR MONITORS
Performance

## Slide 34
34
G. Trad – 9th LHC Operations Evian Workshop
SR MONITORS
From the experiments
during a separation scan
Performance

## Slide 35
35
G. Trad – 9th LHC Operations Evian Workshop
SR MONITORS
Performance
Typical BSRT scan at top energy in 2018
The instrument is well able to resolve very small differences between bunches
Precision (rms) is <5% on emittance

## Slide 36
36
G. Trad – 9th LHC Operations Evian Workshop
SR MONITORS
Plans for Run III
Replace in vacuum extraction mirror
Replace view ports
Replace the mirrors with MgF coating layer on the table.
Pursue studies for improved measurement quality and reliability

## Slide 37
Summary
37
G. Trad – 9th LHC Operations Evian Workshop
The understanding of the system has matured over the 2 LHC runs
The system is operational since 2015 and has improved over the years :
  Reliable SR extraction mechanism
  Stable FESA control (automatic steering and settings management)
  Move to digital cameras and new generation of image intensifiers
  New calibration technique independent of BPM scaling errors
Relative accuracy (B by B) probably of the order of few %, absolute accuracy not easy to estimate:
  from comparison with OP scans and emittances estimated from luminosity probably about 10-20 %
Main problem is the system reliability, the measurement of its accuracy is intrinsically entangled with that of the WS.
Stability of calibration values and repeatability of the calibration Runs.

## Slide 38
Thank you for your attention!

## Slide 39
5th Evian Workshop
Transverse beam size measurement
BSR Interferometry – WHY?
Quadrature correction for the BSRT
Correction includes:
  Diffraction
  Aberration
  Magnification error…
How is determined ?
Calibrating with WS
Uncertainties?
% rel. Error
Uncertainties?
% rel. Error

## Slide 40
Courtesy: R. Tomas, M. Hofer et al.
Optics in IR4

## Slide 41
41
Experience in Run I-II
G. Trad – LHC Beam Size Measurement Review
Viewport
Extraction Mirror Reflectivity
Viewport
Extraction
Mirror
Viewport Reflectivity
Optical properties
BSRT working point @ FT

## Slide 42
Can the EXP tell more?
Horizontal plane
Biggest bunch
Horizontal plane
Smallest bunch

## Slide 43
LHC Profile Monitors Status- 6th Evian Workshop, G. Trad – 16/12/2015
BGI
WS
BSRT
BGV
CROSS-CALIBRATION
Y
slope
X
intercept
Using WS data
for magnification
Using BPM data
for magnification

## Slide 44
Can the EXP tell more?
Vertical plane
Biggest bunch
Vertical plane
Smallest bunch

## Slide 45
Can the EXP tell more?
Horizontal
Vertical
To be investigated further
Emit Scan
WS
Emit Scan
WS

## Slide 46
Emittance from WS vs BSRT
B1
Injection
Good agreement
(H+4% , V+2%)
FT
Bad as expected
B2
Injection
Good agreement
(H+5%, V+6.8%)
FT
Good agreement
(H-11%, V+12%)
Errors quoted for emittance:
2 um
B1H
B1V
B2H
B2V

## Slide 47: Synchrotron Radiation Monitors - BSRT
B1V @ 6.5 TeV
3 Cross-Calibration this year:
April-August-October
August Calibration was “fishy” still investigating
All calibrations made such as no data loss even if a calibration goes wrong
This year observed a dependence of the calibration on the beam mode @ 6.5 TeV
Investigate bigger spread on B1 emittances (Observations correlated with beam behavior?)
Study dependence of the calibration on machine conditions, optics/ beam modes