## Slide 1: Computed Tomography Technology + reconstruction Helge Pettersen Haukeland University Hospital
Notes: Add some details and mathematics on how X-ray systems work

## Slide 2: About me
Researcher HUS (imaging in therapy, pCT, …)
PhD in medical physics (proton CT): 2015 – 2018
Medical physicist working with CT 2012 – 2015
CT @ Haukeland:
Clinical work : Quality Assurance, optimization (better image quality & lower dose), teaching (UiB, HiB etc.), research and «technology pusher»
4 physicists responsible for 130+ x-ray machines at Helse Bergen, Helse Førde and private institutions
Notes: $ QA: check radiation output, radiation to adjacent rooms, dose meter calibration
$ optimization: Doctor complains about IQ / I check dose readings and’m unhappy about a specific lab

## Slide 3: Contents
This lecture (90 min): X-ray CT
  X-rays: Physics & production
  CT Technology
  CT Dose Reduction
  Image Reconstruction
    In 2D and 3D
Notes: Intro: Intro to CT, includes x-rays, geometry and a few examples

## Slide 4
An X-ray photon passing through the body may
Be unaffected (i.e. pass through the body)
Be absorbed (Photoelectric effect)
Be scattered (Compton) – causes image noise

## Slide 5: X-ray imaging
Transmission imaging
X-rays are produced at the X-ray tube
Passing through the patient
Hitting a detector on the other side
The fraction of photons hitting the detector is given by the total mass attenuation µ of the patient:
Notes: Index: differently attenuating organs or tissue

## Slide 6
Convoluted energy dependence in calibration range
Compton scatter dominates in clinical energy levels >30 keV

## Slide 7: X-ray tube
© Paul Frame, Oak Ridge Associated Universities
The Jackson focus X-ray tube: ca 1896
© The Cathode Ray Site
Concave cathode: better focus
Platinum anode

## Slide 8: Modern X-ray tube

## Slide 9: Anodes and cathodes
Separate filament circuit
Low voltage, 2-3 A
Creates electron cloud about cathode
cloud

## Slide 10: Anodes and cathodes
+70 kV
beam
-70 kV
Potential between cathode and anode accelerates to 70—140 keV
Anode is struck by «high energy» electrons

## Slide 11: Anodes and cathodes
Photon beam
X-ray beam is created by interactions between electrons in anode and anode material
Bremsstrahlung
Characteristic radiation by electron energies near the K edge
In addition, heating by ionizations
X-ray focus

## Slide 12: Energy spectrum (Birch & Marshall, 1979)
Electron beam () is gradually stopped in anode, sends out photons with energy and spectrum
Material constants
is the density
is Avogadro’s number
is the atomic number
Normalization
X-ray intensity
per unit energy interval
per incident electron flux
per atom
Material stopping power for incoming electron
Bethe equation
Attenuation for outgoing photon
is the attenuation coefficient
is the distance from γ production to anode exit: , anode angle
Relativistic corrections
Integration
From beam energy
To minimum energy to create photon with
Notes: Integrate the electron energy from incoming energy E0 to the min electron energy to create photon of energy Eγ.
The stopping power for the electron decreases the ENERGY and INTENSITY
the stopping power for the photons decreases the INTENSITY

## Slide 13
Unfiltered spectrum
Characteristic K-edge radiation
Bremsstrahlung radiation
kVp: kV peak

## Slide 14
http://www.qioptiq.com/x-ray.html
Hand Mit Ringen; W Röntgen 1896

## Slide 15
Siemens Definition AS
From siemens.com
Computed Tomography
X-ray tube and detector system rotates about the patient
Makes cross-sectional images

## Slide 16: Introduction
Detectors
X-ray tube
Filtration
Collimation
Dose-meter
Rotation 0.5- 4 Hz
Source: imgur.com 
High voltage generator
Notes: Video: 2 Hz (0.5 sec/rotation)

## Slide 17
Axial projection
Coronal projection
Sagittal projection
Notes: Note: No over-/underlying rib bones, better contrast of the lung anatomy and heart, can choose projection plane

## Slide 18: Timeline
1895: Wilhelm Röntgen discovers X-rays
1917: Johann Radon develops mathematical method of cross-sectional reconstruction from transmission measurements
1963: Allan Cormack describes technique for calculating the absorption distribution of the human body
1971: Godfrey Hounsfield develops the first CT scanner
1975: First whole body clinical CT
1976: Filtered Back Projection introduced
1989: William Kalender + Peter Vock develops helical CT
1998: 4-slice CT scanners
2005: Dual Source CT
2008: Iterative Reconstruction methods; 320-slice CT scanners

## Slide 19: 1971: EMI head scanner by Hounsfield
23.5 cm FOV, 3x3x8 mm voxel size (80x80 image matrix) 5 min scan, 5 min reconstruction NaI-PMT detector

## Slide 20
t=0
+ 3 år
Medical imaging - CT

## Slide 21: CT number scale
Grey levels on CT images represent the mass attenuation in each voxel
Hounsfield Units (HU):
Air = -1000 HU; Water = 0 HU

## Slide 22: CT windowing (4096 greyscale values)
https://sites.google.com/a/wisc.edu/neuroradiology/image-acquisition/the-basics
Notes: Axial images, seen from the bottom
Lung: Bronchial tubes (for air) and blood vessels

## Slide 23: Planar X-ray vs CT
Planar X-ray © Yale University, 2004
Coronal MIP-projection of CT image
RadioGraphics 5, 25 (2005) 1409-1428
Overlying and underlying rib bones, no depth selection
Excellent tissue contrast
Notes: Note: No over-/underlying rib bones, better contrast of the lung anatomy and heart, can choose projection plane

## Slide 24: CT images: Head&neck
https://www.imaios.com/en/e-Anatomy/Head-and-Neck/Head-and-neck-CT
Notes: Axial images, seen from the bottom
Lung: Bronchial tubes (for air) and blood vessels

## Slide 25: Functional imaging
Blood flow analysis – clot or bleed?
Static image (subtraction angiography)
Dynamic image (blood flow/volume time analysis in multiple scans)
Notes: Her satser Toshiba

## Slide 26: Dose calculation in radiotherapy
https://www.wienkav.at/kav/kfj/91033454/physik/eclipse/rapid1.htm
Notes: Her satser Toshiba

## Slide 27: Today: 4 main vendors
SIEMENS
GE
Canon
(aka Toshiba)
Philips

## Slide 28: CT technology

## Slide 29: X-ray detectors
The radiation through the patient must be
measured.
Image basis: Non-absorbed X-rays
X-rays
Scintillator
The rays pass through a scintillator. The x-rays are transformed to visible light and amplified
Visible light
Detector
A detector transform the visible light to electrons (making a current)
Current
ADC
Amplifiers and analogue-to-digital converters (ADC) create a digital signal from the current.
Digital signal

## Slide 30: X-ray detectors
Webb, S. and M. A. Flower (2012)
Anti-scatter grid (Pb/W)
Accept only X-rays from focus
Scintillator (CsI:Tl)
Photo diode (a-Si): light to electrons
The thin film transistor (TFT) chip stores the electrons in a capacitor, until a release current (switch) is applied

## Slide 31: The scintillator
An incoming photon liberates e/hole pairs
  Compton scattering + PE
  Energy gap
Random movement until they find an (doping) activator and are recombined
The activator releases visible light when hit
  e/hole pair is reunited and energy is released
  CT needs very quick scintillators
Nikl M. Meas. Sci. Technol. 17 (2006) R37-R54
Notes: Små mengder: Mellom 1:10^8 og 1:10^4

## Slide 32: CT detectors
Goldman L, Principles of CT and CT Technology. J Nucl Med Technol 2007; 35:115-128
Thin fan-beam
in z direction
Thick fan-beam
in z direction

## Slide 33: CT detectors
https://www.itnonline.com/content/detection-technology-introduces-shelf-tileable-ct-detector-module
https://anz.medical.canon/products/computed-tomography/aquilion_lightning_80_technology

## Slide 34: Image Acquisition Axial vs Helical
https://www.radiologycafe.com
Pitch: Distance between slices
Pitch:

## Slide 35: Dose Reduction Tools in CT ( constant noise)

## Slide 36: The Bow Tie filter
Equal image noise thorughout the image
Usually proprietary information (= hard to do Monte Carlo simulations)

## Slide 38: Automatic Exposure Control
A 4 cm increase in patient size increase the attenuation by 2
45 cm vs 28.8 cm: 16x more attenuated, 4x more noise
Sedlmair, M. (Siemens) «CAREDose4D, CAREkV and SureView» Forchheim, 2013-10-22

## Slide 39: Automatic constant mAs
Constant mAs based on weight or 2D «scout» image
  mAs: milliAmps * irradiation time
Adjusted to the patient
No modulation during scan (in z dir)
Less image quality variation between patients
Söderberg M., PhD dissertation (2012), Lund University

## Slide 40: Z modulation
Can use the 2D scout to see what mAs we need in each slice
Integrated radiation dose
Dose is adjusted in Z direction
Reduced DLP (total dose)
Constant noise in Z direction
Can adjust the mean dose manually
McCollough,H.C. et al., RadioGraphics 2006; 26:503-512

## Slide 41: 3D modulation
By using a lateral scout as well (2 scout images)
3D information about attenuation
Apply a sine shaped mAs modulation to match rotation speed
Keep signal to detector constant at all times
Quickest variation in shoulders
Kortesniemi M., NACP-Foredrag i Bergen (2013)
Lee C. H. et al., RadioGraphics (2008), 28:1451-1459

## Slide 43: Dual Energy CT

## Slide 44: Functional imaging – dual energy
Using a whole spectrum to obtain images gives a good estimate for the total attenuation;
However, the body consists of more than one …
And
At low energies the photoelectric effect dominates
and heavy materials do lots of absorption
At high energies the Compton scattering dominates
Less Z dependence

## Slide 45: Attenuation and energy
Kaza R K et al. Radiographics 2012;32:353-369
Calcium
Water
Iodine: High variation with E

## Slide 46
Need large (spectral separation) to separate low
http://radiologyinthai.blogspot.no/2010/12/dual-energy-ct-2.html
http://radiologyinthai.blogspot.no/2010/12/dual-energy-ct-2.html
Spectrum separation
Notes: Her må jeg skrive mer om Sn

## Slide 47: Spectrum separation
Multiple ways to create a dual energy spectrum
Quick change between high / low kVp (between each rotation or contiuous)
Two physical x-ray tubes
Sandwich detector
Photon couting detector (not yet in clinical use for CT)

## Slide 48: Spectrum separation
With two physical x-ray tubes (dual source)
Can use filtering (tin / Pb / W) on one of the tubes
100 kVp + Sn 140 kVp
More separated effective energies
100 kVp = 58 keV
Sn140 kVp = 81 keV
Johnson T. R. C, AJR2012, 199:S3-S8
Effective energy
Effective energy
Notes: Her må jeg skrive mer om Sn

## Slide 49: Basis Material Decomposition
Each voxel with energy E as a linear combination of two basis materials and :
is the mass attenuation
is the density of the material – a weighting factor

## Slide 50: Basis Material Decomposition
Use this information in
Insert for
We have and
Solve for and , we get
A purely image, e.g. bone
A purely image, e.g. water

## Slide 51
Yu L et al. AJR2012, 199:S9-S15
Bone (Ca, Z = 20)
Water (H, O; Z = 1, 8)
Raw data
Image

## Slide 52: Monochromatic images
We know how the voxels vary with energy
Can make a monochromatic image
Very good HU accuracy and beam hardnening corrections
PolyChr 120 kVp
MonoChr 127 keV
Yu L et al. AJR2012, 199:S9-S15

## Slide 53: What more? Applications?
We know
The materials in the data
Can make three kinds of images
Optimal images: Choose the MChr with best contrast
  E.g. iodine uptake but with low noise
Identification of materials (ionide, bone, …)
  Bone removal
  Virtual non-contrast images
  Color code materials (uric acid)
Quantification of materials

## Slide 54: Quantification of iodine
Johnson T. R. C, AJR2012, 199:S3-S8
Sn 140 kVp
100 kVp
More yellow, more iodine / calcium
Comparison

## Slide 55: Materials removal
Original image
Contrast agent removed
Bone removed
Johnson T. R. C, AJR2012, 199:S3-S8

## Slide 56: Applications
Kidney stone classification
Thomas C. et al., Radiology 2010; 257:402-409

## Slide 57: Applications
Gout
Nicolaou S. et al. AJR2012, 199:S78-S86

## Slide 58: Applications
Plaque (red) in arteries filled with contrast agent (blue)
«Dual Energy CT – SOMATOM Definition» Siemens whitepaper

## Slide 59: Dual Energy CT and particle therapy
A new application for DECT is its application in particle therapy
Particle therapy: Electron density / atomic number, ionization potential
CT: Choose one (by calibration)
Dual Energy CT: Calculate both from the two spectra!

## Slide 60
Accuracy of SECT vs DECT

## Slide 61: Image reconstruction
Notes: Bør har kommet hit til pausen!

## Slide 62: Image reconstruction
Impact group
Use a series of 2D projections of an object to calculate the 3D volume
Notes: OVERVIEW, not explain the concepts

## Slide 63: Image acquisition
Fig: Webb’s Physics of Medical Imaging (2012)
Notes: This is the true value of µ(x,y)

## Slide 64: Image acquisition
The attenuation map corresponds to one row in the raw data / sinogram
When the gantry is rotated 1°, we get a new row in the sinogram
We are dealing with a «rotating geometry»
A small structure in the object is curved in the sinogram
sinogram
Notes: OVERVIEW, not explain the concepts

## Slide 65
Projection direction

## Slide 66: Image reconstruction
The reconstruction problem: How to reconstruct the image from the sinogram?
?
Notes: OVERVIEW, not explain the concepts

## Slide 67: 2D Image Reconstruction
Two types of reconstructions:
  Iterative reconstructions
    Algebratic Iterative Reconstruction
    (Model Based Iterative Reconstruction)
  Projection methods
    Filtered Back Projection (FBP)

## Slide 68: Algebraic Iterative Reconstruction
Historically, the first CTs used Algebraic Iterative Reconstructions from the sinogram
Illustrates the idea of modern iterative algorithms
Solve a set of linear equations
Scales poorly with large images
Noise sensitive
Input data: The projections of the image matrix
Output data: The matrix itself
 | 5 | 6 | 3 | SUM X1
 | 5 | 3 | 4 | SUM X2
 | 5 | 6 | 7 | SUM X3
SUM DIAG | SUM Y1 | SUM Y2 | SUM Y3 | SUM DIAG

## Slide 69
 | 3.3 | 5 |  | 7
 | 3.3 | 5 |  | 20
 | 3.3 | 5 |  | 10
12 | 10 | 15 | 8 | 13
 | 2.75 | 3.65 | 2.5 | 7
 | 5 | 5.85 | 4.7 | 20
 | 3.3 | 4.15 | 3 | 10
12 | 10 | 15 | 8 | 13
 |  |  |  | 7
 |  |  |  | 20
 |  |  |  | 10
12 | 10 | 15 | 8 | 13
 | 3.3 |  |  | 7
 | 3.3 |  |  | 20
 | 3.3 |  |  | 10
12 | 10 | 15 | 8 | 13
 | 3.3 | 5 | 2.7 | 7
 | 3.3 | 5 | 2.7 | 20
 | 3.3 | 5 | 2.7 | 10
12 | 10 | 15 | 8 | 13
 |  |  |  | 7
 |  |  |  | 20
 |  |  |  | 10
12 | 10 | 15 | 8 | 13
 | 2.3 | 2.3 | 2.3 | 7
 | 6.7 | 6.7 | 6.7 | 20
 | 3.3 | 3.3 | 3.3 | 10
12 | 10 | 15 | 8 | 13
 |  |  |  | 7
 |  |  |  | 20
 |  |  |  | 10
12 | 10 | 15 | 8 | 13
2. Apply average projection values along all axes (vertical)
3. Apply average projection values along all axes (horizontal)
 | 2.3 | 2.3 | 2.3 | 7
 |  |  |  | 20
 |  |  |  | 10
12 | 10 | 15 | 8 | 13
 | 2.3 | 2.3 | 2.3 | 7
 | 6.7 | 6.7 | 6.7 | 20
 |  |  |  | 10
12 | 10 | 15 | 8 | 13
1. Projections
4. The iterative process
For each pixel, find the average solution
If RMSE is below a limit: stop

## Slide 70: Projection methods
Projection-Slice theorem: Given infinite 1D projections along infinite lines, we can reconstruct perfectly
  Called the Inverse Radon Transform
Where is a line along and
Is the convolution between the attenuation map and a «ramp filter»

## Slide 71
Or, for the mathematically uninclined…

## Slide 72: Back projection
For each row in the sinogram, project the data onto an empty image (corresponding to the angle )
Sinogram
Back Projection
Notes: OVERVIEW, not explain the concepts

## Slide 73: Back projection
OK image with enough projections (-> 360)
Due to inherent problem with method, edges are blurry and smoothed
  Very evident in complex images
Need to find a way to remove this artifact
Filtering
Notes: OVERVIEW, not explain the concepts

## Slide 74: Filter
A filter is a mathematical operation which adjusts an image matrix
Is itself a matrix
Works on every pixel, but with local information (filter size – 3x3)
Above: For each pixel, 1*current value + 0*every neighbour value
Below: For each pixel, sum row pixel values, ignore above/below. This will smooth horizontally
Notes: OVERVIEW, not explain the concepts

## Slide 75: Filter
Many different possibilities. Normalization ++
We are interested in edge enhancing
-1: Enhance differences between neighbouring pixels
Result: Higher contrast edges
Chart: 
Chart: 
Notes: OVERVIEW, not explain the concepts

## Slide 76: Different filters
Horizontal smoothing
Edge enhancing
Notes: OVERVIEW, not explain the concepts

## Slide 77: Filtered back projection
Every sinogram row is filtered with a edge enhancing filter
Can choose between different filters to obtain smooth or sharp result images
Notes: OVERVIEW, not explain the concepts

## Slide 78: Filtered back projection
The reconstructed image no longer has a «ghost»
Notes: OVERVIEW, not explain the concepts

## Slide 80: Reconstruction filters
Different reconstruction filters are choosen
by technicians at scan time.
Visibility of small lesions
Image texture
  Amount of noise
  Texture of noise
  Low contrast object visibility
Choose filter based on diagnostic task
Notes: OVERVIEW, not explain the concepts

## Slide 81: Caveat emptor… Fan Beam FBP
Path length correction:
: Distance between x-ray focus point and detector*
However: All the information is in the image
Figure: Kachelrieß, M. et al., 2004. Med. Phys. 31, 1623–1641.
Notes: * Depending on your frame of reference, some gives focus point <-> isocenter

## Slide 82: Reconstruction techniques
… in 3D …
Notes: OVERVIEW, not explain the concepts

## Slide 83: Reconstruction techniques
2D reconstructions:
Reconstruction plane = Detection plane
FBP as shown works nicely
3D reconstructions:
Reconstruction plane Detection plane
  Large -direction coverage, thin detectors
  FBP gives image artefacts
Notes: OVERVIEW, not explain the concepts

## Slide 84
z direction
Single-slice
Four-slice
Multi-slice
Cone-beam artefacts

## Slide 86: 3D Reconstruction techniques
Exact Cone Beam CT
  Tuy Conditions: Cannot use axial or helical source path; need two orthogonal rotations
  Katsevich (2003) simplified: Clinically available in Canon ConeXact™
Approximate Cone Beam CT
  Single Slice Rebinning (2D FBP with tilted planes)
  Feldkamp / Davis / Kress (FDK)-type algorithms
Notes: Two more Tuy conditions: X-ray tube cannot enter patient; Must follow a sensible path around the patient

## Slide 87: Single Slice Rebinning
One way of solving the 3D problem (remember the end result is parallel 2D slices): Make approximate 2D slices
2D filtered back projection with tilted plane
Castellano E, Reconstruction techniques for multi-slice and cone-beam CT (2013)
Notes: OVERVIEW, not explain the concepts

## Slide 88: Single Slice Rebinning
SSR is used for < 30 mm helical reconstructions in some CT scanners (Siemens + GE) For < 30 mm + axial, 2D FBP is used. (2013) [Canon + Philips uses modified FDK for < 30 mm; all use modified FDK for 30 mm]

## Slide 89: FDK-type algorithms
A 3D generalization of FBP
Assume parallel rays, but make geometric correction for increased pathlength
Use 1D ramp filter on row-by-row basis (same )
3D back-projection (cone projection)

## Slide 90: FDK-type algorithms
Remember that ,
Where is the filtered 1D attenuation map
Parallel beam:
Fan beam:
Cone beam (FDK):
  Still 1D integral: row-by-row convolution w/ geometric correction
  : Ratio between the path length of parallel and cone beam

## Slide 91: 2008 and onwards: Resurgence of Iterative Reconstruction
Added computational power: «New» IR
Projection – Back Projection workflow
  Statistical noise modelling
  Model system properties for forward projection

## Slide 92: Artifact reduction: Cone Beam artifact
Raupach, R. (Siemens) IRIS/SAFIRE
Weighted FBP
IR + Weighted FBP

## Slide 93
Fro
  : Regularization step: Expected 2D noise image
Figure by Mika Kortesniemi

## Slide 94: Different IR solutions
Vendor | Iterative Image Filter | Statistical Noise Reduction | Model Based IR
GE | VISIR | ASIR-V | VEO
Siemens | IRIS | SAFIRE | ADMIRE
Philips |  | iDose | IMR
Toshiba | AIDR | AIDR 3D | First

## Slide 95: Model Based Iterative Reconstruction
Model the physics in the forward projection
Artifact reduction
Model Based Iterative Reconstruction (MBIR)
Physical size of focal spot
Physical voxel sizes
Correct beam shape: Cone beam
Detector- and scintillator response
Simulate transport of photons through the system
  Full Monte Carlo… Can be very time demanding!

## Slide 96
Philips IMR
GE VEO
Siemens Admire
Canon FIRST

## Slide 97: Thanks for your attention

## Slide 98: Recommended reading
General overview Reconstruction algorithms History of medical imaging
(anno 1988 but still good)