## Slide 1: The Bergen proton CT project
Helge Pettersen, Haukeland University Hospital
Notes: Add some details and mathematics on how X-ray systems work

## Slide 2: 3-slide introduction if Pierluigi did bad Proton Computed Tomography

## Slide 3: X-ray CT
X-ray tube
X-ray detector
,SP

## Slide 4: Proton CT
High energy
proton beam
Energy detector
,SP
Position detectors
Position detectors

## Slide 5: The energy detector
Range telescope
Scintillator

## Slide 6: Proton CT with a Digital Tracking Calorimeter First proof-of-concept detector

## Slide 7
pCT – Digital Tracking Calorimeter
Patient
Measure multiple protons’ vector, position & track protons to find range
Digital Tracking Calorimeter
Beam
Assume proton vector & position

## Slide 8
Beam tests in CERN + Groningen, NL (2014)

## Slide 9
Protons hitting the pixel detectors
Pixel Position Y [Pixel number]
Pixel Position X [Pixel number]
Notes: You have to play «Connecting the Dots», which, incidentally, is also the name of an internationally held workshop specializing on track reconstruction

## Slide 10: Analysis workflow

## Slide 11: Proof-of-concept: ALICE-FoCal
From 2013: Analysis + beamtests with the ALICE-FoCal detector for pCT
Expectedly poor resolution
Track reconstruction possible at high rates
  ~1 million/s across 16 cm2 @ 2 kHz readout

## Slide 12: The ranges from many protons
Pettersen, H.E.S. et al., 2017. Nucl. Instr. and Meth. in Phys. Res. A 860, 51–61.

## Slide 13: Proton CT with a Digital Tracking Calorimeter Second prototype

## Slide 14
Alice Pixel Detector (ALPIDE)
http://bora.uib.no/handle/1956/16041: Design of High-Speed Digital Readout System for Use in Proton Computed Tomography
More on ALPIDE in 34 minutes!

## Slide 15: Monte Carlo simulations
A very important tool in detector development
  Define geometry, materials, beam, readout…
  Simulation of how the particles interact and how energy is deposited

## Slide 16: GATE
Geant4-based MC framework
Macro-based, simple to automate, e.g. to test iterative designs:

## Slide 17: Charge diffusion from proton tracks

## Slide 18
Protons hitting the pixel detectors
Pixel Position Y [Pixel number]
Pixel Position X [Pixel number]
Notes: You have to play «Connecting the Dots», which, incidentally, is also the name of an internationally held workshop specializing on track reconstruction

## Slide 19
Charge diffusion in pixels
Pettersen, H.E.S., 2018. PhD thesis

## Slide 20: Heidelberg beam test: 2018
A list of 15 000 «good» clusters
+ Calculation of the ions’
Grimstad, S., 2019. MSc: Simulation and analysis of clustering for proton CT
Tambave, G. et al., 2019. Nucl Instr And Meth in Phys Res A; 162626

## Slide 21
Tambave, G. et al., 2019. Nucl Instr And Meth in Phys Res A; 162626

## Slide 22: Compare this to MIMOSA (PoC)
Pettersen, H.E.S., 2018. PhD thesis
Notes: So took all the clusters from the experimental data and characterized them regarding a) their residual energy, which gives the energy deposited in that layer and b) their cluster size after the different chips have been calibrated to yield similar cluster size distributions. This is a correlation plot between the two, and with this I’ve applied different models to describe the correlation. The First Principles-model was customized for this usage by the help of a master student.

## Slide 23: MC modeling of Cluster Sizes
Not feasible to directly simulate in MC
MSc of Silje Grimstad:
Use relashionship from Heidelberg data
Use collected library of clusters, sorted by
MC Find in pixel Sample random cluster with correct
Paint cluster around hit position
Do cluster analysis (hit position from cluster…)

## Slide 24
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.

## Slide 25: Accuracy of cluster reconstruction
Grimstad, S., 2019. MSc: Simulation and analysis of clustering for proton CT

## Slide 26: Track reconstruction

## Slide 27: Proton track reconstruction
With a DTC-type detector:
  Possible to reconstruct many protons in a single readout
  All proton hits throughout the detector must be «de-spaghetti-fied» into tracks

## Slide 28: Tracking algorithm
Identify all hits in first/last layer
Find angular change in all possible combinations of hits in first/last + next layer. Accumulate : for pair
Identify all tracklet candidates where . Keep the best (or the two best if sufficiently similar).
Recursively grow the tracklet tree to reconstruct all tracks
Pettersen, H.E.S. et al., 2019. EPJ WoC (in press).

## Slide 29: Stuff we’ve tried
Backwards + forwards tracking
  Backwards slightly better
scoring w/per-layer Highland
Follow all tracks where
  Scales «poorly» with nlayers and ntracks
Track-splitting corrections; cluster merging corrections
Kalman filtering: MCS uncertainties too large compared to measurement uncertainties

## Slide 30
Good track
Incomplete track
Unfiltered secondary
Two tracks confused
Unused cluster
Garcia-Santos, A., 2019. Utrecht University, Delft. MSc: Optimization of the Track Reconstruction Algorithm in a Pixel Based Range Telescope for Proton Computed Tomography

## Slide 31
Pettersen, H.E.S. et al., 2019. EPJ WoC (in press).

## Slide 32
Aluminum absorber thickness (final = 3.5 mm)
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.
Now assuming a pencil beam & 10 µs integration time

## Slide 33: What are the degrading effects?
Pettersen, H.E.S. et al., 2019. EPJ WoC (in press).

## Slide 34: Track filtering
Need filter nuclear interactions in detector/patient
    3 sigma angular filter
    3 sigma filter on range per beam / voxel
    Filter on in the last layer
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.

## Slide 35
Volz, L. et al., 2019. Phys. Med. Biol. 64, 15NT01 uses filtering to remove 2ndies in He beam
Pettersen, H.E.S., 2018. PhD thesis

## Slide 36: What happens to bad tracks?
500,000 protons tracked with
286,925 remaining after filtering (w/ 11.5% 2nd)
84% are correctly reconstructed:
  10% of the tracks are confused (>70 % due to cluster merging)
    Image noise
  5% are incomplete
    Systematic lower WEPL
  1% are confused + incomplete
    Image noise

## Slide 37: Range calculation

## Slide 38
Finding the proton range
Pettersen, H.E.S., 2018. PhD thesis

## Slide 39
The Bragg-Kleeman range-energy relationship, , can be rewritten to give an accurate depth-dose curve for a single proton:
Pettersen, H. E. S. et al., 2018. Radiat. Phys. Chem. 144, 295–297.

## Slide 40
Range resolution from individual tracks
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.

## Slide 41: Range accuracy and precision
Range straggling
Range straggling
+
detector
uncertainty

## Slide 42
Range distribution per beam energy
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.
Notes: The thinner the absorber layers, the more sensor layers the Rang Straggling Distribution cover, increasing the resolution & accuracy

## Slide 43
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.

## Slide 44
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.

## Slide 45
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.
Design choice: 3.5 mm

## Slide 46: Number of layers needed
Pettersen, H.E.S. et al., 2019. Physica Medica 63, 87–97.

## Slide 47: What about images?

## Slide 48: The Most Likely Path scheme
The Most Likely Path (MLP) of the proton through the object
A model fit to the entry and exit angle/position
Use the curves for image reconstruction
Li, T., et al., Reconstruction for proton computed tomography by tracing proton trajectories: a Monte Carlo study. Med Phys, 2006. 33(3): p. 699-706
More on MLP on Thursday!

## Slide 49: How to find the entrance position?
Krah N et al. (2018) PMB 63 (13)
Also Pettersen HES, Volz L et al. Rejected by Phys Med Biol 
MC tracks
Estimated tracks
Use middle value of pencil beam
Estimate from front position (not too bad eh!)
Measure before the patient…
More on single-sided imaging on Friday!

## Slide 50: References
Garcia-Santos, A. (2019). MSc: Optimization of the Track Reconstruction Algorithm in a Pixel Based Range Telescope for Proton Computed Tomography. Utrecht University.
Grimstad, S. (2019). MSc: Simulation and analysis of clustering for proton CT.
Pettersen, H.E.S. (2018). PhD: A Digital Tracking Calorimeter for Proton Computed Tomography. University of Bergen.
Pettersen, H.E.S., Alme, J., Biegun, A., van den Brink, A., Chaar, M., Fehlker, D., Meric, I., Odland, O.H., Peitzmann, T., Rocco, E., et al. (2017). Proton Tracking in a high-granularity Digital Tracking Calorimeter for proton CT purposes. Nucl. Instr. and Meth. in Phys. Res. A 860, 51–61.
Pettersen, H.E.S., Chaar, M., Meric, I., Odland, O.H., Sølie, J.R., and Röhrich, D. (2018). Accuracy of parameterized proton range models; a comparison. Radiat. Phys. Chem. 144, 295–297.
Pettersen, H.E.S., Alme, J., Barnaföldi, G.G., Barthel, R., van den Brink, A., Chaar, M., Eikeland, V., García-Santos, A., Genov, G., Grimstad, S., et al. (2019a). Design optimization of a pixel-based range telescope for proton computed tomography. Physica Medica 63, 87–97.
Pettersen, H.E.S., Meric, I., Odland, O.H., Shafiee, H., Sølie, J.R., and Röhrich, D. (2019b). Proton Tracking Algorithm in a Pixel-Based Range Telescope for Proton Computed Tomography. EPJ Web of Conferences In press.
Pettersen, H.E.S., Volz, L., Sølie, J., Rohrich, D., and Seco, J. A Linear Projection Model to Estimate a Proton’s Position in a Pencil Beam For Single Sided List Mode Proton Imaging. Rejected PMB
Tambave, G., Odland, Odd Harlad, Pettersen, H.E.S., Alme, J., Genov, G., Grøttvik, O.S., Samnøy, A.T., Röhrich, D., Ullaland, K., Ur Rehman, A., et al. (2019). Characterization of Monolithic CMOS Pixel Sensor Chip with Ion Beams for Application in proton Computed Tomography. NIMA.
Krah, N., Khellaf, F., Rit, S., and Rinaldi, I. (2018). A comprehensive theoretical comparison of proton imaging set-ups in terms of spatial resolution. Physics in Medicine and Biology 63.

## Slide 51
Thanks for your attention 