## Slide 1
Technical workshop 2020
23/01/2020
Yann Donon
SmartLINAC
Notes: New project
Samara University, Russia : expertise in control system, in particular in the aerospace

## Slide 2: LINACs mecial applications
SmartLINAC
Electron linear accelerators are today the core of modern radiation therapy facilities
More widespread diffusion of this technology is prevented by maintenance cost
Maintenance cost is conditionned by harsh environments and lack of on-premise technical skills
“CERN-ICEC-STFC-hosted Workshop on Innovative, Robust and Affordable Medical Linear Accelerators for Challenging Environments”
Notes: In research, at CERN, they are used as booster for circular accelerators

## Slide 3: SmartLINAC project goals
SmartLINAC
Increase availability, reduce operational costs:
  Lower the technical skills maintenance requirements by:
    Fault diagnosis allowing local ressources to make guided interventions
  Planned maintenance by:
    Fault prediction (machine learning)
Increase scientific LINACs reliability to 99%
Bring medical LINACs reliability up to scientific’s reliability!!!

## Slide 4: Personalized Preventive maintenance
Initial models based on MTTF and know fault patterns
Models trained on existing data from LINAC2/4
1
5
Initial maintenance plans
Real LINACs data
In different environments and operation constraints
3
Real data used to refine models
4
Personalized plans and models
2

## Slide 5: The LINAC4 RF power source data
Notes: Challenges:
Understanding the nature of the data
Labelling the data
Understanding symptoms of anomalies
Highlighter: jitters (source decay)

## Slide 6: The LINAC4 RF?
Chart: 
| Accelerator Controls | Beam Losses | Cooling and Ventilation | Electrical Network | Machine Interlock System | Operation | Other | Power Converters | Pre-Chopper | Radio Frequency | Source
cdt | 4.4458 | 24.3883 | 1.2333 | 0.81667 | 0 | 4.1492 | 1.65 | 49.9217 | 21.1944 | 80.3622 | 29.1767
Notes: Why those data?
36,98%
Very representative data sample

## Slide 7: The noise
Capture
Human
Jitters
Volatility
Notes: Different kind of noise,
Capture errors
Human interactions
General volatility
Jitters
Jitters are interesing as they have been found to be symptomatic of quality decay in the beam

## Slide 8: The (current) solution: SNiF
Notes: neural network, KDE, ARI, statistical features analysis and Kalman filtering, none could tackle the challenge but all served well in understanding the data.
Proximity areas, prior to jitters with unique statistical properties
Many sources of false positive and false negative
New technique, based on a process developped for feature detection in images (BIM)
outliers filtering : Grubb’s test
Normalization the dataset: and Z-score algorithm
Smoothing using a Gaussian blurring kernel
Threshold: higher than average variance by sliding windows
Stochastic matrix comparison

## Slide 9: Future steps
Testing with more data from different sources
Building new collaborations
Test, train, repeat!
Platform

## Slide 10: Thanks!
yann.donon@cern.ch