## Slide 1: Recap: Beam diagnostics
Hao Zhang (hao.zhang@Cockcroft.ac.uk）

## Slide 2: Outline
Why beam Diagnostics?
Challenges in Diagnostics?
Instrumentation type
  Beam position monitoring
  Beam profile monitoring
  Emittance measurement
  Beam current monitoring
  Beam energy monitoring
  Beam halo monitoring

## Slide 3: Demand for beam diagnostics
Essential for any accelerator
‘Organ of sense’
Deal with real beam with imperfections
Three categories
  Reliable, quick measurements for general functionality, e.g. current transformer
  For a daily check of performance and stability, e.g. profile determination
  Complex instrumentation, e.g. Emittance measurement
P. Forck, Lecture Notes on Beam Instrumentation and Diagnostics, Joint University Accelerator School, January – March 2012
Notes: Reliable, quick measurements to determine the basic parameters of a machine setting used as a fast check of the general functionality
Instrumentation built for a daily check of performance and stability, or for the control of wanted parameter changes in the accelerator setting.
Complex instrumentation for the commissioning of a new accelerator component, for the
development of higher performance and for solving more serious problems in case of a malfunction.

## Slide 4: A ‘typical’ Accelerator Diagnostics
Material sciences
Thermodynamics
Electro-Magnetism
Optics
laser technology
Mechanics
Electronics
Nuclear Physics
Vacuum science
…
transformer
Pick-up
Slit-grid emittance
Residual gas monitor
SEM-grid
transformer
Pick-up
Faraday cup
Pepper-pot
Multi-disciplinary field !

## Slide 5: Challenge in beam diagnostics
Better resolution (time/space)
(single) photon resolution
Improved radiation hardness
invasive
Able to cover variety of beams

## Slide 6: Beam Position Monitoring

## Slide 7: Capacitive Pickups
d
The Capacitive Pickups are based on symmetry
  The beam displacement x is detected by a pair of symmetrical arrange electrodes
courtesy O.R. Jones
–
x
Examples:
Shoebox type
Buttons
Striplines

## Slide 8: Electro-Optic Beam Position Monitors
courtesy S. Gibson

## Slide 9: Beam Current Monitoring

## Slide 10: Faraday cup
The Faraday Cup destructively intercepts the beam
DC coupled! (With just a resistor the signal is Vout = Ibeam*R)
Low current measurements possible e.g., USR fA!
Problem with secondary electrons:
  Use long cup or voltage suppression or magnetic field
If not properly terminated ‐> very high voltage (beam potential)
Must process beam power
Notes: ultra-low energy electrostatic storage ring
Limitations:
Beam energy ?
Sensitivity/noise ?
Antimatter ?

## Slide 11: Faraday cup examples
Simple faraday cup in my lab
For higher intensities
water cooling may be needed

## Slide 12: Beam transformer
Courtecy: W. Blokland

## Slide 13: Wall-Current Monitors
The image current accompany the beam current.
A voltage proportional to beam current can be measured across the resistors
Also the noise issue will limit the resolution

## Slide 14: Cryogenic Current Comparator (CCC)
Non-destructive measurement device for low intensity beams (between 1 nA to 10 μA)
The CCC consists of:
  Cryogenic device
  SC pickup coil,
  Highly efficient SC shield,
  High performance SQUID
  measurement system.

## Slide 15
Beam Profile Monitoring

## Slide 16: Scintillating screen
Simply
Cheap
Effective
Credit: B. Walasek-Höhne, GSI and G. Kube, DESY
Notes: Linked to ionisation
• Photons are emitted by the deexcitation
of atomic states
populated by the passage of
the particle
• Emission time ns to hours

## Slide 17: Scintillating materials
Notes: Phosphors have very high light yields, but can only be
used as thin coating on a rigid support and get damaged
very quickly
• Normally used only for very low intensity beams
• Ceramics, glasses and crystals are a more popular choice
in high energy accelerators
• CHROMOX (Al2O3:CrO2, Aluminium Oxide) is a very
common choice because it is a very robust ceramic
• YAG (Y3Al5O12) is also a very frequent choice (fast)
18

## Slide 18: Some example of scintillating light
Notes: Caesium iodide dotted with Titanium
YAG dotted with Cerium
aluminum oxide dotted with Chromium
Zirconium oxide dotted with Magnesium

## Slide 19: Optical Transition Radiation
Radiation is emitted when a charged particle crosses the boundary of different dielectric properties
Radiation has defined angular distribution
Radiation is radially polarised
Credit: D. Morris, et al., TRIUMF
E. Bravin, Cern Accelerator School on Beam Instrumentation
2-15 June 2018, Tuusula, Finland
Notes: Radiation is emitted when a charged particle crosses the
boundary of different dielectric properties
• Radiation has defined angular distribution
• Radiation is radially polarised
• Thickness of radiator not important

## Slide 20: Wire scanner
E. Bravin, Cern Accelerator School on Beam Instrumentation 2-15 June 2018, Tuusula, Finland
SLAC SLC high resolution 3 axis scanners
CERN “flying wires”
Scans a thin wire or a needle across the beam. Detects secondary emission current or high energy secondary particles (scintillator + PMT)
Notes: • Secondary emission
• Good for low energy beams (no high energy secondary)
• Small signal
• If the wire becomes too hot it can start emitting thermionic electrons
spoiling the measurement
• High energy secondary
• No problem with wire heating (well...)
• Strong signal
• Detection may be non homogeneous leading to distorted profiles

## Slide 21: SEM Grid
Similar to wire scanner
The SE current from each wire or strip is acquired independently
Wire spacing down to a few hundreds' microns
Single shot measurement

## Slide 22: slit scanner
Similar to wire scanner
Beam pass through a slit and then the intensity is measured by faraday cup or scintillator.
Slit could be vertical or horizontal
Useful for low energy and low intensity (I>> ISEM)
Notes: Pass a thin slit across the
beam and detect the beam
current passing trough the slit
• Similar to wire scanner
• Useful for low energy and low
intensity (I>> ISEM)
• Detector typically is a Faraday
cup, but could also be a
scintillator, silicon or diamond
detector.
38

## Slide 23: Scraper
Similar to wire scanner
But deal with lower intensity
Measured CDF instead of PDF
J. Hunts, et al., NIMA, Vol 896, Pages 139-151.

## Slide 24: Beam profile of high intensity beams
Energy deposition
  Signals are often proportional to the deposited energy
  Energy deposition can cause damage to the instrument
Ideally: Non-invasive.
Damaged screen
Damaged SEM Grid
Damaged OTR screen
Notes: Energy deposition is probably the most important aspect
for all intercepting devices
• The Bethe-Bloch formula describes energy losses for
many of our cases
• The energy lost by the particles is not necessarily
deposited in the sensor

## Slide 25: Synchrotron radiation
Synchrotron radiation is broad band.
Visible wavelength
  Optical component is easy to implement.
  diffraction limits the resolution at longer wavelengths when the beam is small
X-ray
  Better resolution
  Hard to form an image
    Pinhole camera
    Zone plates

## Slide 26: Laser wire
First ionization potential for H‐ ions is 0.75eV.
Photons with λ<1500nm can separate H‐ ion into free electron and neutral H.
The H‐ ion has no excited states so the electron is removed into the continuum.
Laser can be used to mark a portion of beam by neutralization.
Once a porton of the beam is marked, measurements can be made on the neutral beam, the removed electrons, or the reduced beam current with beam current transformer or BPM stripline.
T. Shea, Laser Diagnostics for H- beams, , Accelerator Beam diagnostics, USPAS 2009

## Slide 27: Residual gas based Ionization Profile Monitor
Credit: P. Fork, GSI

## Slide 28: Residual gas image of a electron beam
E-Beam direction
Give you a one-dimensional profile!
Secondary ions or electrons are collected.
Magnetic field are used to compensate the profile distortion by self space charge force or wake fields.
Secondary particles are collected by electronics or optical (scintillator plus MCP)

## Slide 29: Beam induced fluorescence (BIF)
Schematic drawing of the BIF-monitor as installed at GSI UNILAC. *
BIF detection principle - How the beam induced fluorescence light is imaged, intensified and detected.*
*Credit: GSI, http://www-bd.gsi.de/dokuwiki/doku.php

## Slide 30: IPM and BIF
The cross section for impact ionization and fluorescence dependent on the beam energy, residual gas species, and certain fluorescent wavelength.
BIF has longer integration time due to lower cross section and viewed at certain solid angle.
Might need gas injection using a leaking valve to increase the local pressure.
Normally N2 has higher cross section, other gases such as Ne, Ar has been used.
Using Supersonic gas jet！

## Slide 31: Motivation to use gas jet
Low temperature
High directional jet
Non-invasive or minimum-invasive
Not dependent on beam species, energy and intensity
No damage threshold
Not affect the vacuum very much

## Slide 32: Gas jet profile monitor in Cockcroft institute

## Slide 33: Gas jet beam profile monitor
Estimated jet property
Parameter | Value
Density | 2.5*1016 particle/m3
Thickness | 0.5 mm
Vertical size | 5 mm
Estimated integration time = 1 ms
MCP
Phosphor screen
Ionization
Ions extraction
Visible light

## Slide 34: Typical results at CI
size | Value
Xrms | 0.42
Yrms | 1.23
Xrms from residual | 1.01
Setting | Value
Energy | 3.75 keV
Current | ~5.0 mA
External field | 7.5 kV/m
Exposure | 70 ms

## Slide 35: Using gas jet to measure LHC proton beam
HIPACE 300L/S
Turbo-molecular pump
HIPACE 700L/S
Turbo-molecular pump
Nozzle Chamber
Skimmer Assembly
Gas Separator
Skimmer Chamber
Imaging System
Interaction Chamber
Moveable Gauge Chamber
Dump Chamber
Third Skimmer

## Slide 36: Emittance measurement

## Slide 37: Multi-slit measurement
(X0, X0’)
L
Slit size = d
Beamlet rms size =

## Slide 38: Pepper pot
Notes: Measure angular spread of the beam and beam profile at the same time.
One option: ‘pepperpot’ and screen.
Blockthe beam apart from a few small holes.
Reduces space charge dominated beam to emittance-dominated beamletsthat drift to screen.

## Slide 39: 3 profile measurement
Transfer matrix
L1
L2

## Slide 40: Quadrupole scan
Thin lens approximation
L
f

## Slide 41: Tomographic Phase Space Reconstruction
y
y’
x
RC6
RC3
RC9
Lens
Screen
Projected / Slice, using Quadrupoles / Solenoids with Space Charge
Experiment
D. Stratakis, R.A. Kishek, S. Bernal, et al., Physics of Plasmas 17, 056701 (2010).
Notes: space charge
solenoids
time-resolved
emittance growth and halo

## Slide 42: Phase space tomography to study halo
x
x’
5 mm
20 mrad
x
y
5 mm
20 mrad
5 mm
20 mrad
x
y
5 mm
20 mrad
Y’
y
0
0.1
0.2
0.3
0.4
0.5
0.6
0.7
0.8
0.9
1
Matched
Mismatched
x
x’
H. Zhang, PhD thesis, university of Maryland, 2014

## Slide 43: Transverse emittance measurement methods
Technique | Pros | Cons
Basic Optics | Simple to implement. Linear space charge OK. Small or large beams. | RMS emittance only. No phase-space.
Pepper-pot | Simple to implement. Space charge OK. No beam optics. | Coarse phase-space. Impractical for small beams.
Quad Scan | Different schemes. Easy computation. | Space charge limits. No phase-space. Linear optics assumed.
Slit-Wire | Phase space info. Integrated or t-resolved. Space charge OK. | S/N problems. Hard to implement. Comp. intensive.
Tomography | Detailed phase space info. Linear optics/sp. charge. | More beam manipulation required. Computationally intensive.

## Slide 44
Beam energy Monitoring

## Slide 45: Using magnets

## Slide 46: Other methods
Faraday cup
Segmented Dump

## Slide 47: Beam halo Monitoring

## Slide 48: Halo Formation and Negative Effects
Beam Halo has many negative effects
  Nuclear Activation of The Transport Channel
  Emittance Growth
  Emission of Secondary Electrons
  Increasing Noise in The Detectors
Causes of Beam Halo
Nonlinear forces due to space charge
Parametric resonances
Scattered Laser Light in Photo-cathodes:
Misalignment, Magnet error, noise
Notes: Beam halo is group of particles outside the beam core.
Beam halo can be caused by parametric resonances or nonlinear forces due to space charge [2], as well as field emission, light scattered on photo-cathodes, or by numerous other phenomena
It brings many negative effects such as emittance growth, emission of sencondary electrons and so on. Because of this negative effects, we want to avoid halo by study the phenomenon.
Experimentally, the challenge to measure faint halos is we need a method with high dynamic range about 10_5 or 10_6. people using image system to detect the halo using coronagraphy method, the mask usually is not adaptive when the beam core shape changes.
Today, we will introduce an adaptive mask method to measure halo using a device call DMD().

## Slide 49: Optics System and Image process
180 Frames
32 mm
900 Frames
32 mm
C.P. Welsch et al.,
Proc. SPIE (2007)
J. Egberts, et al.,
JINST 5P04010 (2010)
H. Zhang, et al.,
Phys. Rev. STAB 15 (2012)
B. Lomberg, et al.
Proc. IBIC (2013)
120

## Slide 50: Demonstration of Adaptive Masking on UMER
70
280
x
y
(a)
(b)
IQ
640
660
250
45
45
60
82.9%IQ
66.3%IQ
49.7%IQ
Quadrupole Current
32mm
Different Beam / Different Mask Shape
6mA beam

## Slide 51: Conclusion
Without beam diagnostics
  it is not possible to even operate a accelerator!!
Non-invasive, single shot, real-time measurement is demanded.
New concept adapted from other field is hugely popular
  Example: laser, non-linear optics, advanced algorithm, machine learning,…
User-based development specifically target to project
  OMA, AWA, EuPRAXIA, EuroCirCol, AWAKE, HLLHC, etc. projects
Also: Good prospects for commercialization (D:Beam).

## Slide 52: Any Question?

## Slide 53: Reference
Workshops, conference and schools
  BIW, IBIC
  DITANET, OPAC, USPAS, JUAS, JAS, CAS
J. Harasimowicz et al., Rev. Sci. Instr. 81 (10), 2010
H. Koziol, BEAM DIAGNOSTICS FOR ACCELERATORS
P. Forck, Lecture Notes on Beam Instrumentation and Diagnostics, Joint University Accelerator School, January – March 2012
J. Hunts, et al., NIMA, Vol 896, Pages 139-151.
U. Raich, Accelerator Beam Diagnostics, USPAS, Albuquerque NM, June 23-26, 2009
D. Stratakis, R.A. Kishek, S. Bernal, et al., Physics of Plasmas 17, 056701 (2010)
H. Zhang, PhD thesis, university of Maryland, 2014
E. Bravin, Cern Accelerator School on Beam Instrumentation 2-15 June 2018, Tuusula, Finland
C. Welsch, Beam Diagnostics, CIPAS, 2015
J. Harasimowicz, et al., PRSTAB, 15, 122801 (2012)
S. Bernal, Emittance and Emittance Measurements, USPAS08, Beam Dynamics Experiments on the University of Maryland Electron Ring