Performance of the FASTPIX sub-nanosecond CMOS pixel sensor demonstrator
Eric Buschmann, Justus Braach, Dominik Dannheim, Katharina Dort, Thanushan Kugathasan, Magdalena Munker, Walter Snoeys, Mateus Vicente
09.09.2021
This project has received funding from the ATTRACT project funded by the EC under Grant Agreement 777222.

---
FASTPIX
- Technology demonstrator for monolithic pixel detector with sub nanosecond timing in 180nm process
- Goal: combine advantages of monolithic CMOS with small capacitance and high field
- 32 matrices of hexagonal pixels:
  - 4 pixel pitches: 8.66, 10, 15, 20 μm
  - 8 pixel designs: combination of process and layout variations
- In each matrix: 4 analog channels, 64 digital channels
- Digital matrix: 1 LVDS fast OR of pixel signals, 2 LVDS ToT / pixel encoding via delay lines
- 1 LVDS for test-pulse injection
T. Kugathasan et al., 10.1016/j.nima.2020.164461

---
Sensor Optimization
- Small collection electrode minimises capacitance and improves signal to noise, but leads to small sensor junction → lower and inhomogeneous electric field, limiting timing precision and radiation tolerance
- Sensor design optimization before FASTPIX
  - Increase and shape field significantly while maintaining small sensor capacitance
- Sensor design optimization for FASTPIX
  - ‘fingers‘ in p-well opening to resolve trade-off between fast charge collection and capacitance
Sensor optimization
in FASTPIX

---
Sensor Optimization for FASTPIX
Comparison hexagonal to square pixel cell charge vs. time for particle incident at pixel corner:
- Hexagonal design reduces the number of neighbors and charge sharing → higher efficiency
- Hexagonal design minimizes the edge regions while maintaining area for circuitry → faster charge collection
- Other optimised parameters include opening of p-wells, size of collection electrode, deep n-implant… to optimise charge collection and capacitance
Simulated hexagonal unit cell –
electrostatic potential:
→see M. Munker at iWoRiD 2021 for details on sensor optimisation
Optimisations important not only for timing, but also for efficiency and radiation tolerance

---
Measurements
- Internal pulsing → test of front-end response and calibration
  - One analog and two digital pixels per matrix can be pulsed
  - Pulse height is set via external voltage
- Measurement with Sr90 source / frontend + sensor test
  - Four analog pixels per matrix
  - Histogram of rise time measured with analog pixels → Estimate of charge collection speed without external trigger
- Test-beam measurements → response to minimum ionising particles

---
Measurement Setup
Lab setup with Caribou DAQ
- Caribou readout system:
  - FASTpix on chip board, connected to CaR board for powering, configuration and test pulsing
- Readout with oscilloscope, controlled by Peary software (Zynq SoC) → automated scans
- Oscilloscope connected to analog or digital outputs
- Offline analysis of raw waveforms
CaR
board
Chip board
Zynq board
FASTPIX on test board
5.3mm
Oscilloscope:
DSO9254A
2.5GHz / 10Gs/s
4.1mm

---
Analog / Digital Readout
Analog output
Pulser input
Digital output
- 4 analog pixel outputs → waveform analysis (rise time, pulse height)
- Three digital outputs from matrix:
  - Combined discriminator outputs via 2 delay lines and direct OR of all channels

---
Digital Readout
2 pixel event
3 pixel event
- Pixel discriminators are combined in the Fast OR output and connected to 2 delay lines running in opposite directions
- TOA is encoded in the edges of Fast OR output
  - Asynchronous readout allows for high time resolution
- ToT is encoded in the distance between peaks, which is used to match pixel signals on the 3 channels
- Decoding is limited by overlap between peaks

---
Digital Readout
Pixel positions from test-beam data.
64 peaks with a gap of 8 stages in the middle of the delay line
- Pixel position on delay line is calculated from Out 1&2
- Position on delay line is then mapped to position in pixel matrix

---
Calibration
- Calibration of threshold with 2 pulseable pixels per matrix
- Expect ~20% uncertainty on measured values due to uncertainty on injection capacitance
- 43e threshold and 5e noise for 10µm matrix on W18 at Ithr=1V
Plan to perform absolute calibration with X-rays or other sources
Wafer 18
-6V/-6V pwell/backside
10x10 µm² matrix

---
Source Measurements
- ⁹⁰Sr source placed above chip
- Readout of single pixel
- Waveforms are recorded and analysed offline
- Pulse heights and rise times are calculated
Small amplitude
Large amplitude

---
Test-beam Measurements - Setup
- Test-beam at SPS in August and September in Timepix3 telescope
- Micro-channel plate as fast time reference (expected resolution below 10ps, see: J. Bortfeldt et al.)
- New improved chip board for FASTPIX
CaR board
MCP
Tpx3 planes
Chip board
Tpx3 planes
Hitmap of TPX3 plane showing events with associated FASTPIX triggers (55x55um² pixels on TPX3)
- Readout of FASTPIX and MCP with oscilloscope
- Synchronisation of oscilloscope triggers to telescope DAQ
- Test-beam still ongoing: only preliminary results

---
Test-beam Measurements
Significantly more hits on edge of matrix:
Hit map, largest ToT pixel in event
Dummy pixels
Pixels connected to readout
Wafer 18
-6V/-6V pwell/backside
20x20 µm² matrix
70e threshold
- Pixels on edge see additional signal from outside of the matrix
- Number of hits depends on the number of neighbouring dummy pixels

---
Test-beam Measurements
Also visible in ToT spectrum:
Second peak at low ToT from
edge effects
ToT spectrum, inner pixels
ToT spectrum, outer pixels
Wafer 18
-6V/-6V pwell/backside
20x20 µm² matrix
70e threshold
outer pixel hits are partially from events with charge lost outside the matrix
Treat inner and outer pixels separately to account for edge effects

---
Test-beam Measurements
- Timing of largest ToT pixel vs ToT
- Combined (top) and for inner region only(bottom)
- ‘slow‘ tail at low ToT disappears in inner region
Wafer 18
-6V/-6V pwell/backside
20x20 µm² matrix
70e threshold

---
Test-beam Measurements
Wafer 18
-6V/-6V pwell/backside
20x20 µm² matrix (left)
70e threshold
10x10 µm² matrix (right)
50e threshold
- Timing after timewalk correction on 20µm (left) and 10µm (right) matrix
- Pixel-by-pixel correction for best results, reaching below 200ps resolution
- 10µm matrix is operated at lower threshold with a few Hz noise rate
- Larger cluster sizes for 10µm leads to lower average seed charge and thereby more time walk and makes decoding more difficult

---
Test-beam Measurements
Comparison of mean cluster size for different pixel pitches
Wafer 18
-6V/-6V pwell/backside
Calculated cluster size does not account for edge effects →biased to smaller values

---
Summary and Outlook
- FASTPIX combines a small collection electrode with fast charge collection in a monolithic design
- Tests show that analog and digital readout are operational
  - Performed source measurements and threshold calibration
- Preliminary test-beam results show MIP time resolution of approximately 120-130 ps
  - very promising timing performance
  - Analysis is work in progress: use tracking information for efficiency, in-pixel, and position resolution measurements

---
Backup

---
Test-pulse measurements: Jitter
Pulsing of digital pixel
Pulser input
Digital output
- Measured jitter between pulser input and digital output: <50ps
- Measurement limited by rise time of signal generator
- Timing not limited by pixel frontend

---
Source Measurements
Spectrum of ⁹⁰Sr source
Rise time 10-30% of peak in pulse height spectrum
- Comparison of wafer 3 (standard process) and wafer 18 (optimized for fast charge collection)
- Visible difference in rise times for standard and optimized process

---
Test-beam Measurements
Wafer 18
-6V/-6V pwell/backside
10x10 µm² matrix

---
Test-beam Measurements
Wafer 18
-6V/-6V pwell/backside
20x20 µm² matrix

---
Test-beam Measurements - Setup
- Parasitic running during June test-beam at DESY
- Setup with FASTPIX and scintillator as coincidence/time reference connected to oscilloscope (MSOS404A, 4GHz, 10Gs/s)
- 2.2GeV beam, ~10KHz beam rate
- All measurements with one sample (W18-2)
Scintillators
CaR board
Chip board
Zynq board

---
Test-beam Measurements
850mV
- Using second scintillator and DESY clock to estimate contribution of scintillators to time resolution
- About 470ps RMS between all references → expect resolution of about 470ps/√2 = 350ps for single scintillator
PMT voltage	PMT1-clk (ns)	PMT2-clk (ns)	PMT1-PMT2 (ns)
800mV	0.472	0.468	0.471
850mV	0.460	0.461	0.465
900mV	0.463	0.460	0.468