## Slide 1
Silicon crystals for steering of high-intensity particle beams at ultra-high energy accelerators
A. Mazzolari, M. Romagnoni
10/03/2020

## Slide 2: Crystals technology at INFN
Collimation
Beam steering
Innovative radiation sources
Pair production studies
Innovative detectors
Beam steering
Innovative radiation sources
Pair production studies
Beam steering
Innovative radiation sources
Pair production studies
CRYSBEAM
SELDOM

## Slide 3: INFN Competencies
INFN has competencies related to
  Development of innovative ideas and research activities (connected to channeling and related effects)
  Crystals manufacturing and characterization
  Holders manufacturing and characterization
  Goniometry
  Trackers and detectors
  Data analysis
  Design of setups for channeling experiments
INFN is a scientific agency encompassing a full set of competencies related to studies of coherent effects between charged particle beams and crystals.

## Slide 4: INFN infrastructure
Laboratory fully equipped for silicon micro and nanomachining
ISO4 certified clean room (130 m2)
High-resolution x-ray diffraction
Dicing and polishing equipment
White light and Fizeau inteferferometers

## Slide 5: INFN infrastructure
Laboratory fully equipped for silicon micro and nanomachining
Fotolitography
Equipment for silicon chemical etching
Nanoimpringing

## Slide 6: Strip crystals for the LHC
4 mm
A primary bending is mechanically imposed along the 50 mm size.
As a result, «anticlastic bending» manifests along the 4 mm size.
50
2

## Slide 7: Choice of crystalline planes
(110)
(111)
Comparable depth of potential well
(110) Planes offers lower nuclear interaction rate
See also W. Scandale et al., Eur. Phys. J. C (2018) 78:505

## Slide 8
2006 Manufacturing based on mechanical dicing followed by chemical etching
2007 Manufacturing based on purely chemical methods
2014 Manufacturing of low-miscut crystals (<10 urad)
2018 Manufacturing based on mechanical dicing followed by chemical etching and polishing of surface entry face Manufacturing of ultra low-miscut crystals (<1 urad)
Brief recap about crystals manufacturing

## Slide 9
Crystal manufacturing – AFM and HRTEM –
Lateral surface (AFM)
Sub-nm roughness was achieved
Entry surface (High Resolution transmission electron microscopy).
Zero nm amorphous layer
High-quality surfaces achieved via anisotropic chemical etching
Sub-nm roughness achieved (0.2 nm)

## Slide 10
2006 First holders version (in collaboration with IHEP), made of Al7075
2007 FEM-assisted optimization of holder geometry
2008 First prototypes of pre-shaped holders (aluminum)
2013 Request for holders with the following specs:
weight<90 g
UHV compatible (bakeable)
electron cloud compatible
2014 First generation of titanium grade 5 holdersnot compatible with bake-out
2018 Pre-shaped titanium grade 5 holders now fits LHC requirements
Titanium.
About 90 commercial alloys, most common is «titanium grade 5» (90% Ti 6% Al 4% V )
Brief recap about holders manufacturing

## Slide 11: Crystals for the LHC
Thickness along the beam: 4.0±0.1 mm
Height < 55 mm.
Weight < 150 g
Channeling plane: (110)
Channeling axis: <111> or <110>
Miscut for axial channeling: 0±18 mrad
Miscut for planar channeling: < 10 μrad.
Torsion: < 1 μrad/mm
Bending angle: 47.5-52.5 μrad
Dislocation density < 1 cm2
Bake out
• Bake out temperature= 250°C
• heating ramp= 50°C/h
• Bake out time =48h00
• Number of thermal cycles: 3 at least
• Maximum allowed total outgassing after each bake out: 1*10-7 mbar·l·s-1
The compliance of the mechanical properties of the crystal assemblies to the specifications of Table 1, in particular the bending angle, must be assessed after at least 3 thermal cycles.

## Slide 12: Challenging aspects
Planar miscut: < 10 μrad MRF finishing
Torsion: < 1 μrad/mm  ultra precise machining and assembly.
Bending angle: 47.5-52.5 μrad  ultra precise machining and assembly.
Dislocation density < 1 cm2  standard for microelectronics is 1 order of magnitude higher purchase of large quantity of silicon wafers and selection of the best ones.

## Slide 13: MRF finishing
150mm MRF Wheel w/ ribbon of MR fluid
6” vacuum wafer chuck
4” Si Wafer
5-axis CNC machine moves wafer
Technique developed for astronomy applications (glasses and ceramics, i.e. amorphous materials)
After a 2 years R&D the tecnique was adapted to operate on crystalline materials.

## Slide 14
Starting surface
PV 2.18 um
RMS 0.39 um
Surface after first treatment
PV 0.615 um
RMS 0.14 um
Final surface
PV 0.177 um
RMS 0.01 um
MRF is a deterministic polishing process
Does not induce any lattice damage
Miscut < 1 urad
MRF finishing

## Slide 15
X-ray source
Detector
Sample
Autocollimator
Mirror
Monochromator
High resolution x-ray diffractometer (PANALYTICAL)
Miscut: angle between optical surface and atomic planes
Requested miscut < 10 μad.
MISCUT REDUCTION

## Slide 16: Recently developed crystals
Developed «pre-shaped» holders made of titanium grade 2 and grade 5.
High-resolution x-ray diffraction characterizations: analysis of correlation
Investigated thermal stability of holders+crystal assembly.
Crystals for the LHC

## Slide 17: Pre-shaped holders
Ultra-high precision machining provides holders with crystal-supporting surfaces inclined at a proper angle to impart to the crystal the desired deformation.
  At first step, the holders are manufactured with approaches typical of conventional precision mechanics.
  selected holders are treated with super-finishing techniques to adjust surface inclinations.
Approach already developed for crystals installed in the SPS and for studies of multiple-volume reflections (2010÷2015)
Design is now revisited:
  holders made of titanium grade 2 and grade 5 (previous version was made of aluminium).
  removed torsion-adjustment mechanism.
A pre-shaped holder for LHC. Holder is made of titanium. Flexure for torsion adjustment have been removed.
A pre-shaped holder manufactured in 2011 for UA9. Holder was made from aluminium and torsion correction mechanism still used.
Base for development of current holders.

## Slide 18
Crystal | Deflection angle (µrad) |  | 
 | High resolution x-ray diffraction | Channeling | Consistency
STF47 | 33±2 | 35±2 | YES
STF48 | 144±2 | 142±2 | YES
STF49 | 247±3 | 246±2 | YES
STF50 | 142±5 | 143±2 | YES
STF51 | 33±2 | 33±2 | YES
STF70 | 56±2 | 55±2 | YES
STF71 | 60±5 | 62±2 | YES
STF99 | 119±3 | 120±2 | YES
STF100 | 67±6 | 63±2 | YES
STF101 | 170±6 | 165±2 | YES
STF102 | 45±3 | 42±2 | YES
STF103 | 52±5 | 54±2 | YES
STF104 | 95±5 | 91±3 | YES
STF105 | 49±3 | 50±2 | YES
STF106 | 42±2 | 42±2 | YES
STF107 | 56±2 | 56±2 | YES
STF110 | 52±3 | 54±2 | YES
STF110 | 56±10 | 62±2 | YES
STF112 | 64±3 | 63±2 | YES
STF113 | 46±3 | 45±1 | YES
STF114 | 52±3 | 52±1 | YES
STF117 | 53±3 | 50±1 | YES
STF118 | 52±3 | 53±1 | YES
STF119 | 54±3 | 52±1 | YES
STF120 | 54±3 | 52±1 | YES
STF121 | 48±3 | 48±1 | YES
STF122 | 50±3 | 46±1 | YES
SFT123 | 52±3 | 52±1 | YES
PL08 | 715±30 | 706 ± 2 | YES
PL09 | 1040±50 | 1067 ± 2 | YES
Determinantion of crystal bending angle with high-resolution x-ray diffraction and channeling of high-energy particles are in good agreement.

## Slide 19: Holders Thermal stability
Crystal bending angle is measured through high-resolution x-ray diffraction before and after thermal cycles for 41 holder+crystals assemblies.

## Slide 20
Chart: Crystal #1
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 9 | 12 | 14 | 18
 | 58 | 56 | 58 | 59 | 57 | 58 | 59 | 56 | 59 | 58 | 60
Chart: Crystal #2
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 9 | 12 | 14 | 18
 | 37 | 35 | 36 | 36 | 38 | 38 | 37 | 36 | 37 | 35 | 37
Chart: Crystal #3
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 9 | 12 | 14 | 18
 | 47 | 48 | 45 | 46 | 47 | 47 | 48 | 47 | 46 | 47 | 48
Chart: Crystal #4
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 9 | 12
 | 46 | 45 | 46 | 46 | 45 | 47 | 46 | 46 | 45

## Slide 21: Holder stability vs bake-out: conclusions
Manufactured 41 holders (31 titanium grade 2+10 titanium grade 5).
41 holders subjected to bake-out thermal cycles. Performed a total of 243 thermal cycles.
For all the cases crystal bending angle is stable.
Newly developed holders (both titanium grade 5 and titanium grade 2) are thermally stable with respect to bake-out cycles.
Chart: Crystal #1
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 9 | 12 | 14 | 18
 | 58 | 56 | 58 | 59 | 57 | 58 | 59 | 56 | 59 | 58 | 60

## Slide 22: CRYSTALS FOR THE LHC
Chart: Crystal STF 117
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 6 | 7
 | 54 | 55 | 52 | 53 | 54 | 53 | 53 | 50 | 51
0 1 2 3 4 5 6 6 7 | 54 | 55 | 52 | 53 | 54 | 53 | 53 | 50 | 51
Chart: Crystal STF 118
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 6 | 7
 | 50 | 54 | 50 | 53 | 52 | 54 | 51 | 53 | 53
0 1 2 3 4 5 6 6 7 | 50 | 54 | 50 | 53 | 52 | 54 | 51 | 53 | 53
Chart: Crystal STF 119
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 6 | 7
 | 50 | 49 | 50 | 47 | 47 | 46 | 48 | 48 | 48
0 1 2 3 4 5 6 6 7 | 54 | 55 | 52 | 53 | 54 | 53 | 53 | 53 | 53
Chart: Crystal STF 120
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 6
 | 53 | 55 | 53 | 53 | 53 | 55 | 56 | 52
0 1 2 3 4 5 6 6 | 53 | 55 | 53 | 53 | 53 | 55 | 56 | 52

## Slide 23: CRYSTALS FOR THE LHC
Chart: Crystal STF 121
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 6 | 7
 | 50 | 49 | 50 | 47 | 47 | 46 | 48 | 48 | 48
0 1 2 3 4 5 6 6 7 | 50 | 49 | 50 | 47 | 47 | 46 | 48 | 48 | 47
Chart: Crystal STF 122
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 6 | 7
 | 50 | 49 | 50 | 47 | 47 | 46 | 48 | 48 | 48
0 1 2 3 4 5 6 6 7 | 53 | 49 | 48 | 53 | 52 | 49 | 49 | 46 | 46
Chart: Crystal STF 123
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 6 | 7
 | 50 | 49 | 50 | 47 | 47 | 46 | 48 | 48 | 48
0 1 2 3 4 5 6 6 7 | 54 | 51 | 52 | 51 | 54 | 53 | 52 | 52 | 51
Chart: Crystal STF 124
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 8
 | 50 | 48 | 48 | 49 | 50 | 50 | 50 | 50 | 49 | 49
0 1 2 3 4 5 6 7 8 8 | 50 | 48 | 48 | 49 | 50 | 50 | 50 | 50 | 49 | 49

## Slide 24: CRYSTALS FOR THE LHC
Chart: Crystal STF 125
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 8
 | 50 | 52 | 50 | 49 | 51 | 53 | 52 | 52 | 51 | 50
0 1 2 3 4 5 6 7 8 8 | 50 | 52 | 50 | 49 | 51 | 53 | 52 | 52 | 51 | 50
Chart: Crystal STF 126
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 6
 | 51 | 51 | 50 | 48 | 50 | 52 | 51 | 50
0 1 2 3 4 5 6 6 | 51 | 51 | 50 | 48 | 50 | 52 | 51 | 50
Crystal code | Bending angle | Miscut | Bakeable
STF117 | 51±1 | 6±1 | YES
STF118 | 53±1 | 6±1 | YES
STF119 | 53±1 | 6±1 | YES
STF120 | 52±1 | 6±1 | YES
STF121 | 47±1 | 6±1 | YES
STF122 | 46±1 | 6±1 | YES
STF123 | 51±1 | 6±1 | YES
STF124 | 49±1 | 6±1 | YES
STF125 | 50±1 | 6±1 | YES
STF126 | 50±1 | 6±1 | YES

## Slide 25: CRYSBEAM EU project
Motivation: extraction of the LHC circulating beam, delivering extracted particle beams to users
Need of a crystal bent at ~1 mrad, bending radius ~80 m
Deflection ~0.3 mrad

## Slide 26: SELDOM EU project
P.I. Nicola Neri

## Slide 27: Germanium crystals
We own the technology to machine also germanium crystals
Mechanical machining
Chemical etching for lattice damage removal

## Slide 28: Most recent developmentS
Crystals with miscut less than 1 urad
Assessed a correlation between x-ray and particle measurements
Thanks to the experience already gained in development of holders for the SPS, pre-shaped holders made of titanium (grade 2/grade 5) are now available
Holders/crystals assembly are now bakeable to 250 °C
Tight requirement on bending angle can be routinely satisfied
Recent progresses on torsion reduction

## Slide 29: COnclusions
INFN owns the expertise, technical resourses, manpower to perform channeling experiments at any wordwide facility.
Established a protocol for manufacturing of crystals free from crystalline defects.
Possibility to deliver crystals with miscut lower than 1 urad.
Developed a protocol to manufacture crystals matching requirements for installation in the LHC.
Possibility to realize germanium crystals
Possibility to realize cm-thick uniformly bent crystals