## Slide 2: BLM Software Changes during LS2
David Medina Godoy
BE-BI-SW

## Slide 3: Outline
Projects I Am Involved With
Agile Approach
BLMINJ System Overview
BLMINJ LS2 Changes
Challenges of Software Development
Initial Results & Future Work
Conclusion

## Slide 4: Projects Where I Am Involved
BLM Projects
  LHC
  Injectors (except SPS)
Activities:
  FESA
  Graphical User Interfaces (JavaFX/Swing)
  Ad-hoc assistance to develop and diagnose electronics and installation
Awake

## Slide 5: An Agile Approach
Close collaboration BL-SW
Weekly follow-up meetings
Sprints
Pair programming
Code reviews
Mockups
Unit-tests

## Slide 6: BLMINJ System Overview
Acquisition Electronics
Processing Electronics
RT Software (FESA)
Driver (ENCORE/EDGE)
Expert App
OP App
CALS/NXCALS
D1
D2
D3
D4
D5
D6
D7
D8
Tunnel
Surface
Measures lost particles from the trajectory
Machine protection (HW Thresholds)
Machine optimization (SW Thresholds)
Software
Hardware

## Slide 7: Electronics Synchronisation
With Basic Period (BP) event:
  Received via MTG (Main Timing Generator)
  Broadcasted in backplane by a dedicated timing card
Done at the processing level (i.e. 2 samples jitter between cards).
Each BP:
  Electronics:
    Continuously calculate the integration period values (2 s, 400 s, 1 ms)
      Hardware thresholds exceeded -> interlock
      Acquisition frequency is 2 s
      Implemented with moving sum windows.
  CPU:
    Reads-out and publishes data

## Slide 8: System Functionalities per USER
Uses data accumulated over 1..N BP
  3 Running Sums (2 s, 400 s, 1 ms)
  Loss Integrals over Beam Presence & Cycle
  Loss Evolution over the cycle
    μs or ms sampling
  Loss Capture function
    On-demand, with adjustable start of recording and sampling

## Slide 9: Loss Evolution Function
Evolution of the losses over the entire basic period
  1200 samples
  1ms/sample → PSB, PS
  2 μs/sample → LN4, TT10, TT2

## Slide 10: Loss Capture Function
  512 samples
  Settings:
    Start of recording adjustable relative to basic period
    Configurable sampling period (2 μs, 400 μs, 1ms)
  To be used mainly in PSB and PS

## Slide 11: Sum of Losses over Beam Presence
  Timing signal connected via a cable to electronics for Beam In/Out

## Slide 12: Hardware Changes
Motivation:
  Generic processing board for most BI instruments
  Allow the use of common libraries with the other developments
  Simplify transfer of code to newer electronics
Old System

## Slide 13: Hardware Changes
To be replaced with Remote Control Unit
Replaced by VFC-HD
Old System
Motivation:
  Generic processing board for most BI instruments
  Allow the use of common libraries with the other developments
  Simplify transfer of code to newer electronics

## Slide 14: Hardware Changes
New System
Motivation:
  Generic processing board for most BI instruments
  Allow the use of common libraries with the other developments
  Simplify transfer of code to newer electronics

## Slide 15: Keep BLMINJ Homogeneous
BLMINJ
Installation | PSB | PS | TT2 | TT10 | LN4
Type | ring | ring | linear | linear | linear
Cycle | 1 BP | 1..5 BP | 1..20 BP | 1..20 BP | 1 BP
Interlock | BIS | Tail clipper event sent to scraper by BLMSYNC | BIS | BIS | BIS
OP requirements |  | Concatenates all losses from 2 crates |  | Cycle summary acquisition Acquire around injections | 
One FESA class to rule them all

## Slide 16: LS2 Software Changes
Summary Acquisition (TT10)
BLMSYNC (PS)
Software Interlocks
Changes in register map

## Slide 17: Summary Acquisition for TT10
New feature publishes several injections
Validated with mock-up in the lab
Gy/s
Injection 1
Injection 2
Injection 3
Simulation Data

## Slide 18: PS Concentrator (BLMSYNC)
Publish together data acquired from all monitors coming from 2 VME crates
Need to synchronize acquired data
  BLMSYNC: generic solution to concentrate data
Interlocking capabilities
  Protection becomes an issue due to higher brightness after LIU (LHC Injectors Upgrade)
  Calculate total ring losses with sum of all monitors from both crates
  No interlock system in the PS
    Beam inhibit by tail clipper event sent to scraper

## Slide 19: Software Interlocks
Limit radiation levels with software implementation comparisons.
The threshold values are set per user and per monitor
  If losses > threshold → block user’s injections
The blocked beam permit signal will be latched until an operator acknowledges.
Synchronization with current user at the start of cycle.
Possibility to mask software interlock per user and monitor
Added watchdog functionality → Interlock if the software is not active

## Slide 20: Software Interlocks
Overview of losses vs thresholds displayed in Expert Application
Highlights in red when a threshold is breached
Display in yellow when a mask is active

## Slide 21: LS2 Software Changes
Summary Acquisition (TT10)
BLMSYNC (PS)
Software Interlocks
Changes in register map

## Slide 22: LS2 Software Changes
Summary Acquisition (TT10)
BLMSYNC (PS)
Software Interlocks
Changes in register map
Assist with firmware validation
Small change at register map level can propagate to many changes in upper levels (encore->FESA->Expert App.)
“Small” on-the-fly requests
Working with legacy code
(re) Commissioning the system

## Slide 23: Unit Tests: Improving Quality
Validation of software interlocks
  Systematic setting of thresholds and testing system behavior
Use of simulation data to assert correct data handling
  E.g. New TT10 injection API
  Systematic Unit test
  Visual test in the navigator
Reproducing system anomalies using Python script
  To assist hardware team to diagnose
  And to avoid the anomaly returning in the future!

## Slide 24: Initial Results
Revised BLMINJ system deployed during 2019 LINAC4 run
Confident in the system
Software now matches OP requirements
  After last-minute API changes for masking, interlock behavior,…
Workaround in software for data-consistency problem
  BLM Team will continue to work on final fix during Q1 2020
  Some technical debt to catch up on
(re)Commissioned with beam with OP
New system used successfully!!

## Slide 25: Future Plans
Finish PS specific concentrator
Finalise strategy for TT10
2 New deployments in TT2, LN4
Many improvements for Expert Tools
Refocus on LHC system during 2020!

## Slide 26: Conclusion
New BLMINJ system validated in LN4
One BLMINJ system for all installations
SW experts provide assistance for system validation as well as the FESA class and Expert GUIs
Small changes in the driver level can propagate to many changes in upper layers
Agile methodologies are our allies when working with complex systems

## Slide 27: Questions
Thank you

## Slide 28: Backup Slides

## Slide 29: Hardware changes

## Slide 30: BLMINJ System Overview

## Slide 31: BLM Systems at CERN
LINAC 4
LHC
SPS
Injectors*
*Except SPS

## Slide 32: BLM systems in the Injectors (BLMInj)*
LINAC 4
*Except SPS
Notes: Put where the crates ate

## Slide 33: Integration Periods
Continuously the processing electronics will calculate 4 integration period values for each channel:
2 s, 400 s, 1 ms and 1.2 s (full cycle)
  implemented as moving sum windows in the hardware
  calculation refreshed at acquisition frequency
Compare with predefined thresholds
  Machine protection with hardware implementation comparisons on each refresh
  Limit radiation levels with software implementation comparisons at end of cycle
  See also next slide.
Calculate for each channel the maximum values recorded on each integration period during the cycle
  Publish them for the online displays and
  the long-term logging

## Slide 34: Software Thresholds
All maximum integration period values recorded on the cycle will be checked against a second set of threshold values. The outputs will be used for repeated over threshold function
  Additional threshold values for the same integration periods will also be required.
  In the case found to be over threshold repeatedly n times it will be required to block that user’s injections.
  The blocked beam permit signal will be latched until an operator acknowledges.
  The repeat value n will be settable per monitor in the range of 1 to 16.
The threshold values will need to be unique per user and per channel:
  Each CPU will process 8 cards x 8 channels
  The information of the current user has to be obtained from the telegram per cycle -> dedicated timing card
  Memory for 32 users will be reserved.

## Slide 35: FESA & FPGA Upgrades
Large upgrade program
  BLM INJ: processing module changed to VFC
    New memory map, driver, FESA and firmware
    New diagnostic/status data (consolidation of two firmware)
    New interlocking features (HW interlocks & expert mode for testing)
    Addition of concentrator with interlocking abilities (for PS mainly)
    Addition of persistent data at VFC level (EPROM)
  BLM LHC: FESA3 and new firmware
    New VMEbus core with MBLT mode
    New memory map optimized for block transfers
    Changes for UFO Buster
    Add Injection Interlock Inhibit functionality
Restructure of the FW code to industry standards
  Allow the use of common libraries with the other developments
  Simplify transfer of code to newer electronics