## Slide 1: The Timepix4 design parameters
  X. Llopart,
  On behalf of the Medipix4 Collaboration
11th February 2020

## Slide 2: Medipix4 Collaboration (from 2016)
CEA, Paris, France
CERN, Geneva, Switzerland,
DESY-Hamburg, Germany
Diamond Light Source, Oxfordshire, England, UK
IEAP, Czech Technical University, Prague, Czech Republic
JINR, Dubna, Russian Federation
NIKHEF, Amsterdam, The Netherlands
University of California, Berkeley, USA
University of Houston, USA
University of Maastricht, The Netherlands
University of Canterbury, New Zealand
University of Oxford, England, UK
University of Geneva, Switzerland
IFAE, Barcelona, Spain
University of Glasgow, UK
15 members
Timepix4: A 4-side tillable large single threshold particle tracking detector chip with improved energy and time resolution and with high-rate imaging capabilities

## Slide 3: Timepix4 applications
Data-Driven applications:
  HEP:
    Very high rate pixel telescope
    Sensor studies (high speed)
    Beam gas interaction (PS SPS/LHC?)
    AeGIS/ASACUSA
    ATLAS background rad monitor and TRD detector
    MOEDAL
    GEMPIX / large area TPC
    Test vehicle for next gen LHCb-VELO with tens of ps time resolution
  Time-of-flight mass spectrometry
  Neutron time-of-flight imaging
  Radiation monitors
  Electron microscopy
  X-ray and powder diffraction
  Compton camera for medical diagnostics
  Sub-pixel resolution imaging
  Gamma and neutron imaging for nuclear industry and Homeland Security
Frame-based imaging applications:
  X-ray imaging in synchrotrons with extreme high rates > 109 particles/mm2/s
20th Anniversary Symposium on Medipix and Timepix
  https://indico.cern.ch/event/782801/

## Slide 4: Timepix4 main requirements
Use of a commercial 65nm CMOS process
Build a 4-side buttable “as large as possible” pixel array
Particle identification and tracking (Data-driven and zero suppressed)
  Sub-ns time binning
  Improve the energy resolution
Imaging (frame based with CRW sequential readout)
  Increase particle count rate
Designers:
  CERN: R. Ballabriga, T. Poikela, E. Santin, V. Sriskaran, N. Egidos and X. Llopart
  Nikhef: V. Gromov and A. Vitkovskiy
  IFAE: R. Casanova

## Slide 5: Timepix3  Timepix4
 |  |  | Timepix3 (2013) | Timepix4 (2019)
Technology |  |  | 130nm – 8 metal | 65nm – 10 metal
Pixel Size |  |  | 55 x 55 µm | 55 x 55 µm
Pixel arrangement |  |  | 3-side buttable 256 x 256 | 4-side buttable 512 x 448
Sensitive area |  |  | 1.98 cm2 | 6.94 cm2
Readout Modes | Data driven (Tracking) | Mode | TOT and TOA | 
 |  | Event Packet | 48-bit | 64-bit
 |  | Max rate | 0.43x106 hits/mm2/s | 3.58x106 hits/mm2/s
 |  | Max Pix rate | 1.3 KHz/pixel | 10.8 KHz/pixel
 | Frame based (Imaging) | Mode | PC (10-bit) and iTOT (14-bit) | CRW: PC (8 or 16-bit)
 |  | Frame | Zero-suppressed (with pixel addr) | Full Frame (without pixel addr)
 |  | Max count rate | ~0.82 x 109 hits/mm2/s | ~5 x 109 hits/mm2/s
TOT energy resolution |  |  | < 2KeV | < 1Kev
TOA binning resolution |  |  | 1.56ns | 195ps
TOA dynamic range |  |  | 409.6 µs (14-bits @ 40MHz) | 1.6384 ms (16-bits @ 40MHz)
Readout bandwidth |  |  | ≤5.12Gb (8x SLVS@640 Mbps) | ≤163.84 Gbps (16x @10.24 Gbps)
Target global minimum threshold |  |  | <500 e- | <500 e-
3.5x
8x
5x
2x
33%
8x
4x
32x

## Slide 6: Timepix4 floorplan arrangement
512 x 448 of 55 x 55 µm pixels
3 “hidden” peripheries with TSV (Through-Silicon-Vias):
  TOP Edge: Data Readout & Slow Control
    147 TSVs/WBs: 106 POWER/GND and 41 IO
  BOTTOM Edge: Data Readout & Slow Control:
    147 TSVs/WBs: 106 POWER/GND and 41 IO
  CENTER: Slow Control and Analog Blocks (DACs, ADC, Band-Gaps…)
    147 TSVs: 124 POWER/GND and 23 IO
On-chip bump to pixel redistribution layer (RDL)
Chip size:
  With WB (wirebonds extenders): 29.96 mm x 24.7 mm
    >93.7% active area (28.16mm x 24.64mm)
  Without WB (TSV Only) : 28.22 mm x 24.7 mm
    >99.5% active area (28.16mm x 24.64mm)
Control architecture allows to operated Timepix4 from any of the 3 peripheries:
  i2C protocol
  Custom Slow Control protocol
  Interface to DAQ:
    Through 3xTSVs
    Through 2xWB
Fast readout requires at least 1 serial link enabled in each edge periphery:
  Serial links are highly configurable 40MBps  10GBps
24.7 mm (448 pixels)
28.22 mm (512 pixels)
29.96 mm (512 pixels)

## Slide 7: 4-side buttable pixel arrangement
Target to build large area detectors by combining smaller modules
The through-silicon vias (TSVs) is the key technology for this paradigm shift
Sensor
ASIC
FC
Wire Bonding
PCB
FC
BGA

## Slide 8: Expected Timepix4 Yield
Chart: die size [mm2]
Timepix3/Medipix3
Yield=73%
Chart: die size [mm2]
Timepix3/Medipix3
192/261 good dies
Timepix4
Yield=37%
Timepix4
30/73 good dies
*Based in the foundry yield model for perfect dies

## Slide 9: Timepix4 Pixel Schematic
Global threshold
Front-end
Leakage
Current
compensation
Preamp
Input
pad
1 pixel
VCO
@640MHz
Super pixel
Control
voltage
8 pixels
640MHz
TpA
TpB
TestBit
MaskBit
Front-end
Counters
&
Latches
Time stamp
16-bits
Synchronizer
&
Clock gating
OP Mode
5-bit Local
Threshold
~50mV/ke-
2.5fF
3fF
Sync
Readout
Data out
to EOC
8-bits
clock
(40MHz)
8-bits
4
8-bits
ADB
32 pixels
8
Clock
local
 | bits
ADDR | 9
ToA | 16
ufToA_start | 4
ufToA_stop | 4
fToA_rise | 5
fToA_fall | 5
ToT | 11
Pileup | 1
SPG

## Slide 10: Pixel Operation in TOA & TOT [DD]
65534
X
TOA (16 bits)=65534
Tpeak < 25ns
TOT (10 bits) =4
Preamp Out
Disc Out
ClkRef (40MHz)
VCO Clk (640MHz)
TOT Clk (40MHz)
TOA (16-bit)
65535
0
2
3
4
65534
65533
1
Global TOA (16-bit)
FTOA_RISE (4 bits)=7
FTOA_FALL (4 bits)=11
0
X
UFTOA_Start (4-bits)
8
X
UFTOA_Stop (4-bits)

## Slide 11: Timepix4 FE
The Timepix4 chip front-end has to be versatile in order to be used:
  With different types of sensors:
    Bipolar detection and <10nA leakage currents (high-Z materials)
  In tracking and as well in imaging:
    Imaging at high rates (e.g. synchrotrons) imposes a minimum pulse width. In order to be at the state of the art in terms of rate capability, the dead time of the front-end should be smaller than ~200ns.
    Tracking applications impose an input jitter in the front-end below 50 psrms (time bin 195 ps)
FE architecture follows the “classic” Medipix approach (Krummenacher FE)
  Gain adjustment (log gain in hole collection)
  Improved for low jitter and fast count-rate when required  power
Qin=2.2ke-
IKRUM=20nA
R. Ballabriga

## Slide 12: Timepix4 Slow control and Readout
Timepix4 has separated readout and control data paths
Slow control (SC):
  Protocols:
    I2C protocol
    Custom serial slow control protocol (< 40 MHz)
  Slow control operation can be done from any of the 3 available peripheries
  Registers:
    Pixel matrix configuration > 2 Mb (8-bit/pixel + 24-bit/spgroup)
    Peripheral registers (DAC settings, operation modes, e-fuses, ….)
Data readout (DAQ):
  Each half matrix readout through its correspondent edge periphery
    8 serializers in each periphery with < 81.92 Gbps
  Uses a 64/66b standard protocol
  Zero-suppressed and data driven readout  64-bit data packet/hit
    TOA/TOT mode
    24b data-driven Photon Counting (PC) mode
  Full frame with Continuous-Read-Write (CRW):
    8-bit or 16-bit Photon Counting (PC) mode

## Slide 13: Timepix4 Readout Modes : Data-Driven
Zero-suppressed continuous data-driven
  Output bandwidth from 40 Mbps (2.6 Hz/pixel) to 160 Gbps (10.8 KHz/pixel)
  Uses 64b/66b standard encoding communication protocol
Acquisition time
64bit
64bit
64bit
64bit
64bit
64bit
Shutter
Qin
DataOut
Address: 18 bits
Time: 29 bits
Energy: 21 bits

## Slide 14: Timepix4 Readout Modes : Frame-based
Full frame readout with continuous read-write (CRW):
  No zero-suppressed  no pixel address sent
  8-bits or 16-bit counter depth
  Uses 64b/66b standard encoding communication protocol
  Frame rate limited by bandwidth:
    16-bit @ 8Ghits/mm2/s > 1.28 Gbps  ~370 fps up to 44.6 Kfps @ 163.84 Gbps
    8-bit @ 1Ghits/mm2/s > 20.48 Gbps  ~11.8 Kfps up to 89.2 Kfps @ 163.84 Gbps
16-bit frame
8-bit frame

## Slide 15: On-pixel TDC simulations  TOA resolution
Chart: Qin [ke-]
| 0.5 | 0.6 | 0.8 | 1.2 | 2 | 3.6 | 6.8 | 13.2 | 26 | 51.6 | 102.8
TOA_tpix3 | 13.337 | 10.498 | 6.003 | 3.428 | 2.093 | 1.334 | 0.972 | 0.959 | 0.731 | 0.716 | 0.662
±σrms | 17.47 | 13.656 | 7.403 | 4.246 | 2.738 | 1.953 | 1.587 | 1.427 | 1.168 | 1.149 | 1.085
 | 9.204 | 7.34 | 4.603 | 2.61 | 1.448 | 0.715 | 0.357 | 0.491 | 0.294 | 0.283 | 0.239
Chart: Qin [ke-]
| 0.5 | 0.6 | 0.8 | 1.2 | 2 | 3.6 | 6.8 | 13.2 | 26 | 51.6 | 102.8
TOA_tpix4 | 12.671 | 9.796 | 5.471 | 2.966 | 1.561 | 0.771 | 0.384 | 0.195 | 0.081 | 0.034 | 0
±σrms | 16.757 | 12.937 | 6.729 | 3.432 | 1.817 | 0.937 | 0.507 | 0.265 | 0.136 | 0.091 | 0.057
 | 8.585 | 6.655 | 4.213 | 2.5 | 1.305 | 0.605 | 0.261 | 0.125 | 0.026 | -0.023 | -0.057
~451 psrms ()
~56 psrms ()
Threshold set at 500 e- and FE with ENC 60 e-rms
Timepix4 TOA resolution limited by TDC when Qin > ~10 ke-
Timepix4
Timepix3

## Slide 16: On-pixel TDC simulations  TOT resolution
Qin =1 .2 Ke-, RZ=25ns/ Ke- and FE with ENC 60 e-rms
Chart: Event Number
Chart: Qin [ke-]
| 0.5 | 0.6 | 0.8 | 1.2 | 2 | 3.6 | 6.8 | 13.2 | 26 | 51.6 | 102.8
TOT_tpix3 noise (e-) | 912.8965 | 659.396 | 357.7785 | 421.7135 | 356.4896 | 354.6324 | 225.2439 | 488.748 | 432.9907 | 318.6708 | 296.8315
TOT_tpix4 noise (e-) | 278.2575 | 176.3683 | 81.95 | 62.6234 | 66.8925 | 67.6504 | 72.3298 | 65.5425 | 58.9064 | 60.1072 | 60.8506
Chart: Event Number
Timepix4 TOT measurement dominated by ENC and not by sampling error:
~700 e- FWHM in Si Sensor @ENC 60 e-rms

## Slide 17: Digital (dynamic) Power Consumption [data-driven]
Chart: Pixel Hit Rate [Hz]
| 1 | 1.2589254117941673 | 1.5848931924611136 | 1.9952623149688797 | 2.5118864315095806 | 3.1622776601683795 | 3.9810717055349727 | 5.0118723362727229 | 6.3095734448019343 | 7.9432823472428176 | 10 | 12.58925411794168 | 15.848931924611136 | 19.952623149688804 | 25.118864315095799 | 31.622776601683803 | 39.810717055349755 | 50.118723362727238 | 63.095734448019364 | 79.432823472428197 | 100 | 125.89254117941677 | 158.48931924611153 | 199.52623149688802 | 251.18864315095806 | 316.22776601683825 | 398.10717055349761 | 501.18723362727269 | 630.95734448019323 | 794.32823472428208 | 1000 | 1258.925411794168 | 1584.8931924611156 | 1995.2623149688804 | 2511.8864315095811 | 3162.2776601683804 | 3981.0717055349769 | 5011.8723362727324 | 6309.5734448019384 | 7943.2823472428154 | 10000 | 12589.254117941671 | 15848.931924611146 | 19952.623149688792 | 25118.86431509586 | 31622.77660168384 | 39810.717055349742 | 50118.723362727294 | 63095.734448019342 | 79432.823472428237 | 100000
Power cm2 | 0.52861 | 0.52861 | 0.52861 | 0.52862 | 0.52862 | 0.52863 | 0.52864 | 0.52865 | 0.52867 | 0.52869 | 0.52871 | 0.52874 | 0.52878 | 0.52883 | 0.52889 | 0.52896 | 0.52906 | 0.52918 | 0.52933 | 0.52952 | 0.52976 | 0.53006 | 0.53044 | 0.53092 | 0.53152 | 0.53228 | 0.53323 | 0.53443 | 0.53594 | 0.53784 | 0.54023 | 0.54325 | 0.54704 | 0.55182 | 0.55783 | 0.5654 | 0.57493 | 0.58693 | 0.60204 | 0.62105 | 0.64499 | 0.67513 | 0.71307 | 0.76084 | 0.82098 | 0.90181 | 0.98226 | 1.1027 | 1.2631 | 1.4637 | 1.7045
DIG Power cm2 | 0.13191 | 0.13192 | 0.13192 | 0.13192 | 0.13193 | 0.13194 | 0.13195 | 0.13196 | 0.13197 | 0.13199 | 0.13202 | 0.13205 | 0.13209 | 0.13213 | 0.13219 | 0.13227 | 0.13236 | 0.13248 | 0.13264 | 0.13283 | 0.13306 | 0.13337 | 0.13375 | 0.13422 | 0.13482 | 0.13558 | 0.13653 | 0.13773 | 0.13925 | 0.14115 | 0.14354 | 0.14655 | 0.15035 | 0.15513 | 0.16114 | 0.16871 | 0.17824 | 0.19024 | 0.20534 | 0.22436 | 0.2483 | 0.27844 | 0.31638 | 0.36415 | 0.42428 | 0.50512 | 0.58557 | 0.70598 | 0.86642 | 1.067 | 1.3078
ANA Power cm2 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669 | 0.39669
Simulated power consumption density to be ~20 % less of Timepix3:
  Digital power consumption 25% less
    Improved pixel matrix clock distribution (DDLL)
    130nm  65nm
  Analog consumption is ~5% more
    Minimize jitter  < 50psrms
    Compensate increase in input capacitance
Chart: Pixel Hit Rate [Hz]
Timepix3
Timepix4

## Slide 18: Timepix4 (2019)
29960 µm
Pixel size | 55 x 55 µm
Array | 512 x 448
Pixels | 229376
APLL | 19
ADLL | 16
10 Gbps serializers | 16
On-pixel VCO | 28672
dDLL Columns | 448
Biasing DACs | 13
ADC | 1
Transistors in pixel | ~6000
Transistors in chip | ~1.5 bn
24700 µm
55 µm
51.4 µm
55 µm
ANALOG
DIGITAL
VCO
VCO
ADB Station

## Slide 19
29960 µm
6 wafers received at the end of January
First diced chips available last week
First in the coming days in Nikhef
24700 µm

## Slide 20: Conclusions
Timepix4 is the new particle-tracker and photon counting hybrid pixel detector designed with the support of the Medipix4 collaboration:
  Large area hybrid pixel detector with 6.93 cm2 sensitive area
  4-side buttable with <0.5% dead area (not sensitive)
  TOA: 23-bit dynamic range (1.6ms) with 195 ps LSB
  TOT: 15-bit dynamic range with 60-100 e- LSB
  PC: 8-bit or 16-bit CRW up to 8*109 hits/mm2/s
  Very configurable architecture to accommodate a large number of different applications
Total design time of ~3 years with 9 engineers  ~13 year FTE
Timepix4 was submitted to fabrication on the 7th of November
First tests in the coming days @ Nikhef

## Slide 21: Thank you !
More detailed information in: https://indico.cern.ch/event/788037/
http://medipix.web.cern.ch/
Timepix4
Timepix2

## Slide 22: SPARE

## Slide 23: Analog (static) power supply distribution
 | Total I (chip) |  | 1 WB | 2 WB | 3 TSV
Nominal Analog Power [10 µA/pixel] | ~2300 mA | Vdrop [max-min] | 68 mV | 19.6 mV | 6.9 mV
 |  | Imax pad | 118 mA | 60 mA | 57 mA
Low Analog Power [1 µA/pixel] | ~230 mA | Vdrop [max-min] | 6.8 mV | 1.96mV | 0.69mV
 |  | Imax pad | 11.8 mA | 6 mA | 5.7 mA
1 WB
2 WB
3 TSV