## Slide 1: Development of Multi-Via (MVIA) pixel with signal-charge division to realize high effective 105 dynamic range for X-ray Free-Electron Laser applications
T. Hatsui
XFEL Project Head Office
RIKEN

## Slide 2: Outline
X-ray Free Electron Laser Project at SPring-8
  Detector Requirements
Multi-via Pixel Concept
Experimental Results
  Gain
  Depletion depth
  Radiation hardness
System Development

## Slide 3: SPring-8 in Japan
Kyoto
Tokyo
SPring-8

## Slide 4: SPring-8 XFEL (2011-)
Multiport CCD: mechanical model
MVIA detector by using SOI sensor technology

## Slide 5: SPring-8 X-ray Free-Electron Laser
XFEL
High Peak brilliance
x109
Ultrafast pulse
< 10 fs
x103
Spatial
coherence
100%
x103

## Slide 6
X-ray Microscopy
Coherent Diffraction Imaging by using XFEL
100 nm
10 nm
1 nm
0.1 nm
M. Takata et.al, NATURE | VOL 408 | 23 NOVEMBER 2000 p.426
Miyano et.al.,Science Vol. 289 (2000) 739.
X-ray Diffraction
Yamauchi Lab.,(Osaka Univ. )
Dr. Shimura (International Medical Center of Japan) et.al.,
Elemental Mapping of Cell
Ultrafast motion of atoms

## Slide 7: Look at atomic structure in microscopy way Coherent X-ray Imaging by XFEL
XFEL
Crystallization
Notes: 分子量: 1.5 MDa, 総電子数: 0.7 × 106

## Slide 8: Spring-8 XFEL X-ray 2D detector Requirements
X-ray Source Characteristics
60 Hz
Variation in each shot
Pulse intensity, Sample variation, etc.
Requirements
Frame rate > 60 Hz
Single Photon Detection Capability
What would be the definition?
Device life > 1 year
Optimization Parameters
Quantum Efficiency (Q.E.)
Full Well (FW) Capacity
Pixel Size

## Slide 9: Motivation What is limiting the Full Well Capacity?
Charge generated by single x-ray photon
For 6-12 keV,
1600-3200 electrons
In optical region
～1 electron
Charge/photon ratio: too large → smaller full well

## Slide 10: Multi-Via Concept for XFEL 3.2 um pitch via
-
+
+- +- +- +- +- +- +-
+- +- +- +- +- +- +- +- +- +- +- +-
+- +- +- +- +- +-
X-ray
Via-Via: 3μm
Distance

## Slide 11: Multi-Via Concept for XFEL
+- +- +- +- +- +- +-
+- +- +- +- +- +- +- +- +- +- +- +-
+- +- +- +- +- +-
Via-Via: 3μm
High Gain
Mid Gain
Low Gain

## Slide 12: Layout in 2009-1 chip
Via Array
Pixel 85.8 x 85.8 um
A TEG with 32 x 32 pixels
(5 x 5 mm2)

## Slide 13: Pixel Core
Gain (∝N/Cf) | Via | Cf | Relative Gain
High | 270 | 50 | 480
Mid. | 54 | 200 | 24
Low | 9 | 800 | 1
Charge Amps and S/H
Most of charge to high-gain amp.
lower kTC noise
A little to low-gain amp.
Smaller Feedback capacitance
Smaller Pixel Size

## Slide 14: Signal gain of FY09-1 chip Experiment with electrical input
100 Me-
high
mid.
low
FZ SOI

## Slide 15: MVIA with FZ SOI wafer 2010-8-10：taken by Omodani
Single Photon of Mo K (18-19 keV)
Leak Current
Hor. 20 usec/div
Vert. 20 mV/div

## Slide 16: Depletion Depth Measurement preliminary bias 23 V for FZ SOI
photons/pixel/sec
Mo Kα
Q.E.
0.88 uV/e- for high gain
Henke table
Mo Kα 60 %
Eph,effective = 15.3 keV
83 mV/msec
Theory for 260 um
QEtheory = 37%
Experiment
QEexp = 38.4 %

## Slide 17: Radiation Damage and Q.E.
Thicker Depletion solve both the issues
Annual Fluence 30 Mrad
  Active Layer radiation hardness: 150 krad
Photodiode (depletion)
Active Layer (transister etc.)
Front-illumination
Back-illumination
Depletion Depth
レントゲン1000万～1億回分
Experiment
= 260 um @ bias 23 V
Taget
= 500 um @ bias 85 V
Notes: 写真撮影時の照射線量 1Rとして、10^14 photon/mm^2 -> 8.6 x 10^6 R （75000Gy）
大体1000万回程度

## Slide 18: Q.E. and Radiation Hardness vs. Depletion Depth (), and Photon Energy
∝√(V/N)
Phase I target
FZ wafer
CZ wafer

## Slide 19: SYSTEM Development
Phase I: Single module detector
2010.4-2011.3
Phase II: Tiled Detector
2011.4-2012.3

## Slide 20: 50 x 50 um Dual Gain MVIA pixel
Pixel | Shape | 50 μm square
 | Type | 2 gain Multi-via pixel
 | Noise | 100 e- (0.061 eq. photon@6 keV)
 | Full Well | 50 Me- (30 000 eq. photons@6 keV)
 | Depletion Depth | > 500 μm

## Slide 21: MVIA Detector (Dual Gain)
 |  | Target Performance
Sensor | Size (max) | 32 mm x 64 mm
 | Pixel Number | 640 x 426 (273k) / reticule 640 x 1278 (818 k) / die
 | Input (address) | 5 Serial line @ 100 MHz
 | Input (timing) | 3 (start, store, idle start)
 | Input (bias) | 1
 | Output node | 2 / reticule (high & low) 6 / module
 | off-chip ADC | 6 ch ADC (ENOB 11bit) 20 MS
I/O pads for wire-bonding
X-ray
(Flip chip)
Die has 3 reticule domains.
Notes: 縦
*print (512+2^7)*0.050
32
*print 512+2^7
640
横
*print (512-2^6)*0.050
22.4
*print (512-2^6)
448
読み出すべきピクセル数/sec
*print 640*448*60
1.72032e+07
17 Mpixel/sec

## Slide 22: Packaging
Analog Front-end
ADC and Control Logic onto flexi cables
Die (32 mm x 64 mm)
3 reticule units
I/O pad on top and bottom
Flip chip
Package
low height profile to meet industry standard

## Slide 23: MVIA Phase I (-2012.3)
Ultra-high Vacuum
Air
Serial Data Transfer
Trigger
Power
FPGA
Cameralink base config.
11 bit ADC + gain info. (1bit) = 12 bit /pixel
0.82 M pixels
60 frame/sec
0.59 Gbps
CPLD & ADC
Notes: *print 640 * 426 *3*60*12
5.88902e+08

## Slide 24: MVIA Detector (-2013.3) 32 Array
 |  | Target Performance
System | Image Area | 256 mm x 256 mm
 | Tiling | 4 x 8
 | Pixel Number | 5.1 k x 5.1 k
 | Dead area | < 0.8 mm
 | Vacuum | <10^-7 Pa
 | Cooling | Water
Die
Notes: 横
15 x 20 mm 0.076 mm pixel
0.076*256 = 19.456
60 mm -> 256*3 = 768
768*4=3072
縦
192*0.076=14.592
192*

## Slide 25: Conclusion
X-ray Free Electron Laser Project at SPring-8
  Detector Requirements
Multi-via Pixel Concept
  Charge division for higher dynamic range
Experimental Results
  Gain
  Depletion depth
    260 um achieved
    500 um anticipated
  Radiation hardness
    30 Mrad (Si) with some limitation
System Development

## Slide 26: Collaborators
KEK
  Yasuo Arai, and SOIPIX collaboration
Private Sector
  OKI Semiconductor
  Rohm
  A-R-Tec Corp
  Kyocera

## Slide 27: Detector Development Members
T. Kudo
M. Omodani
T. Tatsumi
K. Kobayashi
T. Hatsui
Y. Kirihara
T. Kameshima
H. Miyazaki

## Slide 28: SPARE SLIDES

## Slide 29: Remaining Issues
IR drop in large area format

## Slide 30: IR drop
Power consumption
Each amp. 100 uA 1.8 V
Heat
Operation < 0.1 % yield 0.3 W
IR drop
0.5 V with current process
300 W (peak)
2 amp/pixel
VDD & GND
Add 5th Metal Layer
Quasi differential output
Lower current operation
IR drop < 130 mV
Notes: print 818000*100*10^-6*2*1.8
294.48

## Slide 31: 2010-1 Submission
Experiment the large area format to solve
  IR drop
  dummy output
RIKEN
MVIA chip
VDD & GND
16 mm
15 mm

## Slide 32: Radiation Hardness
Requirement
Device Life > 1 year
1.6 x 1014 photons (12 keV) /mm2 = 30 Mrad (Si)

## Slide 33: FY09
Guard ring widht 300 um
I/O pads 100 um x 100 um
M2-M3-M4 only

## Slide 34: QE Improvement
Phase III
  MPCCD (high res.)
    Feasibility study has just started with free samples
    Targeted to be ready within FY2011
  MVIAPIX
    Now under feasibility study
    Production date
      Not clear
MPCCD Phase III
This Develop.

## Slide 35: QE Improvement
MPCCD Phase III
This Develop.
50 um
200-300 um
600 um

## Slide 36: Leakage Current of guard ring