## Slide 1: SC cavity measurements
Walter Venturini Delsolaro
CERN
EASISchool on superconductivity,
29 September 2020

## Slide 2: Topics
Introduction: measurements, errors, uncertainty
Quantities of interest
    Resonant frequency, operational bandwidth  beam requirement
    Qo value heat load on cryogenics
    Accelerating field Ea beam requirement
RF basics
  Transmission line theory; characteristic impedance, reflection coefficient
  Standing waves, resonators
  S parameters
  RF components: cables and connectors, matched loads, attenuators, splitters, directional couplers
  Measurement equipment: network analyser, spectrum analyser, power meters
Superconducting RF cavity measurements
  Qo and Ea measurement method(s)
  Sources of uncertainty, evaluation and mitigation
  Cool down precautions and effects
  Multipacting
  Field emission
  Quenches
  Q switches
  Disentanglement of Rs “components”

## Slide 3: Measurement
Determining the/a value of a measurand (the quantity to be measured), implies:
  Definition of the measurand
  Definition of the measurement unit (the value of the measurand is a real number expressing the ratio between the measurand and the unit)
  Definition of the principle of measurement (example: thermal contraction)
  Definition of the method of measurement (example: by difference)
  Definition of the measurement procedure (the sequence of actions to carry out the measurement)

## Slide 4: Errors and uncertainty
Error is an idealised concept: the “true” value of a quantity is metaphysical, sometimes even conceptually hard to define (think to the length of an object…)
Exception: conventional true values (speed of light, Avogadro number, etc..)
In science we believe “something is out there”, but “errors” cannot be known
Traditionally errors are viewed as having two components: systematic and random. The distinction has to do with their behaviour in time. However:
A “random” error in determining an influence variable or a calibration constant becomes “systematic” in the end result
Known systematic errors should be eliminated by correction factors; these are in turn affected by “errors”
Uncertainty reflects the lack of knowledge on the value of the measurand, of which the measurement results is only an estimate

## Slide 5: Sources of uncertainty
  Incomplete definition of the measurand
  Imperfect realization of the definition
  Under sampling (non representative samples)
  Uncertainty on measurement of influence variables
  Unknown effect of all environmental conditions
  Operator bias, human error
  Errors on calibration standards
  Errors on parameters in the data reduction algorithm
  Numerical errors
  Approximations in the models
  “Unknown unknowns”

## Slide 6: Standard uncertainty
The standard deviation of the probability distribution of measured values around the “true” value(s)
Type A and Type B uncertainties: the distinction refers to the method of deriving the uncertainty, conceptually they are the same, and both are accepted by the ISO norm
  Type A is obtained with statistical procedures applied to repeated observations (frequentist probability)
  Type B is obtained from an assumed probability density function based on the degree of belief that an event will occur (subjective probability)

## Slide 7: Combined standard uncertainty
The uncertainty on a derived quantity is assessed by the uncertainty propagation formula:
If
Then:
This simple formula assumes no correlations between the input quantities, correlations are often present and must be accounted for
The formula is a Taylor expansion: depending on the degree of non linearity of the model (the function f), it may be needed to include higher order terms
Simple Monte Carlo calculations can be set up to compare with the analytical formula

## Slide 8: Superconducting cavities

## Slide 9: Electromagnetic spectrum

## Slide 10: Frequency
Resonance frequency vs operating frequency
Accelerators dictate the operating frequency to the cavities, not vice versa: tuners shift the resonance frequency toward the operating frequency
Frequency control is an essential part of cavity production knowledge of frequency shifts all along, tuning range
Precise frequency not being critical for cavity testing, it can easily be forgotten. Keep in mind this is the simplest and yet essential “figure of merit” in cavity test

## Slide 11: Accelerating field
The energy gained by a charged particle traversing the cavity (on axis and with the right phase wrt to the cavity field) is the charge times an integral involving the E field over the transit
The E field in the cavity changes along the path and in time (transit time factor).
The accelerating field (aka gradient)is defined as the maximum possible voltage gain divided by a suitable (but conventional) length representing the accelerating gap
The observable from which the accelerating field is derived is usually the power transmitted by the cavity to a calibrated pick up antenna
From Ref. 7
Notes: The useful quantity is the accelerating field. We pay MV/m with Watts of dissipated power at the operational temperature (4.2 K or 1.8 K) Carnot and refrigerator efficiencies have to be taken into account to get the “real” power consumption

## Slide 12: Quality factor
Defined for every oscillator (also mechanical, optical, etc)
The time constant of the damped oscillations is Q/ω
For linear systems (in electromagnetism – ohmic) , Q is a constant
In our cavity, the stored energy is U=
And the dissipated power is Pc=
Rs is the surface resistance, which encompasses the power losses due to the RF currents on the cavity walls

## Slide 13: From Q0(H) to Rs(H)
Usually, using the definition of Q0 , the surface resistance is factorized out of the integral representing the cavity dissipated power:
Q0= = ,
where Г is a geometric constant known from RF simulations.
Then Rs(B)= . This is not correct because, if Rs(B), it can’t be factorized out of the integral. Indeed, since the cavity has a field distribution from 0 to Bpeak, Rs depends on s through B. The error is significant especially for non elliptical cavities.
When fitting models of Rs versus B (Q slopes), it is possible to extract the exact field dependence using a numerical procedure (see for example D. Longuevergne paper at SRF13): The idea is to divide the cavity surface in parts where the field is approximately constant. Then make a change of variables in the integral from s to B, using the distribution function S(B) of the fractions of cavity surface were the magnetic field is between B and B+dB

## Slide 14: Example of application
QS16 2K
<Rs>(H)
QS16 2K Rs(H)
QSS2 2K Rs(H)
QSS2 2K <Rs>(H)
If the local surface resistance uniform is Rs(x, y, z) = Rs[H(x, y, z)]
Fraction of B on the RF surface
The correction is significant when the Q slope is large, it can be neglected if the Q slope is small

## Slide 15: How to decompose the surface resistance?
Traditionally: Rs(T, B, ω)= RBCS(T, ω)+ Rres(B, ω)
Rs(T0)= Rfl+Rres=Rfl+Rres0+Rres(B)
Rfl=(Rfl0+Rfl1 Hrf) Hext
(see Physica C 351 (2001) 421-37)
RBCS(T, ω), from Matthis-Bardeen theory, is a low field approximation
Disentanglement at high field is not obvious!
Consequently, Rres is also only known at low field

## Slide 16: RLC circuit, amplitude and phase responses
R
C
L
Circuit is voltage driven, the current is the response. The modulus of the current is maximum at resonance 
Notes: Circuit is voltage driven, the current is the response. Modulus of Z is minimum (current is maximum) for ω=1/Sqrt(LC) (resonance)

## Slide 17: Driven oscillator, amplitude and phase responses
V=cos(ωt)
L
R
C
In a 1.3 GHz SC cavity with Qo ~ 1010
Δf ~ 0.1 Hz
Δf ~ 0.1 Hz
Close to resonance the phase undergoes a dramatic change, sometimes exploited to “find” the cavity

## Slide 18: RF basics- transmission lines
From lumped circuits to distributed parameters: telegraph equations
R
L
G
C
V(z)
I(z)
V(z+Δz)
I(z+Δz)
j ω L) I(z) Δz+ V(z+ Δz)=V(z)
I(z) - V(z+ Δz) j ω C) =I(z + Δz)

## Slide 19: Telegraph equations
For lossless lines
The same equation holds for the current: travelling voltage and current waves. In the other limit RC>>>LC wave propagation is suppressed the equation takes the form of a diffusion equation

## Slide 20: Propagation constant
Sinusoidal oscillations:
Differentiate the first equation and
plug in the second:
The same equation holds for the current.
Travelling voltage and current waves
k is the complex propagation constant

## Slide 21: Line impedance
Characteristic Impedance Z0 of the line: ratio of voltage to current in the forward or in the backward wave.
)=)

## Slide 22: Impedance (cont.)
=
Thus:
For a Lossless line (often good approximation) is a real number.
and is purely imaginary
Propagation without attenuation.

## Slide 23: Transmission lines: reflection coefficient
ZL
Z0
z
0
-L
Define the reflection coefficient as
The line impedance is Z0 but at the load the ratio V/I must be ZL so there must be a reflected wave
Notes: Gamma is the ratio of the complex amplitude of the reflected wave to that of the incident wave
The power actually delivered to the load thus reduced by 1 − | Γ | 2

## Slide 24: Transmission lines: Standing waves
Z0
0
Z=L
=
Case of lossless line, k=jω/v
Change z -z; ω/v=𝛽=2π/λ
=
…use Euler’s identity
=2sin()
v
…go back to real:
v
Standing wave: time and space are decoupled, no propagation

## Slide 25: A peculiar example of transmission line: the LHC dipole string

## Slide 26: Standing waves, resonators
Z0
0
Z=L
v
Can cut open the line at z=n ; (for n=1: half wave resonator)
Can also cut at λ/4 (quarter wave resonator)

## Slide 27: Quarter Wave Resonator
V
C
L
I
l~/4
l ~/4
Electric Field
Magnetic Field
f= 101.28 MHz
λ0=2.96 m
̴ βλ

## Slide 28: Transformation of the reflection coefficient and standing wave ratio
Lossless line:
Is the reflection coefficient at the distance d
SWR
Define the standing wave ratio SWR
Looking at the extremes of the modulus of V(d):
Remember: defined only for a lossless line
Distance between two maxima of the reflection coefficient= λ/2

## Slide 29: S parameters for a 2 port network
Sxy: power enters port x
Define: a1=V1+, b1=V1-, a2=V2+, b2=V2- complex voltages
S11 Input port voltage reflection coefficient S21 Forward voltage gain S22 Output port voltage reflection coefficient S12 Reverse voltage gain
a1
a2
b1
b2
│S21│ is the scalar linear gain, that is │V2- /V1+│ , the phase information is dropped.
The insertion loss is expressed in decibels, that is IL= -20 Log10(│S21│)
Notes: IL is the extra loss produced by the introduction of the device under test (DUT) between the 2 reference planes of the measurement. The extra loss may be due to intrinsic loss in the DUT and/or mismatch.

## Slide 30: Power units
Relative power levels are expressed in decibel
P(dB)=10 Log(P1/P2)
Absolute power measurements often are expressed in dBm: 1 mW is taken as a power reference level
P(dBm)=10Log(P in mW)
P(mW) = 10[P(dBm)/10]
Less than 30 dBm is less than 1 W
A negative number of dBm is less than 1 mW!
10000

## Slide 31: RF components: cables and connectors
Coaxial cables: Characteristics
  Shunt capacitance per unit length, in farad per meter
  Series inductance per unit length, in henrys per meter
  Series resistance per unit length, in ohms per meter
  Shunt conductance per unit length, in siemens per meter
  Characteristic impedance in ohms
  Attenuation (loss) per unit length, in dB per meter
  Velocity of propagation, in meters per second
  Single-mode band (in coax, TEM have zero cut-off frequency, but other modes set in at high freq.)
  peak voltage ( set by the breakdown voltage of the insulator)
Connectors
Many types: BNC, N, SMA, 7/16 DIN (rated according to max frequency)
Decreasing the diameter of the outer conductor increases the highest usable frequency
RF and microwave connectors are precision-made parts, easily damaged by mistreatment: as a rule, the only thing which has to be rotated is the threaded sleeve
Connectors give important contributions to loss and mismatch on the line

## Slide 32: RF components: attenuators, DC stoppers, matched loads shorts, open
Attenuators: reduce power (absorbing it) without distorting the waveforms.
Characteristics:
  Attenuation expressed in dB of relative power
  Frequency bandwidth
  Power dissipation
  SWR for input and output ports
  Accuracy
  Repeatability
DC blocks: prevent the flow of DC (and audio) power, used to protect sensitive equipment (for example power meters), or when these components are undesired
Loads: calibrated at the characteristics impedance of the line.
Shorts, open: mainly as calibration standards.

## Slide 33: RF components: splitters, directional couplers
Power splitters (or power combiners)
Directional couplers: characteristics
  Coupling factor in dB :-10 Log (P3/P1). always < -3 dB
  Insertion Loss in dB :-10 Log (P2/P1)
  Isolation: -10 Log (P4/P1); -10 Log(P3/P2) (they can differ)
  Directivity: -10 Log (P4/P3)= -10 Log (P4/P1)+10 Log (P3/P1)
  Directivity= Isolation-coupling factor: should be as high as possible
  Directivity is “narrow band”, it relies on destructive interference

## Slide 34: Power measurements, power meters
Below ~ 10 kHz power is measured by measuring V and I: it is a derived quantity
Up to 100 MHz direct measurements of voltage and current are still possible; but it is easier and more accurate to measure power directly with a power meter
In the GHz range and beyond, power is the fundamental quantity: voltage and current are the derived quantities
Thermistors (bolometers), Thermocouples, and diode detectors are the most common sensors used to measure RF power

## Slide 35: Spectrum analyzer

## Slide 36: Resolution Bandwidth
Larger RBW | Smaller RBW
Smaller FFT size; fewer samples; requires less acquisition and computation time; often unable to resolve two closely spaced tones in a spectrum. | Larger FFT size; more samples; requires more acquisition and computation time; tones are easily resolved.

## Slide 37: Network analysers
Measure S parameters as a function of frequency
S parameters are complex quantities
Two types:
  Scalar Network Analysers (SNA): measure amplitudes
  Vector Network Analysers (VNA): measure amplitudes and phases
VNA are the most common nowadays
Usually two ports (4 ports and more are also available but expensive)
With a two ports one can measure S11, S22, S12, and S21 (reflection and transmission from or through a DUT).
Measurement output in several formats
Require calibration performed by the user at each use on top of the periodic calibration at the factory

## Slide 38: Basic measurement circuit
RF power amplifier
LLRF drive
Matched load
Circulator
Directional coupler
Cryostat top plate
Forward power meter
Reflected power meter
Transmitted power meter
Input coupler
Pick up antenna
LHe

## Slide 39: Attenuation measurements (cable calibrations)
The readings of the power meters on the rack must be referred to the powers at the cavity, taking into account the attenuation of the lines:
The total attenuation values are:
FWD= (A→FWD) - (A→C) - Cold part (coupler)
RFL=(B→RFL ) + (B→C) + Cold part (coupler)
TXM=(D→TXM) + Cold part (antenna)

## Slide 40: Cavity measurement analysis
Equivalent circuit model (for one mode)
Gc
C
L
Gg
1:n
Generator + Circulator
Transmission line
Input Coupler
Cavity

## Slide 41: Cavity measurement analysis: undriven cavity:
;  Loaded Q:
; U decreases exponentially, the time constant is
++; defines intrinsic and external Q’s
Coupling parameters: ; :
they tell how strongly the couplers interact with the cavity.
=+
While external Qs only depend on geometry ( overlap of field lines between cavity and coupler), the betas depend on the cavity Q.

## Slide 42: Cavity measurement analysis: driven cavity with one coupler
From the equivalent circuit, the reflection coefficient of the cavity seen from the transmission line can be computed:
The conductance seen from the generator, relative to is and the susceptance is
Then, the complex reflection coefficient is
and 𝛿 is the detuning factor
The power flowing in the cavity is
The reflected power can be cast in the form
The stored energy obeys the equation:
)  )
(detailed derivations in Ref [8], p. 150)
Equilibrium value of U for a given Pf

## Slide 43: Steady state
from the expression of the reflected power
we derive
the coupling coefficient is thus determined from the measurement in CW of the reflected and incident power provided it is known if 𝛽>1 or 𝛽<1
From the differential equation for E, by setting proper initial conditions, we can analyse the reflected power waveforms at RF turn on and off.
These allow to distinguish the type of coupling

## Slide 44: RF off - transient
)
If we turn off the RF power at t=0, then U0 =0 ( the equilibrium value of the stored energy becomes zero
,
 the time constant for the exponential decay of the field is 2
When measuring the time constant make sure you know if you are measuring the power or the amplitude
Now the RF is off, so the reverse travelling power is only the emitted power.
Therefore we can write (definition of Qe)
Pr=Pe= , U(t)=U0 exp(-t/), and is the equilibrium value of U before the RF is turned off.

## Slide 45: Coupling determination from RF off transient
Another way to determine the coupling factor is to compare the forward power before RF is turned off with the peak of the emitted power immediately after:

In this case we don’t need to know beforehand if the system is under coupled or over coupled

## Slide 46: Determination of the coupling factor
Reflected power from a square pulse when 𝛽>1
Voltage waveforms

## Slide 47: Locking the cavity: generator driven systems
Block diagram of a VCO PLL System. The encircled elements can be replaced by an RF generator which accepts frequency modulation from the DC output of the mixer (phase error-> frequency error around resonance).
From Ref. 8

## Slide 48: Locking the cavity: Self excited loop
High gain, positive feedback loop: the cavity selects its own resonance frequency from the noise floor and the circuits starts to oscillate. It only works with correct phase shift on the return path.
Finds and tracks the cavity resonance

## Slide 49: Real measurement circuit

## Slide 50: Test stand

## Slide 51: Measurement procedure
1) Lock the cavity on resonance
2) Calibration point:
Determine input coupling condition
Measure in CW: Pi, Pr and Pt
Measure the time constant of U
From 𝛽 and 𝜏L Q0=ω0 𝜏L (1+ 𝛽)
From the measurement in CW, obtain the dissipated power Pc=Pi-Pr-Pt
Find the stored energy by the definition of Q0:
Find the accelerating field
Find the calibration constant of the pickup
2) Q-E scan
Increasing the power level, U is determined by the transmitted power, knowing hence the field is determined knowing
The corresponding value of Q is determined by the definition, knowing the stored energy and the dissipated power
MP
ke is obtained by means of computer simulations

## Slide 52: Warm measurement of high Qext (pickup calibration)
The external Q of the pickup probe is essential for the measurement of Eacc
As good practice it should be chosen ~ 100xQ0
It is derived when measuring Q0 with the decay time (calibration)
It is good to measure it at warm as a consistency check (it is a geometric parameter)
Qext can be determined from the reflection coefficient and the loaded Q
A VNA can display a polar plot of the reflection coefficient, normalised to the characteristic impedance (Smith chart)
Remember , and

## Slide 53: Warm measurement of Qext with network analyser
1st method (used for not too far from 1):
Determine from the Smith chart if the cavity is under- or over-coupled
Measure the Standing Wave Ratio (SWR) on resonance
,
Determine QL from the 3dB bandwidth:
Q0=(1+),

## Slide 54: Warm measurement of high Qext with network analyser (pickup calibration)
2nd method, used for pickup calibration, must use a coupler with not too far from 1
Measure from S11 (see method 1)
Measure S21 and QL
: (Ex. derive it from Pin= Pf - Pr=Pc+ Pt )

## Slide 55: Measurement Setup
Reference Antenna
( critical coupled)
Pick-Up Antenna
S21 -> Qext,Pick-Up
input
Pick uo

## Slide 56: Interpretation of cavity behaviours
Fitting Q(T) data with the standard BCS expression reveals the so called “non BCS losses”
The global transition of the cavity to the normal conducting state is a catastrophic event, can happen to bulk Nb cavities in case of large beam losses
Multipacting is easily recognized by the “field lock” feature. Also is recurs at the same field levels for the same cavity geometry
Field emission most striking mark is the emission of X rays. The endpoint energy of the spectrum corresponds to the maximum kinetic e- energy
Many theories available, a lot still to understand
Q switches are small (local) quenches. The cooling power is sufficient to prevent global quench. Typical hysteretic behaviour

## Slide 57: Multipacting: SEY of Nb

## Slide 58: Multipacting (Cavity Top)
70ns
87.5ns
50ns
30ns
8ns
# of initial electrons: 742
Kinetic energy: 0-4eV (uniform)
Emission angle: 0-180 deg. (random)
Max time steps: 4000
Eacc=1MV/m (30 deg)
Pei Zhang

## Slide 59: Regions of Potential Multipacting
E-field
H-field
Pei Zhang

## Slide 60: QWR multipacting Summary
MP happens at low fields for antenna region and outer wall
- Eacc = 0.05-0.1MV/m
MP happens at higher fields for cavity top region
- Eacc = 0.8-2 MV/m
MP-free region
MP region
Pei Zhang

## Slide 61: Cavity spectrum during multipacting conditioning

## Slide 62: Field Emission
Chart: Eacc (MV/m)
| 2.9758977859847199E-2
calibration point | 505428415.9404
Q3_3 second meas | 111015284.4603 | 104759397.4243 | 217962828.1257 | 197298140.002 | 242598546.8752 | 269210131.1253 | 294518067.3046 | 320824431.4508 | 343805363.761 | 359805318.4325 | 376083402.396 | 388630106.4109 | 399600280.8639 | 412328255.668 | 425650109.6223 | 449187590.5377 | 460221662.0894 | 478500426.6602 | 489342683.1478 | 504068344.3157 | 514251141.4018 | 526930676.5184 | 535821321.3149 | 545491351.7696 | 553445484.7961 | 560462744.468 | 562008038.2524 | 569888584.5231 | 534968649.9258 | 189214994.5395 | 173912775.3393 | 158742170.3783 | 147264156.4736 | 136452870.4931 | 126930791.0886 | 111580730.616 | 103885419.364 | 98760596.295 | 91275387.6182 | 87453517.2911
10W | 1317.2673 | 5269.0691 | 32931.6822 | 64546.097 | 84305.1063 | 131726.7286 | 526906.9146 | 1185540.5577 | 3293168.2159 | 6454609.7032 | 13172672.8638 | 22261817.1398 | 29638513.9435 | 52690691.455 | 82329205.3985 | 118554055.7738 | 161365242.5811 | 210762765.8202 | 266746625.4912 | 329316821.594 | 398473354.1288 | 474216223.0954 | 556545428.4939 | 645460970.3243

## Slide 63: Temperature & Radiation
The cavity bottom was heated up exponentially and the radiation level went up exponentially as well, a clear indication of field emission.

## Slide 64
X rays spectrum (bremsstrahlung)
Chart: Energy, keV
In principle it is possible to calibrate field measurements with X rays spectrum, but it is quite tricky: only a small fraction of the emitted electrons experience the full accelerating voltage. One needs to resolve the upper tail of the spectrum from the background….

## Slide 65: Electron field emission
Eon(2nA) =140 MV/m
= 31, S = 6.8·10-6 µm2
Al Mg Nb
Eon(2nA) = 132 MV/m
 = 27, S = 7·10-5 µm2
Eon(2nA) > 120 MV/m
= 46, S = 6·10-7 µm2
Al
Si Nb
500 nm
2 µm
2 µm
Typical particulate emitters

## Slide 66: He processing
Chart: Eacc (MV/m)
| 2.9758977859847199E-2
calibration point | 505428415.9404
Q3_3 first meas | 403099123.2731 | 391694511.1917 | 374923653.4462 | 361147764.4913 | 346113536.3188 | 330190774.6121 | 314024130.2465 | 293816352.0444 | 268745789.3104 | 245893421.2108 | 217753392.6486 | 195350874.166 | 174888012.6218 | 159272979.6891 | 153330048.2718 | 410130210.7662 | 426168002.8887 | 438743710.6255 | 448024892.384 | 461774585.7394 | 471574238.0554 | 481795494.2013 | 491296069.9315 | 500050061.5277 | 507871204.2584 | 516528236.2341 | 521011544.8198 | 530161116.6174 | 537064058.0711 | 542949927.7265 | 546458992.6071 | 550773022.9274 | 551746409.2159 | 555945407.9081 | 563128506.6771 | 565284853.7069 | 566910987.5935 | 572115704.6742 | 562082173.363 | 545168134.1341
Q3_3 second meas | 111015284.4603 | 104759397.4243 | 217962828.1257 | 197298140.002 | 242598546.8752 | 269210131.1253 | 294518067.3046 | 320824431.4508 | 343805363.761 | 359805318.4325 | 376083402.396 | 388630106.4109 | 399600280.8639 | 412328255.668 | 425650109.6223 | 449187590.5377 | 460221662.0894 | 478500426.6602 | 489342683.1478 | 504068344.3157 | 514251141.4018 | 526930676.5184 | 535821321.3149 | 545491351.7696 | 553445484.7961 | 560462744.468 | 562008038.2524 | 569888584.5231 | 534968649.9258 | 189214994.5395 | 173912775.3393 | 158742170.3783 | 147264156.4736 | 136452870.4931 | 126930791.0886 | 111580730.616 | 103885419.364 | 98760596.295 | 91275387.6182 | 87453517.2911
10W | 1317.2673 | 5269.0691 | 32931.6822 | 64546.097 | 84305.1063 | 131726.7286 | 526906.9146 | 1185540.5577 | 3293168.2159 | 6454609.7032 | 13172672.8638 | 22261817.1398 | 29638513.9435 | 52690691.455 | 82329205.3985 | 118554055.7738 | 161365242.5811 | 210762765.8202 | 266746625.4912 | 329316821.594 | 398473354.1288 | 474216223.0954 | 556545428.4939 | 645460970.3243
X ray dose rate data
Uses RF power in ignite a low pressure plasma in the cavity filled with gaseous helium
He pressure must be high enough for this but not higher, risk of breakdown and damage to cavity and couplers
Used as a last resort against field emission when cavity can’t be easily removed and rinsed with high pressure water or other chemical or mechanical treatments

## Slide 67: Low Field Q switch
Q0 lost 2%
ΔRs = 0.5 nΩ
Rs ~ 25 nΩ

## Slide 68: Example of giant Q switch (NC transition of the tuning plate in a QWR)
Chart: Eacc(MV/m)
| 6
HIE-ISOLDE specification | 470000000
10 W | 1317.2673 | 5269.0691 | 32931.6822 | 64546.097 | 84305.1063 | 131726.7286 | 526906.9146 | 1185540.5577 | 3293168.2159 | 6454609.7032 | 13172672.8638 | 22261817.1398 | 29638513.9435 | 52690691.455 | 82329205.3985 | 118554055.7738 | 161365242.5811 | 210762765.8202 | 266746625.4912 | 329316821.594 | 398473354.1288 | 474216223.0954 | 556545428.4939 | 645460970.3243

## Slide 69: Sources of uncertainty in SC cavity measurements
Power meter accuracy and resolution
  Determination of coupling β
  Direct measurement of power in Q - E scans
Cross contamination of forward and reflected power due to poor directivity of directional couplers
Detuning
Dependence of cable attenuations on power level
Uncertainty on time constant measurement due to
  Q slope (non exponential decay)
  Multiple reflections during decay (circulator mismatch)
  Fit error
  Instrument accuracy
Uncertainty on geometric constants (simulations assume perfect geometry)

## Slide 70: Q0 measurement uncertainties
…from =
()2=()2 ()2 + ()2 ()2
under ideal conditions this leads relative uncertainties of a few %.
See for example: O. Melnychuk, A. Grassellino, and A. Romanenko, Review of Scientific Instruments 85, 124705 (2014)
However this assumes:
perfect directivity of the directional couplers,
perfect matching of the circulator protecting the amplifier
no detuning during the decay measurement
These effects can be very important in some conditions

## Slide 71: Directivity error experiment
In theory, the measurement of the forward power from the directional coupler port should not depend on the phases, but…
Picture from J.P. Holzbauer, Yu. Pishalnikov, D. A. Sergatzkov, W. Schappert, S. Smith, Nucl. Instrum. Meth.A 830 (2016) 22-29

## Slide 72: Directivity error experiment
From J.P. Holzbauer, Yu. Pishalnikov, D. A. Sergatzkov, W. Schappert, S. Smith, Nucl. Instrum. Meth.A 830 (2016) 22-29

## Slide 73: Bias on τL due circulator mismatch
Picture from J.P. Holzbauer, Yu. Pishalnikov, D. A. Sergatzkov, W. Schappert, S. Smith, Nucl. Instrum. Meth.A 830 (2016) 22-29

## Slide 74
Improving measurement tools: variable coupler
Constant (and minimal) measurement uncertainty independent of cavity Q0
2) Easiness of MP conditioning

## Slide 75: Bias from reflected power

## Slide 76: Simple cross checks
The external Q of fixed couplers should be constant at all field levels, and equal to the values found at the calibration step. It can be computed from Q0 and 𝛽 all along the Q-E scan.
with the phase locked-loop open and the generator drive frequency off resonance, Pr/Pf should stay close to unity at all power levels, and constant over a small bandwidth ()

## Slide 77: Double-check the measurement
Four different calibrations were conducted
- over coupled, under coupled, critical coupled

## Slide 78: Cool down precautions
The way the cavity is cooled down can influence the result: main effects are:
«H disease»
Flux trapping
Thermo electric currents

## Slide 79: Effects of thermal gradient when crossing Tc
First cool down with high thermal gradient
Thermal cycle above Tc, second cool down with more homogeneous temperature

## Slide 80: Residual Resistance due to Flux trapping
)
External magnetic field: Bout , Ba ,BNC , Bavg , Bo
Expulsion inefficiency, or trapping efficiency, or trapping ratio, t , τeq , Btrapped/Bout.
Related to the observable field enhancement BSC/BNC at cavity equator when crossing Tc ( expulsion ratio, Posen ratio, r, εeq )
Sensitivity (in nΩ/mG): S, R0/Btrap , rfl

## Slide 81: Temperature and field Q scans
“Field cooled”
Akira Miyazaki

## Slide 82: Analysis of Q slopes
B-field compensated when cavity crossed Tc
B-field enhanced when cavity crossed Tc
enlarged
Intrinsic Q-slope is temperature dependent curvature component
Q-slope by trapped vortex is close to linear at low fields and weakly temperature dependent
2.4K
4.6K
2.4K
4.6K
Akira Miyazaki

## Slide 83: An empirically found formula to fit the non-linear Q-slope
from Rs vs T data
Fit Rs vs H and determine
2.3K
4.5K
T [K] | a [T-1]
2.3 | 37
4.5 | 27
Such an exponential dependence has been reported by others (bulk Nb and Nb/Cu)
R. L. Geng (Cornell) ”Thermal analysis of a 200MHz Nb/Cu cavity” SRF2001
D. Longuevergne (IPNO) ”Magnetic dependence of the enegy gap:…” SRF2013
Akira Miyazaki

## Slide 84: Cryomodule testing
Cryomodules are complex units ready for installation in accelerator, the power coupler is already adapted to the beam loading (usually β>>1)
Besides RF measurements, static heat loads and alignment are important
Precise measurement of Q0 with RF methods is impossible (without beam, almost all power is reflected back)
Calorimetric methods are applied
Field measurement are still possible (if pickup calibration was preserved in cavity assembly)

## Slide 85: Example of dynamic heat load measurement on a cryomodule
RF on (total ~ 31 W)
RF off : mass flow reduces by ~ 1.5 g/s
DC heaters on (31 W DC)

## Slide 86: First beam acceleration with a HIE ISOLDE superconducting QWR

## Slide 87: Linac phasing with 4 cavities: first try
~3 MeV/Cavity
~3.9 (±0.16) MeV/u (12C4+ )

## Slide 88: Sources
[1] Wikipedia.org
[2] Guide to the expression of uncertainty in measurement, BIPM, IEC, ISO, 1995
[3] J. C. Slater: Microwave electronics, Rev. Mod. Physics, 18, 4, p. 441, Oct. 1946
[4] Jackson: Classical Electrodynamics John Wiley & Sons Ltd. 1962
[5] S. Ramo, J. R. Whinnery, T. V. Duzer: Fields and waves in communication electronics, John Wiley & Sons Ltd. 1965
[6] D. M. Pozar, Microwave Engineering, John Wiley & Sons Ltd. 2005
[6] A. Facco, Tutorial on low beta cavity design, proceedings of the 12th International workshop on RF superconductivity, Cornell University, Ithaca, New York, USA
8] T. Powers, “Theory and Practice of Cavity Test Systems” 2005 SRF Workshop.
[9] H. Padamsee, H. Knobloch, T. Hays, RF Superconductivity for Accelerators. Wiley 2007
[10] O. Melnychuk, A. Grassellino, and A. Romanenko, Review of Scientific Instruments 85, 124705 (2014)
[11] J.P. Holzbauer, Yu. Pishalnikov, D. A. Sergatzkov, W. Schappert, S. Smith, Nucl.Instrum.Meth.A 830 (2016) 22-29
[12] J.P. Holzbauer, et al, https://arxiv.org/abs/1804.04747