## Slide 1
1. Superconducting magnets
2. Field control issues and solutions
3. Power supply ripple
Superconducting magnets and field control for medical accelerators
Marco Buzio, Stephan Russenschuck Magnet, Superconductors and Cryostats Group, Technology Department, CERN
Notes: Previous tutorial: Metrological basics, SI and CGS units, relationship to electrical quantities, standards, quantities of general interest for accelerator magnets (not repeated here)

## Slide 2
Superconducting magnets

## Slide 3
Hard superconductors – a recap
  Strong magnetic fields can penetrate by discrete flux quanta trapped by pinning sites
  Flux gradient → non-dissipative DC transport current at critical current density JC
  De-pinning → electrical field E → strongly non-linear dependence JC(E )
  Flux penetration creates persistent currents → magnetization with hysteresis → DC field distortion + AC losses
  Thermal stability + reduction of persistent currents → small filaments in a Cu matrix → local and long-range dynamic coupling → unique decay and snapback phenomena
  transition to resistive state (quench) triggered by local heat deposition or excess current/field
Nb3Sn - brittle intermetallic compound - higher performance but strain-dependent, difficult to produce - many prototype, models and one-off projects (e.g. ITER TF coils) - first large-scale accelerator magnet deployment planned for HiLumi LHC (11 T bending in LS2, IR quads in LS3)
NbTi: - ductile alloy - industry standard - many applications in accelerator and experimental magnets, MRI, electrical distribution and machines
(Ec = 1 µV/cm)
Normal conductor
Critical State Model
Notes: HTC still not viable for deployment in accelerator magnets
Nb3Sn for accelerator magnets: some wiggler, maybe magnets deployed at BNL/FNAL ?

## Slide 4
Normal - vs superconducting magnets (1/2)
  SC magnets are used for regular operation in synchrotrons since the 70s, but no practical applications to hadronteraphy synchrotrons exist yet. Why ?
  For High Energy Physics, superconductivity is an enabling technology so high complexity and cost are more easily tolerated
  For medical applications, solutions already exist; the goal is lowering total costs → choice is less clear-cut
  Magnet total cost:
  Capital cost: direct: design, prototyping & testing manufacture indirect: special power supplies (e.g. purely inductive loads for SC) installation constraints (size, weight)
  Operation cost: direct: excitation power (only clear winning point for SC) cooling power maintenance indirect: downtime operational flexibility (cycling constraints)
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 5
Normal - vs superconducting magnets (2/2)
 | Normal conducting | Superconducting
Current density | 2~5 A/mm2 (Cu, air-cooled) 5~10 A/mm2 (Cu, water-cooled) | 9500~3000 A/mm2 (NbTi @ 4.2K, 1~5 T) ~400 A/mm2 (LHC NbTi cable @ 1.9 K, 9 T)
Field limitations | 1.5~2.3 T (linear range) ∞ (above saturation) | ~10 T (NbTi @ 1.9K) ~16 T (Nb3Sn @ 1.9K)
Magnetic forces | Magnetic pressure σ[MPa]  0.4 B[T]2 (practical limit 45~50 T) | 
Dissipation | Ohmic losses RI2 (e.g. 140 kW/magnet in CERN SPS) | Heat influx + AC/beam losses Losses @ 4.5 K  240 Losses @ 1.9 K  900
Cooling power | Demineralized water flow + cooling plant (typically ~1 kW/magnet) | e.g. 40 MW LHC cryogenic system (~2.5 W/m @ 1.9K, 20 kW/cryomagnet)
Field quality | Geometry + iron effects (saturation, hysteresis, eddy currents) | Geometry + iron (if applicable) + persistent currents (magnetization, ramp rate, decay, snapback)
Interconnects (electrical / vacuum) | Easily accessible | Cold (inside cryostat)
Resistive (DC) voltage drop per magnet | Typically a few V (R ~ 10 m) | << V (R ~ n) Interconnections, splices
Protection | Thermal switch interlock (T<60 deg) | Quench detection + Active/passive coil protection system
Notes: Nonlinear phenomena are often mixed (some iron in all SC magnets)
Quench protection / diode MUST NOT fail in LHC
Passive protection: nothing to do with temperature margin, but rather with cooling efficiency (e.g. force flow) or energy dump

## Slide 6
Iron-dominated dipole designs
  C-shape (η=3) Simple, easy access; odd harmonics
  H-shape (η=4.6) improved field quality, less weight
  Superferric H-shape (η=4.5) corner saturation even at low field
  Window frame (η=2) best field quality, less/uniform saturation
  Iron efficiency
  Dipoles = main driver for the ring size (and civil engineering costs)
  Field quality determined mainly by pole shape + saturation
Notes: Consider dipoles as the main drivers for overall synchrotron size
LEP dipole
LHC warm TL

## Slide 7
Superferric dipole example: SIS 100
  Largest superferric synchrotron project, currently in construction at GSI (FAIR)
  based on internally force-cooled JINR Nuclotron cable (AC losses)
  Efficient cold iron yoke: less magnetic leakage, coil closer to iron, but more complex, larger cold mass, limited access
  1.9 T @ 13.1 kA
  Fast-ramped at 4 T/s
  Bending radius 52.6 m
  Magnetic length 3 m
Notes:   5.7 Tm  5 units could bend C12 @ 400 MeV/u, IF curvature would allow

## Slide 8
Coil-dominated dipole designs
  LHC (η=1.12) Best field quality
  racetrack coils are comparatively easy to make (ideal for models and demonstrators)
  compatible with SC tapes
  Field quality hardly adequate for accelerator purposes
  Cos- design approximates 2D theoretical J distribution with no leakage
  Most efficient: peak field @ coil  central field + 5%
  Best choice for high-quality, high-field large magnets; but long and costly optimization/prototyping
  Iron still gives non-negligible contribution (~1 T in LHC dipoles) + suppression from leakage and cross-talk; also gives measurable saturation
  Nb3Sn Racetrack Model Coil (16.2 T)
Notes: Notice the lower field in the iron yoke compared to the window frame

## Slide 9
Example: SIS 300 bent cos- dipoles
  FAIR SIS300 prototype dipoles
  4.5 m long, 4.5 T
  Fast-ramping at 1 T/s  filament  3.5 m
  Max curvature ever achieved in Cos- design: 66 m bending radius
  Single-layer coil wound directly on a bent mandrel, supported by the collars only
  Difficult construction, design seems hardly extrapolable to stronger curvature …
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 10
Canted Cos- design
  Proposed by Meyer in 1970 , commercialized as Double HelixTM by AML Superconductivity
  Not used as yet in accelerators, but for posible applications see e.g. L. Bottura’s talk (gantry) and E. Benedetto’s talk (synchrotron ring)
  Design based on cancellation of longitudinal field components  intrinsically inefficient, possibly sensitive to mechanical errors
  Pros: multipolar field possible with nominally ideal geometry up to several Tesla, even on a strongly curved former; magnets can be nested; distributed mechanical support
  Major drawback: magnetically inefficient, peak field at conductor  central field + 25%
* Meyer, Flasck, Nucl. Instr. and Methods, 80, 1970
Notes: Simple and clean design made possible by very precise numeric CNC
Projected azimuthal current distribution is of cos-theta type

## Slide 11
Example of compact SC synchrotron: Aurora
  Compact SC electron SR for X-ray lithography made at Sumitomo (1989), at Ritsumeikan university since 1995
  150 MeV to 650 MeV, 1 T to 4.3 T (Bρ=0.5 to 2.2 Tm)
  Solenoid design, no quad (field index changed via saturation of the iron poles and yoke)
Notes: Demand for X-ray lithoonvegraphy suprceded since thanks to improvements in conventional optics

## Slide 12
Example of compact SC synchrotron: Helios
  Two compact electron synchrotron X-ray sources for microchip lithography built at Daresbury in 1992 and 1998
  Max energy 700 MeV (Bρ=2.3 Tm)
  Two-sector racetrack with two 180°, 4.5 T air-core, cold-bore bending dipoles
  97% beam availability observed at IBM
  25 tons including girder. Transports on a single lorry
Notes: Very little is said in the literature about the magnet (M Wilson’s baby)

## Slide 13
Magnetic field control: issues

## Slide 14
Issues in magnetic field control: iron hysteresis
Example: ISOLDE HIE TL bending
Arbitrary minor loops tend to wander inside the limit (major) hysteresis loop
As a rule, prediction using established methods (FEM/BEM, Jiles-Atherton, Preisach …) is currently not possible to the precision usually required
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 15
Issues in magnetic field control: eddy currents
  Impact of eddy currents on ramps: dissipation increases the are of the hysteresis loop
  Linear extrapolation of measurements to quasi-static condition is possible but not always precise
  Classical solution to reduce impact of eddy currents at the end of ramp-up: current overshoot
  Easily saving 1.5
  Moves response onto new limit hysteresis loop
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 16
Issues in magnetic field control: superconductor hysteresis
Superconducting filament magnetization (persistent eddy currents)
large hysteresis with relative errors of the order of 10-3 at low field (injection)
hysteresis depends on temperature, current and current history
main field and multipoles affected in different ways
Iron saturation
affects only small area in the collar (B>2T)
relative errors  10-2 at high field
additional multipoles generated
Linear regime (geometric contribution)
field is proportional to the current (can be computed with Biot-Savart’s law)
the T.F. depends only on the coil geometry
injection
Similar issues for superconducting magnets: different orders of magnitude may make feed-forward prediction possible (e.g. LHC’s FIDEL semi-empirical model)
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 17
Coupling currents
finite inter-filament and inter-strand resistance (RC) gives rise to loops linked with changing flux
multipole errors  Ḃ, RC-1
hysteresis depends upon field level, temperature and excitation history
Decay and snap-back
superconductor magnetization and coupling currents interact in a complex way  long-term logarithmic time dependence effects (field decay)
hysteresis branch switching at the end of decay  sudden current redistribution and additional multipole errors (snap-back)
Dynamic effects in type 2 superconductors

## Slide 18
Issues in magnetic field control: pre-cycling
  Blah blah
  Blah blah …
Stable hysteresis loop achieved at the third repeated cycle (result well in line with broader experience on similar dipoles)
eddy currents decay on current plateaux
return to previous hysteresis loop branch when cycle resumes
ELENA reference dipole
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 19
Issues in magnetic field control: remanent field
degaussing
transient
stable cycles
reference state
NMR marker triggers
(Nov 2016) Imax= 400 A: 0.45  0.02 mTm (~25:1)
(Oct 2017) Imax= 326 A: 0.86  0.03 mTm (~29:1)
In both cases: (Bd)r  5-8·10-5 of full range (needs to be taken into account)
Measurement on the pole surface of an ELENA MB
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 20
Magnetic field control: solutions

## Slide 21
Options for real-time magnetic measurements
LHC field model: FIDEL
SC snap-back (at restart of ramp-up)
linear component generated by excitation coils
SC magnetization  to critical current density Jc and  filament
Residual field (constant)
Iron saturation (includes also Lorentz force displacement effects  I2
SC field decay during plateaux
SC AC coupling currents  ramp-rate
parametric model (all magnets): DC components AC components
ideal magnet (1st approx.)
Source: N. Sammut, FIDEL - The Field Description for the LHC, LHC-C-ES-0012
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 22
Options for real-time magnetic measurements
Single-sensor setup | B-train system | Multi-sensor setup
used in ISOLDE GPS/HRS/HIE local feedback to FGC3 converters Bandwidth: - few kHz for Hall probes - less than 0.1 Hz for NMR (but much higher precision) main limitation: difficult calibration ofm, accurate only on a given limit hysteresis loop | At CERN in LEIR, PSB, PS, SPS, ELENA; all major hadrontherapy machine distributed to RF, FGC3, BE/BI Bandwidth: > 100 kHz (RF requirement) main limitations: cost, software strongly coupled to control system Ideally with sensors in the middle of the gap  reference magnets | Under development for TL dipoles general-purpose analog/digital local feedback to any controller based on n inexpensive Hall probes in the fringe field region + parametric m always converges to field integral with constant m for n Bandwidth: few kHz
Single-sensor setup | B-train system | Multi-sensor setup
used in ISOLDE GPS/HRS/HIE local feedback to FGC3 converters Bandwidth: - few kHz for Hall probes - less than 0.1 Hz for NMR (but much higher precision) main limitation: difficult calibration ofm, accurate only on a given limit hysteresis loop | At CERN in LEIR, PSB, PS, SPS, ELENA; all major hadrontherapy machine distributed to RF, FGC3, BE/BI Bandwidth: > 100 kHz (RF requirement) main limitations: cost, software strongly coupled to control system Ideally with sensors in the middle of the gap  reference magnets | Under development for TL dipoles general-purpose analog/digital local feedback to any controller based on n inexpensive Hall probes in the fringe field region + parametric m always converges to field integral with constant m for n Bandwidth: few kHz
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 23
Single sensor setup limitation: magnetic length reproducibility
Courtesy G. Golluccio
large fluctuations due to history-dependent residual field reproducibility degrades at low field
saturation tends to erase previous magnetic history  better reproducibility at high field
eddy current decay (=0.2 s)
m drops due to saturation in the ends
central coil replaced by NMR (DC cross-check)
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 24
B-train systems
Real-time knowledge of the magnetic field in the main bending magnets is essential for longitudinal and transversal beam control
Mathematical modelling is an option … (e.g. FIDEL, synthetic B-trains …)
… but precise predictions in iron-dominated magnets are hard  today we must measure ! (hysteresis, eddy currents, steel ageing, temperature drifts, I(t) measurement errors …) B-train systems at CERN: LEIR, PSB, PS, SPS, AD, ELENA; BNL: AGS; hadrontherapy: HIT, CNAO, MedAustron
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 25
B-train consolidation at CERN
  Blah blah
  Blah blah …
The CERN-wide B-train consolidation project FIRESTORM (Field In REal-time STreaming from Online Reference Magnets)
Optical fiber serial transmission of absolute field values (non-incremental) via CERN’s White Rabbit protocol ( Ethernet + nanosecond synchronization) (now an IEEE standard adopted by National Instruments et al.)
Industrial PICMG1.3 PC (SLC3) 64 bit GNU/Linux Bmeas FESA3 class for remote configuration and diagnostics
SPEC PCIe carrier card + FMC mezzanine VHDL can be updated remotely quickly

## Slide 26
New PS B-train commissioning
Short-term repeatability ~ 0.02 G
Δ w.r.t legacy: < 0.2 G @ injection, < 2 G @ extraction
with FIRESTORM: 27 Hz ripple disappears, effective field resolution from 0.1 to < 0.02 G, frequency resolution from 8 to 1 Hz
http://elogbook.cern.ch/eLogbook/event_viewer.jsp?eventId=2371101, http://elogbook.cern.ch/eLogbook/event_viewer.jsp?eventId=2371081
feedback from legacy system Bup/Bdown
feedback from new FIRESTORM B(t)
0.2 G ripple
0.02 G resolution
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 27
Multi-sensor setup: motivation
PS BOOSTER switching dipole on the extraction line to ISOLDE/PS
Cycling conditions: essentially random switching between ±1.4 GeV and ±2.0 GeV, 0.5 s rise time, 0.5 s plateaux
±2·10-4 peak non-uniformity in a 38  42 mm2 GFR center on the nominal beam path
“regular” B-train not easily feasible, sensors must be installed next to vacuum chamber
Similar requirements by dozens other magnets in CERN transfer lines
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 28
Multi-sensor setup
Intermediate alternative between single-sensor setups and full-fledged B-train
Meant mainly for individual beam-line magnet control  no integral coils, small sensors to be installed in the gap next to vacuum chamber
stand-alone instrument + optional add-on interface to control system
analog/serial output for flexible interface to converters
Off-the-shelf electronics and sensors (commercial Hall probes)
proof-of-principle testing under way
Final cost expected be one order of magnitude less than a B-train
Hall probes and/or other sensors (coil, NMR/FMR markers)
Dynamically switching between arrays for L/R bunches (if L/R asymmetry important)
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 29
Magnetic sensors in the fringe field
  CERN LEIR B-train
  Coils and FMR markers in the fringe field of LEIR bending dipoles give results comparable with legacy B-train within ±5 G (sufficient to accelerate and extract beam)
HIT B-train
coil wound around the yoke + a single Hall probe market at flat top
field error reduced 6 
dead time by 30%
reproducibility 50 ppm
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 30
Power supply ripple

## Slide 31
Example: APOLO on ELENA MB
  CERN-made APOLO_2p switched-mode power converter, ±450 V/ ± 450 A
  Two converters tested with different loads: 1MB (bldg. 311), 7MB in series (ELENA ring)
Notes: Superconducting synchrotrons: ESCAR (Berkeley), only testbed [Frontiers Of Accelerator Technology - Proceedings Of The Joint Us-cern-japan ]
Flex = freedom with cycling

## Slide 32
Spectral analysis
Test cycle (Hall 311)
Pbar machine cycle
12.5 kHz + harmonics < 1 A
DCCT filter side lobes
50 Hz + harmonics + spurious noise (cable pickup)
6.5 kHz + harmonics < 50 mA
13, 21, 31 kHz .. < 1 nTm
6.5 kHz + harmonics < 3 nTm
  Preliminary comparison between two slightly different configurations
  Switching frequency and harmonics visible in both I(t) and B(t), negligible amplitude
  7  higher load  10  less current noise @ 100 Hz
additional coil output filter
Notes: Magnet time constant = 1.4 s
Acquisition done with 16-bit NI ADC USB @ 200 kHz

## Slide 33
Damping resistor
  Classic technique to damp high current frequencies: resistor in parallel with excitation coil
  Works in addition to damping provided by eddy currents in the yoke and vacuum chamber
I
Im
Rd
10 
SPS MBB: Rm=3.2 m, Lm=7.7 mH
18 
50 
100 
1000 
Notes: design time: fill in information missing in computer models: - material properties, mechanical tolerances ()
complement models: provide hard-to-get information
- time constant of eddy currents (difficult) - hysteresis effects (practically impossible)
magnet acceptance: go-no go decision - accept/reject/do further tests - spot checks (series tests)
fiducialization data for installation and alignment magnetic axis (transversal), longitudinal center
characterization for beam physics - input data for beam optics codes - operating information for the control room - one-off questions (e.g. measure a specific harmonic)
Examples: ELENA magnet model (dilution), magnet with refurbished coils, LHC series

## Slide 34
Conclusions

## Slide 35
Conclusions
compact synchrotrons with SC magnets in the range of 2~3 Tm beam rigidity have some historical precedents
strongly-curved extrapolations of well-established designs such as cos- look difficult; R&D needed to qualify new designs such as CCT
real-time dipole field feedback needed for high accuracy control; many applications to CERN and medical synchrotrons
new measurement solutions are being developed at CERN to decrease costs and fit beam line magnets
measurement of field harmonics for real-time correction might allow to relax field quality constraints in future magnets
switched-mode power converters provide negligible ripple