Ion transport

---
Simulating ion motion
Lines found in most Garfield++ avalanche programs:
// Load the ion mobilities.
gas->LoadIonMobility("/afs/cern.ch/user/r/rjd/GemGain/Charge/mob_Ar_Ar+");
What do they do ?
Why are they often inappropriate ?

---
The importance of ions
Avalanches produce not only electrons, but also ions – at least as many.
Detectors like Micromegas and wire chambers get their signal mostly from ion motion.
- Hence we better know the basics of ions:
  - which ions are produced in the avalache ?
  - which ions generate the signal ?
  - how fast do the ions move ?
  - are they subject to diffusion ?

---
Mobility
The velocity with which ions move is usually expressed in the form of a mobility:
  - :	drift velocity
  - K:	mobility
  - E:	electric field
Reason: for moderate fields, the mobility is approximately constant.
size +15 {v sub func D}
size +15 {v sub func D = K E }

---
Electron-impact cross sections
CO2+ and Ar+ compete.
Ne+ has higher threshold and is produced less.
[Pure gases, data from LXcat]

---
Avalanche products
Ar+ dominates in Ar-CO2, CO2+ and Ne+ in Ne-CO2

---
Avalanche products in Ar-CO2
This tells us that avalanches primarily produce Ar+ ions.
Does this mean that the signals are generated by Ar+ ions ? Maybe.
If so, we need the mobility of Ar+ in Ar-CO2 mixtures.
This can in principle be calculated with Blanc's law.

---
Avalanche products in Ne-CO2
This tells us that avalanches produce mostly CO2+ ions and a few Ne+ ions.
Does this mean that the signals are generated by CO2+ and Ne+ ions ? Maybe.
If so, we need the mobility of CO2+ and Ne+ in Ne-CO2 mixtures.
This can in principle be calculated with Blanc's law.

---
Blanc's mobility interpolation
size +15 {1 over {K sub 0 sup func mix} = f sup {func CO sub 2} over {K sub 0 sup {func CO sub 2}} + f sup func Ar over {K sub 0 sup func Ar}}
1/K0
Mobility Ar+
in pure CO2
1/K0 = E/v
reciprocal
of mobility
for E  0
Mobility Ar+
in pure Ar
Fraction of CO2
1
0
[A. Blanc, Recherches sur les mobilités des ions dans les gaz, J. Phys. Theor. Appl. 7 (1908) 825-839, 10.1051/jphystap:019080070082501]

---
Known & unknown mobilities
- Known mobilities:
  - Ar+	in Ar,	Ne+	in Ne,
  - Ar2+	in Ar,	Ne2+	in Ne,
  - CO2+ in Ar,	CO2+ in Ne.
- Not published:	why ?
  - Ar+ in CO2,	Ne+ in CO2.
- Published, but controversial:	can you guess why ?
  - CO2+ in CO2.

---
Ar+ and Ne+ mobility ≡ vD(E) / E
Noble gas ion mobilities are well known:
Ar
Ne
HW Ellis et al., At. data nucl.
data tables 17 (1976) 177-210
(compilation)
[Speaker Notes: We still need to link the ion density in the transfer gaps with the incident electron rate. This passes through the ion mobility, which is well known for pure noble gases but much less for mixtures. Neglecting the effect of the quencher, Ar+ ions need around 40 µs to traverse a gap. Ne+ ions need only 8 µs.]

---
Ion induced signals
Current induced on the pads of a TPC by ions moving from the anode wires to the pads and the field wires:
e-
Cathode
3 mm
Field + Anode
ions
3 mm
Pads

---
Signals in Ne-CO2
Ne+ in Ne comes nearest to the measurements ...
we've used it for years ...
but does it make sense ?
NA49 TPC: Ne 90 % CO2 10 %.
[Data: Rainer Renfordt]
Time taken by an ion from anode wire to field wire [µs]

---
Principal reactions involving CO2
- Ar+: charge exchange, τ ≈ 0.85 ns
  - Ar+ + CO2  Ar + CO2+
- Ne+: charge transfer in 2-steps, τ ≈ 8 ns
  - Ne+ + CO2  Ne + CO+ + O
  - CO+ + CO2  CO + CO2+
- CO2: 3-body association, 7-20 ps
  - CO2+ + 2CO2  CO2+CO2 + CO2
[For 10 % CO2, atmospheric pressure, room temperature]

---
Situating cluster ions
- Chemically bound molecules:	0.75 - 11.1 eV
  - covalent or ionic bond
- Cluster ions:	0.09 - 1.7 eV
  - bound by charge-induced dipole forces
  - constituents retain their identity
- van der Waals molecules:	0.0009 - 0.1 eV
  - bound by van der Waals forces
  - observed at low temperatures
[B.M. Smirnov, “Cluster Ions and Van Der Waals Molecules,” CRC press]

---
Binding energy of CO2 cluster ions
- Binding energy:
  - CO2+CO2:	0.60 eV	(16.2 kcal/mol)
  - CO2+2CO2:	0.26 eV	(6.0 kcal/mol)
[M. Meot-Ner and F.H. Field, J. Chem. Phys., 66 (1977) 4527]
  - CO2+CO2	0.51 eV	(11.8 ± 1.0 kcal/mol)
  - (CO2)2+CO2	0.14 eV	(3.3 ± 1.4 kcal/mol)
  - (CO2)3+CO2	0.12 eV	(2.8 ± 1.4 kcal/mol)
[S.H. Linn and C.Y. Ng, J. Chem. Phys. 75 (1981) 4921]
(Conversion: 1 kcal/mole = 0.043 eV)

---
Life cycle of CO2+(CO2)n
- CO2+CO2 has a dissociation energy of 0.6 eV far above thermal energies at 1 bar. It is a so-called long-lived cluster:
  - calculated lifetime = 5 ns
      [B.M. Smirnov, “Cluster Ions and Van Der Waals Molecules,” CRC press]
- much longer than the formation time τ = 7-20 ps via 3-body association in 10 % CO2 with Ar + CO2 as “helpers”.
- Any isolated CO2+ rapidly binds again.
CO2+(CO2)n probably lives shorter but will recombine. The cluster size n will therefore fluctuate at the ns time scale.

---
Reactions of Ar+ and Ne+
- Resonant charge exchange: (why does this matter ?)
  - Ar+ + Ar ↔ Ar + Ar+
  - k = 4.6 10-10 cm3/s, τ ≈ 100 ps (for NAr = 2.45 1019/cm3)
- Ar2+ dimer formation:
  - Ar+(2Po3/2) + 2 Ar  Ar+Ar + Ar
  - k = 2.3 ± 0.1 10-31 cm6/s, τ ≈ 7 ns
- Ne2+ dimer formation:
  - Ne+ + 2 Ne  Ne+Ne + Ne
  - k = 0.6 10-31 cm6/s

---
Ions drifting in pure Ar
In pure Ar dimers are formed:
  - Ar+(2Po3/2) + 2Ar  Ar+Ar + Ar
 (k = 2.3 ± 0.1 10-31 cm6/s, 7 ns)
  - Note: dimers move faster than ions due to Ar ↔ Ar+ resonant charge exchange
[PNB Neves et al. 10.1063/1.3497651]
Argon
p = 7.04 Torr
E/N = 30 Td
size +20 color blue {func Ar sup "+"}
size +20 color red {func Ar sub 2 sup "+"}
Contaminants
Arrival time spectrum [µs]

---
Thermal collision frequency
Mean relative velocity (µ = reduced mass):
Multiplying with the cross section σ gives the rate constant:
Combine with the number density to get collision time:
size +15 {{bar v} sub func rel ~=~ sqrt{{8 k sub func B T} over {%pi %mu}} ~approx~ 570 ~ func m/func s }
size +15 {k ~=~ %sigma {bar v} sub func rel ~approx~ 9 ~ 10 sup -10 ~ func cm sup 3 / func s}
size +15 {color red %tau ~=~ 1 over {N %sigma {bar v} sub func rel} ~=~ {k sub func B T} over {p} 1 over{%sigma {bar v} sub rel} ~=~ 1 over {p %sigma} sqrt{ {%pi %mu k sub func B T} over {8 }} ~approx~ color red {45 ` func ps}}

---
Reaction time and Rate constant: 2-body
Consider a charge transfer reaction A+ B  A B+:
  - rate  density of B molecules NB [1/cm3];
The proportionality factor k is called rate constant:
  - rate = k [cm3/s] NB [1/cm3].
  - The reaction time τ is the reciprocal of the rate.

---
3-body reactions
- Reactions resulting in a bound state, require the evacuation of excess energy & momentum
  - via internal degrees of freedom (rotation, vibration ...), or
  - via a 3-body reaction, with helpers.
The rate constant in 3-body reactions has the unit of cm6/s.
- Example 1: Ar+ + Ar + Ar  Ar+Ar + Ar
  - k = 2.3 10-31 cm6/s, assuming N ≈ 2.45 1019/cm3
  - rate = k N2 = 1.3 108/s, τ = 7 ns
- Example 2: CO2+ + CO2 + CO2  CO2+CO2 + CO2
  - k = 2.4 10-28 cm6/s
  - rate = k N2 = 1.4 1011/s, τ = 7 ps

---
mathematica
Start from the number of Ar+ and CO2+ produced in the avalanche.
Calculate evolution of Ar+, CO2+, CO2+CO2 clusters and Ar2+ dimers.

---
Evolution in Ar-CO2 (93-7)
- Initial ions:
  - Ar+	
    - CO2+
    - Ar2+
  - CO2+ 
    - CO2+.CO2

---
Blanc diagram for Ne-CO2
Like in Ar-CO2, CO2+ forms clusters in Ne-CO2.
ALICE
NA49
30 %
20 Td
20 Td
CO2+ in Ne and in CO2
Coimbra
Coimbra

---
Blanc diagram for Ar-CO2
Coxon:	pure CO2
Schultz:	1 atm
Coimbra:	0.01 atm, smaller clusters ?
NA49:	1 atm
ALICE:	1 atm, water clusters ?
20 Td
CO2+ in Ar and in CO2

---
How about alkanes ?
- Ar 90 % - C2H6 10 %, at low pressure.
- Expect Ar+ or C2H6 but ... none are seen – why ?
[André Cortez et al. 10.1088/1748-0221/8/12/P12012]
C4Hn+
C3Hn+
C5Hn+

---
C3+
Ar-CH4
C3H+
C4H3+
Ar2+
C2+
C3H2+
C3H3+
C2H+
Short lived
Long lived
C+
C2H2+
C3H4+
No further
reactions
reported
CH+
C2H3+
C3H5+
e-
15 %
CH2+
Ar
Ar+
C2H4+
C3H6+
42 %
slow
CH3+
C2H5+
C3H7+
e-
CH4+
CH4
C2H6+
C3H8+
CH5+
ion-CH4 and ion-Ar interactions
e- impact
85 %
23 %

---
Avalanche ions in Ar-CO2-CH4 (90-7-3)
 Ion	Energy [eV]	Rate [GHz]	Fraction
Ar	Ar+	15.75961	53.73	88.3 %
CO2	CO2+	13.776	3.119	5.1 %
       CO2+*	17.314	0.3758
       CO2+*	18.077	0.2218
       O+	19.07	0.1324
       CO+	19.47	0.1525
CH4	CH4+	12.65	1.959	3.2 %
       CH3+	14.25	1.115	1.8 %
 CH2+	15.2	0.07018
Magboltz 11.2bis,
E = 100 kV/cm
1 atm, 20 C

---
Reactions in Ar-CO2-CH4
- Ar+ + Ar + M	 Ar2+ + M	100 %	k = 2.3 10-31	± 10 %
- Ar+ + CO2	 CO2+ + Ar	100 %	k = 5 10-10	± 10 %
- Ar+ + CH4	 CH3+ + Ar + H	85 %	k = 1.1 10-9	± 10 %
-  CH2+ + Ar + H2	15 %
- CO2+ + CO2 + M  CO2CO2+ + M	100 %	k = 2.4 10-28
- CO2+ + CH4	 CH4+ + CO2	25 %	k = 1.1 10-9	± 10 %
-  HCO2+ + CH3	75 %
- CH3+ + CH4	 C2H5+ + H2	100 %	k = 1.1 10-9	± 10 %
- C2H5+ + CH4	 C3H7+ + H2	100 %	k = 1 10-14	± 10 %
- CH4+ + CH4	 CH5+ + CH3	100 %	k = 1.15 10-9	± 10 %
- CH5+ + CH4	 products	100 %	k = 3 10-11	± 30 %
- CH4+ + CO2	 HCO2+ + CH3	99 %	k = 9.6 10-10	± 10 %
-  CH3CO+ + OH	1 %
- CH5+ + CO2	 HCO2+ + CH4	100 %	k = 4 10-11	± 10 %
- HCO2+ + CH4	 CH5+ + CO2	100 %	k = 6 10-10	± 15 %
[B.M. Smirnov, Cluster Ions and Van Der Waals Molecules]
[V.G. Anicich and W.T. Huntress Jr., Astrophys. J. Suppl. 62 (1986) 553-672.]

---
Reaction chains
Ar2+
Ar+
CH3+
C2H5+
C3H7+
CO2+
CO2CO2+
HCO2+
CH4+
CH3CO +
CH5+
products

---
Evolution in Ar-CO2-CH4 (90-7-3)
- Initial ions:
  - Ar+	
    - CO2+,
    - CH3,
    - Ar2+
  - CO2+ 
    - CH4+
    - CO2+.CO2
    - HCO2+
  - CH4+ 
    - CH5+,
    - HCO2+
    - CH3CO+
  - CH3+ 
    - C2H5+

---
Products after 10 ns
C2H5	+	 C3H7+	slow reaction, k = 1 10-14
[Hiraoka & Kebarle https://doi.org/10.1063/1.431116]
- CH5+	 “products”	slow reaction, k = 3 10-11
- CO2 CO2+ grows to n = 3-4, reacts with CH4 to HCO2 ?
[Y. Kalkan et al. 2015 JINST 10 P07004]
Ar+	 Ar2+	drifts faster than Ar+
HCO2+	 CH5+	HCO2+ decays at k = 6 10-10
protonated
carbon dioxide
ethenium
methanium

---
Summary ions
Avalanches ionise the constituent gases, and the initial ions undergo a staggering sequence of reactions.
  - In Ar-CO2 and Ne-CO2 mixtures, the signal ions are CO2+(CO2)n clusters;
  - water forms larger clusters, further reducing the mobility;
  - pure noble gases form dimers, Ar2+, Ne2+ which are faster than Ar+ and Ne+ due to resonant charge exchange;
  - Xe forms dimers, trimers and probably bigger objects;
  - alkanes combine to form heavier molecules;
  - CH4 leads to exotic molecules.
There is room for theses in this field.

---
Overview
- Xe-CH4
  - CH5+, C2H5+, C3H7+, Xe2+
- Xe-C2H6
  - C3Hn+, C4Hn+, Xe2+
- Xe-CO2-C2H6
  - C2H4+, C2H5+, C2H6+, XeC2H6+, CO2CO2+
- Xe-CF4
  - CF3+, Xe2+
- Xe-CF4-CO2
  - Xe+, Xe2+, CF3+, CO2CO2+
- Ar-CH4
  - CH5+, C2H4+, C2H5+, C3H4+, C3H5+, C3H6+, C3H7+, Ar2+
- Ar-C2H6
  - C3Hn+, C4Hn+, Ar2+
- Ar-iC4H10
  - C4Hl0+, C8Hn+, C12Hm+
- Ar-CF4
  - CF3+
- Ar-CO2
  - CO2  CO2+, Ar2+
- Ar-CO2-CH4
  - C2H5+, CH5+, CO2  CO2+, Ar2+, HCO2+, CH3+
- Ne-CF4
  - CF3+
- Ne-CO2
  - CO2  CO2+, Ne2+
- N2
  - N2 N2+
- C2H6
  - C2H4+, C3H5+, C3H7+, C3H8+, C3H9+, C4H9+, C4H10+, C4H12+
- Xe-N2
  - Xe+, Xe2+
- iC4H10
  - C8Hn+, C12Hm+

---
Not used

---
Ionisation potentials
- Xe:	12.129843 eV
- Ar:	15.759 eV
- Ne:	21.56454 eV
CO2:	13.777 eV
- CH4:	12.61 eV
- C2H6:	11.52 eV
- C3H8:	10.94 eV
- iC4H10:	10.68 eV
- nC4H10:	10.53 eV
- C2H5:	8.4 eV
- CH5:	8.3 eV ?
- HCO2:	8.3 eV ?
- C3H7:	7.5 eV

---
Cluster mobility in Ar-CO2 & Ne-CO2
- One needs as ingredients:
  - the Blanc interpolation formula;
  - the mobility of the ions in the pure gases.
- There are measurements of:
  - Ar+ in Ar, Ne+ in Ne,
  - CO2+ in Ar and in Ne;
  - CO2+, CO2+CO2 and CO2+(CO2)n in CO2.
- But we have not found measurements of:
  - mobility of CO2+(CO2)n neither in Ar nor in Ne;
  - we can get an idea from the mass-mobility relation.

---
28 < t < 228 µs
CO2+(CO2)2
CO2 at 1 bar
CO2+CO2
CO2+(CO2)3
At 1 bar clusters are observed to emerge and then decay:
  - Until 200 µs: CO2+(CO2)n and CO+(CO2)n
  - Later: only clusters from contaminants, e.g. H2O, OH ...
CO2+
CO2+(CO2)4
CO+CO2)3
CO+(CO2)2
0.2 < t < 18 ms
[Yasumasa Ikezoe et al.,
10.1016/0146-5724(82)90034-6]

---
Experimental check
Remains showing that experiments indeed observe CO2+(CO2)n and not CO2+.
- This we do by comparing:
  - measured mobility as function of mix, with
  - mobility of CO2+, Ar+ and Ne+ from literature.

---
Avalanche products & by-products
- At very low pressure, there are traces of O+, C+ and CO+:
  - O+ + CO2  O2+ + CO	k = 1.03 ± 0.10 10-9 cm3/s
  - O2+ + CO2 + M  O2+CO2 + M	k = 0.5 ± 0.1 10-30 cm6/s
  - C+ + CO2  CO+ + CO	k ≈ 1.1 10-9 cm3/s
  - CO+ + CO2  CO + CO2+	k ≈ 1.0 10-9 cm3/s
These traces probably also exist at higher pressure, but they disappear rapidly, except the cluster and O2+.
[H.W. Ellis et al., J. Chem. Phys. 64 (1976) 3935-3941, 10.1063/1.432024]

---
He, Xe, N2
[J. A. Hornbeck, J. Phys. Chem. 56 (1952) 829–831 10.1021/j150499a003, copied from
J.H. Mitchell and K.E.W. Ridler, Proc. Roy. Soc (London) A 146 (1934) 911.]
He+
He

---
Ne+ in Ne … did it make sense ?
The avalanche produced little Ne+ to begin with;
IPNe+ > IPCO2+ : Ne+ took 8 ns to generate a CO2+;
in nearly pure Ne, there could be some Ne2+.
Instead, we obtained CO2+CO2.

---
CO2 at 1 atm: cluster (dis)appearence
Not clear why CO2+ forms at t > 50 µs.
Clusters with CO2+ ion core form by t = 100 µs.
React with contaminants τ ~190 µs.
CO2+
[Yasumasa Ikezoe et al., 10.1016/0146-5724(82)90034-6]
CO2+CO2
CO2+(CO2)2
CO2+(CO2)3
CO2+(CO2)4

---
Life cycle of CO2+(CO2)n
- CO2+CO2 has a dissociation energy of 0.6 eV far above thermal energies at 1 bar. It is a so-called long-lived cluster:
  - calculated lifetime = 5 ns
      [B.M. Smirnov, “Cluster Ions and Van Der Waals Molecules,” CRC press]
- much longer than the formation time τ = 7-20 ps via 3-body association in 10 % CO2 with Ar + CO2 as “helpers”.
- Any isolated CO2+ rapidly binds again.
CO2+(CO2)n probably lives shorter but will recombine. The cluster size n will therefore fluctuate at the ns time scale.

---
Mobility of CO2+ and CO2+(CO2)nin CO2
At atmospheric pressure, one almost inevitably measures CO2+CO2 or CO2+(CO2)n.
Cluster mobility 10-15 % smaller than ion mobility.
Two measurements exist of clusters with large n.

---
Mass-mobility in CO2
Langevin polarisation limit not appropriate for many ions.
Langevin elastic limit can fit molecular ions.
Uncertainty for cluster ion mobility.
Little resonant charge exchange effect.
10 Td

---
Mass-mobility in pure Ar and Ne
10 Td
10 Td

---
Atlas TRT signals
- Data:
  - Xe-CO2-CF4 70/10/20
  - Straw tube
  - Vw=1530 V
  - rw = 15 µm, rt = 2 mm
- Fit:
  - 1/(t+t0)
Electron signal
Ion tail
Time [ns]
[Data from Романюк Анатолий Самсонович]

---
Adding C2H6
C2H6 makes the tail steeper and the mobility larger:
Xe 70 %, CO2 10 %, CF4 20 %,
Xe 70 %, CO2 27 %, C2H6 3 %
[Plot and data from Романюк Анатолий Самсонович]

---
Ionisation potentials
C2H6:	11.52 eV
- Xe+(2P3/2):	12.129843 eV
- Xe+(2P1/2):	13.44 eV
CO2:	13.776 eV
CF4:	15.70 eV for CF3+ (main ionisation channel)
     16.2 ± 0.1 eV for CF4+

---
Reactions in Xe-CO2-CF4
Xe+ + Xe + M  Xe2+ + M	k = 2.0 ± 0.2 10-31
[A.P. Vitols and H.J. Oskam, Phys. Rev. A 8 (1973) 1860-1863.]
- CO2+ + CO2 + M  CO2CO2+ + M	k = 2.4 10-28
[B.M. Smirnov, Cluster Ions and Van Der Waals Molecules]
CO2+ + Xe	 CO2 + Xe+	k = 6.0 10-10 ± 30 %
[V.G. Anicich and W.T. Huntress Jr., Astrophys. J. Suppl. 62 (1986) 553-672.]
Note: CF4+ is not produced, (only CF3+)

---
Evolution of Xe-CO2-CF4 (70-10-20)
Initial ion mix for 100 kV/cm;
Xe+ and Xe2+ dominate from 10 ns on, Xen+ for n > 2 are not shown: rates are not known.
CF3+ is an avalanche product; with its low IP, it does not react;
CF4+ is not produced, CF4 has a high IP, is not attacked by ions.
CO2+ transfers to Xe+ and rapidly forms clusters.

---
Reactions in Xe-C2H6
- C2H4+ + C2H6	 C3H6+ + CH4	7 %	k = 5.3 10-12	± 10 %
-  C3H7+ + CH3	93 %
- C2H5+ + C2H6	 C3H7+ + CH4	14 %	k = 3.8 10-11	± 10 %
-  C4H9+ + H2	86 %
- C2H6+ + C2H6	 C3H8+ + CH4	42 %	k = 0.19 10-10	± 10 %
-  C3H9+ + CH3	58 %
[V.G. Anicich and W.T. Huntress Jr., Astrophys. J. Suppl. 62 (1986) 553-672,
A.F.V. Cortez et al. 2013 JINST 8 P07013, J.M.C. Perdigoto et al. 2017 JINST 12 P09003]
Xe+ + Xe + M  Xe2+ + M	k = 2.0 ± 0.2 10-31
[A.P. Vitols and H.J. Oskam, Phys. Rev. A 8 (1973) 1860-1863.]
- Xe+ + C2H6	 Xe + C2H4+ + H2	55 %	k = 9.2 10-10	± 20 %
-  Xe + C2H5+ + H	10 %
-  Xe + C2H6+	35 %
- Xe2+ + C2H6	 Xe C2H6+ + Xe	k = 6.8 10-10	± 20 %
[Kevin Giles et al., J. Phys. B: At. Mol. Opt. Phys. 22 (1989) 873-883.	N.G. Adams et al., J. Phys. B: At. Mol. Phys. 13 (1980) 3235-3246.]

---
Mathematica

---
Evolution of Xe-C2H6 (97-3)
Initial ion mix for 100 kV/cm;
after 10 ns, the only Xe-related ion that remains is XeC2H6+ which is made from Xe2+;
the diagram does not show the C3Hx+ and C4Hx+ ions subsequently generated by C2H4+ and C2H5+;
Xen+ for n > 2 are not shown because the rates are not known.

---
Reactions in Xe-CO2-C2H6
All reactions from Xe-C2H6 and in addition:
CO2+ + CO2	 CO2 + CO2+	100 %	k = 3.7 10-10 ± 10 %
[V.G. Anicich and W.T. Huntress Jr., Astrophys. J. Suppl. 62 (1986) 553-672.]
CO2+ + CO2 + M  CO2CO2+ + M	k = 2.4 10-28
[B.M. Smirnov, Cluster Ions and Van Der Waals Molecules]
CO2+ + Xe	 CO2 + Xe+	100 %	k = 6.0 10-10 ± 30 %
[V.G. Anicich and W.T. Huntress Jr., Astrophys. J. Suppl. 62 (1986) 553-672.]
- CO2+ + C2H6	 CO2 + C2H5+ + H	32 ± 2%	k = 7.8 ± 1.2 10-10
-  CO2 + C2H4+ + H2	68 ± 2 %
[MasaharuTsuji, 10.1016/0168-1176(94)03996-8]

---
Evolution of Xe-CO2-C2H6 (70-27-3)
Initial ion mix for 100 kV/cm;
CO2+ rapidly forms clusters due to the large CO2 fraction;
CO2 does not affect the dominance of CxHy over Xe.
Xen+ for n > 2 are not shown because rates are not known.

---
Xe2+ - Xe3+
We have not found the rate comstant for Xe3+ production in the literature, but H. Helm has measured [293 K]:
where kf is the rate coefficient for the transformation of Xe2+ to Xe3+, and kr the rate for the reverse reaction.
Given that [Xe3+] / [Xe2+] = 3.6 ± 0.6 10-19 N, the ratio of concentrations is 9.7 ± 1.6 at 293 K, atmospheric pressure and zero field.
size +10 { K sub func e = {[func Xe sub 2 sup "+"] [func Xe][func Xe]} over {[func Xe sub 3 sup "+"] [func Xe]} = { k sub func r } over { k sub func f } = 2.8 +- 0.5 ~ 10 sup 18 }
[H. Helm, 10.1103/PhysRevA.14.680]

---
What have we learned ?
After 10 ns, the ion with the lowest IP remains:
  - In Xe without C2H6, this is Xe+ and related;
  - if C2H6 is present, CxHy+ dominates.
In the long run, Xe+ is likely to form Xen+

---
Mobility and stability of small Xen+
- Mobilities are known for the smallest Xe clusters;
- these are remarkably stable;
- ∆H for n ≥ 5 clusters is nearly constant at ~0.1 eV.
[Kenzo Hiraoka et al. 10.1063/1.457751, M. Amarouche et al. 10.1063/1.454267]
Xen+	Mass	µ (E = 0, T=300 K)	∆H n-1  n	References
	[Da]	[cm2/V.s]	[eV]
Xe+	131.293	0.55	-	Helm, Viehland-Mason
Xe2+	262.586	0.61	1.05 (5 %)	NIST, Helm 1976
Xe3+	393.879	0.57	0.29 (5 %)	NIST, Helm 1976
Xe4+	525.172	?	0.26 (3 %)	NIST, Hiraoka
Xe5+	656.465	?	0.11 (5 %)	NIST, Hiraoka

---
Mobility of ions in Xe
E/N = 10 Td, extrapolated from higher E/N where needed (Xe+).
Polarisation limit	assuming αXe = 4.01 D.
Xe2+ and Xe3+ are below the polarisation limit.
[From the H.W. Ellis et al. compilations except Xe+ and Xe2+, which are from P.N.B. Neves, 10.1063/1.3497651]

---
Large Xen clusters
- Much larger clusters, with a size of 104, have been observed.
- They are produced by “supersonic adiabatic expansion through a nozzle.”

---
Avalanche ions in Ar-CO2-CH4 (93-4-3)
 Ion	Energy [eV]	Rate [GHz]	Fraction
Ar	Ar+	15.75961	56.65	91.1 %
CO2	CO2+	13.776	1.815	2.9 %
       CO2+*	17.314	0.2250
       CO2+*	18.077	0.1261
       O+	19.07	0.08213
       CO+	19.47	0.08430
CH4	CH4+	12.65	1.996	3.2 %
       CH3+	14.25	1.138	1.8 %
 CH2+	15.2	0.07641
Magboltz 11.2bis,
E = 100 kV/cm
1 atm, 20 C

---
Evolution in Ar-CO2-CH4 (93-4-3)
- Initial ions:
  - Ar+	
    - CO2+,
    - CH3,
    - Ar2+
  - CO2+ 
    - CH4+
    - CO2+.CO2
    - HCO2+
  - CH4+ 
    - CH5+,
    - HCO2+
    - CH3CO+
  - CH3+ 
    - C2H5+