## Slide 1: HNFS setup improvements
SM, 21/4/2020

## Slide 2: Where to improve
Cameras (SNR, resolution)
Scintillator & optics (resolution, light yield)
Sample (SNR)

## Slide 3: Camera
In October 2018: e- well occupied at 426/2**12 = 10% for 50 ms exposure time.

## Slide 4: Camera – signal simulation
  Single particle pattern
  z = 300 mm, sigmas = 8 um, 12 keV (0.1 nm wl)
  Overlap 5x104 patterns on magnified (x20) pixel matrix
  Normalise amplitude

## Slide 5: Camera – signal simulation
  Add pedestal according to 10 % full well occupation (and QE)
  Estimate signal amplitude wrt pedestal at 2.7%:
    Oct 18 image:
  Add Shot Noise (Poisson distributed) and Readout Noise (negligible) according to camera specs

## Slide 6: Camera – signal simulation
Measured
Simulated

## Slide 7: Camera – signal simulation
Measured
Simulated

## Slide 8: Camera - comparison
Use procedure to compare Basler Scout and Basler Ace based on specs (QE, full well capacity, pixel size, sensor size, dark noise) with equivalent photon flux.
difference due to smaller pixel size / larger sensor
FW occupation:
10% B. scout
4% B. Ace

## Slide 9: Camera - comparison
Different camera does not help much if SNR is not increased. Here comparison of present setup (blue) with Basler ace and 2x signal (orange)

## Slide 10: Camera - summary
SNR is limited by shot noise, readout noise negligible.
Imperative to increase signal, better camera does not help
Smaller pixels / larger sensor size improves sp. resolution
…but what about YAG + optics resolution?

## Slide 11: YAG + optics
Considerations based on paper: T. Martin et al.“Recent developments in X-ray imaging with micrometer spatial resolution”, J. Synchrotron Rad. 13, 180–194 (2006)

## Slide 12: YAG
Spatial resolution in scintillator limited by:
  Photoelectric effect (absorption of X ray photon, isotropic re-emission of K band electron that causes fluorescence)
  Scattering of X rays photons (Rayleight, compton)
Ph. effect is dominant up to approx. 100 keV

## Slide 13: Optics
However:
“…high spatial resolution is [thus] not limited by the scintillator but by the resolving power of the optics, i.e. in particular of the microscope objective.”

## Slide 14: Optics
Spatial resolution affected by
  depth of field (dz)
  diffraction (l)
  spherical aberrations (t)

## Slide 15: Optics
Thicker YAG does not improve light collection: limited by depth of field (in our case a few um).
For 100 um thick YAG, 0.4 NA objective, resolution is approx. 3 um
Thinner YAG => less spherical aberrations.

## Slide 16: Optics
Simulate resolution by filtering (in freq. space) signal with a 2D triangular function with 50% modulation at q_typ
I
q
0
1
0.5
qtyp = 2p/Res

## Slide 17: Optics
From Introduction to Fourier optics (Goodman)

## Slide 18: Optics
Present system, infinite resolution

## Slide 19: Optics
Present system, inf vs 3 um resolution

## Slide 20: Optics
Present system, inf vs 2 um resolution. Qualitatively it looks similar to experimental data

## Slide 21: Optics
Present system, inf vs 1 um resolution

## Slide 22: Optics
At 2 um resolution, changing camera (smaller pix) does not have a major impact

## Slide 23: Optics
At 1 um resolution, changing camera starts to make sense

## Slide 24: Conclusion
Ways to improve quality of data:
  Increase signal: try to increase thickness of cuvette / capillary
  Increase resolution: thinner YAG (50 um? less?)
  If resolution is increased: change camera (Basler Ace)
2x signal
2x resolution
Basler Ace camera