## Slide 2
PicoTDCv2: Pico-second TDC for HEP
Moritz Horstmann, Jorgen Christiansen, Samuele Altruda, Gill Lumer-Klabbers, Jeffrey Prinzie (KU Leuven)
CERN/EP-ESE

## Slide 3
3
PicoTDC architecture
Samuele Altruda
25/05/2020

## Slide 4
4
PicoTDC Architecture
Samuele Altruda
64 channels, 3ps or 12ps time binning, 200us dynamic range
Trigger buffer: 512
Trigger buffer: 512
Trigger buffer: 512
Buffer: 512
Counters
TDC channel
1.28GHz
320MHz
160MHz
80MHz
40MHz
40MHz
Ext ref.
1.28GHz …
x16
Readout buffer: 512
8 bit 40,80,160,320M
x8
40MHz
320MHz
x8
x8
Buffer: 512
TDC channel
x16
Readout buffer: 512
8 bit 40,80,160,320M
x8
Trigger matching
Trigger matching
PLL
Trigger buffer: 512
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
D
Interpolation decoder
Hit register
x16
x16
x16
Control,
Programming,
Monitoring
Hits
x64
Leading
Trailing
TOT
Trigger,
Resets
I2C
Readout
1/4 Ports
Data rate
320Mbps
-10Gbps
(5MHit/s
per channel)
Trigger interface
Sync
Ch 0
Pulse out
Pulse generator
x16
25/05/2020
Notes: 64 inputs in 4 groups
1.28Ghz main frequency for running the delay locked loop 781ps, max frequency of hit acceptance.
Each channel has hit registers and interpolation logic decoder to store timing value.
Each channel can store 4 measurement, in a FIFO buffer derandomizer for going to a frequency of 320MHz. Second part driven at 320Mhz, ¼ data rate, buffering needed, 512 word(32 bit) buffer also as trigger delay buffer.
Can keep 16x512+512 measurement per channel group = 8704 per channel group
Bottleneck in readout buffer, 1hit/25ns using 1 channel at 320Mhz, 4 clock cycle for 1 measurement
Bandwidth:
Min 320Mbits/s (8bits/25ns(40MHz) = 1 bits/3.125ns = 320Mbits/s)( ~0.15 Mhits/s per channel) (320MBits/32 bits/64ch)
Max 10Gbits/s (8bits/3.125ns(320MHz) = 1 bits/0.39ns = 256Mbits/s*4channel(max)= 10Gbits)( ( ~4 Mhits/s per channel) (10GBits/32 bits/64ch)
Trigger interface, 4 buffer , 1 buffer per channel for triggering.
I2C configuration interface
4

## Slide 5
5
Two Stage Time Interpolation
Samuele Altruda
1.28GHz
12.2ps delay
(Coarse time)
3.05ps delay
(Fine time)
Phase
Adjustment
32 steps each with 0.6ps resolution
25/05/2020
Notes: Unexpected mismatch between flipflops, which leads to leading trailing misdetection, problem in fine time mode acquisition.
Mismatch also between channels, visible in top part acquisition, causing some issues at the moment.
5

## Slide 6
6
Samuele Altruda
PLL
Top hit receivers
Bottom hit receivers
DLL
Logic
25/05/2020
Notes: One DLL line, top bottom part with different fine adjustment buffers
6

## Slide 7
7
Samuele Altruda
PicoTDC on Test Card first version
25/05/2020

## Slide 8
8
Samuele Altruda
PicoTDC on Generic Package
25/05/2020

## Slide 9
9
PicoTDC testing system
Samuele Altruda
25/05/2020

## Slide 10
10
Instrumentation for testing
Keysight Pulse generator
Power supply
Trombone
oscilloscope
PicoTDC carrier board
Samuele Altruda
07/05/2020
FPGA Board VC707
Silicon Lab clock generator
Samtec CCH-J-02 (2 Pin)
Samtec CCH-J-04 (4 Pin)
Female connector​
(Supplied by SAMTECH directly)
Adapter cable to SMA:
Samtec RF25S-01BJ1-SCC21-0200
(Supplied by SAMTECH directly)
Notes: Mention different instruments and purpose of its use
10

## Slide 11
PCB mezzanine test board
11
Samuele Altruda
Differential signal connection through SAMTEC connector
25/05/2020
Notes: Present briefly what we have
11

## Slide 12
picoTDC
1 Gbit Ethernet
Pulse generator
SMA input
Common Hit signal
IPBUS
Firmware
SDA / SCL
Reset
FIFO
Readout
Sync
TDC data
FPGA firmware
Buffer
PicoTDC I2C
Writing FIFO
Reading FIFO
40 Mhz clock
Timing
Sequencer
Trigger
Hit
ODelay
Trigger
SMA input
FPGA Board (VC707 Xilinx commercial board)
12
Samuele Altruda
25/05/2020
Notes: IPBUS communication to computer through Ethernet.
Readout interface with FIFO to IPBUS block.
Trigger tester with fine delay generator inside FPGA.
I2C communication for picotdc c
Configuration register programming.
12

## Slide 13
FMC2
Reset
picoTDC
FPGA
GND
+1.20v
I2C
Read out Sync
Read out 1
Clk C2M_1
Clk C2M_2
Differential signal connection through SMA
Ch 0 -63
VCC3V3 VCC1V2 3A
Power module
1DIF
1DIF
8DIF
1DIF
1DIF
17 DIF LPC
Read out 2
Read out 3
Read out 4
8DIF
8DIF
8DIF
8DIF
24 DIF HPC
FMC1
GND
64DIF + 6 DIF
DIF HPC + DIF LPC
ADAPTOR CARD WITH SMAs CONNECTORS
1DIF
(1 USER CLK)
(USER CLK)
Trigger
40 Mhz Clk
1DIF
1DIF
2 SIG
EID Reset
BX Reset
2 SIG
I2C Board
I2C expander
Mux
Mux
Mux
Reset
EID Reset
BX Reset
40 Mhz Clk
Trigger
Picotdc pulse generator
3.3V I2C level
PLL HF
PLL CTRL
DLL CTRL
I2C
I2C expander
Mux
New Mezzanine Card
13
Samuele Altruda
VCC3V3 VCC1V2 3A
25/05/2020

## Slide 14
New Mezzanine Card
All SMA connectors (NO Samtec CCH-J-02 , or Samtec CCH-J-04)
Power lines available on FMC1 directly from FMC2
Future version with final package will include some connections between FMC1 and FMC2, a bridge between the FPGA and the Adaptor card. Currently users can have some generic output pin availability from XADC Header or FMC2 from the VC707 evaluation board.
Availability:
  With chip in generic package: ~June
  With chip in final package: ~ autumn (depending on when final package available)
Share point link for the new mezzanine card and adaptor card PCB and schematic:
http://cern.ch/PicoTDC
14
Samuele Altruda
25/05/2020

## Slide 15
PicoTDC test routines
15
Samuele Altruda
25/05/2020

## Slide 16
PicoTDC: Delay sweep test
Sweep test is performed to measure effective single shot RMS resolution covering linearities and jitter
The silicon lab board (Si5341-D-EVB with sub-100 fs rms phase jitter) is used to perform sweep test providing the clock signal and the hit signal which is then delayed by the trombone, or by the board itself.
The limitation of the silicon lab board is the fact that it is able only to generate clock, which are then shifted in phase of 0.3ps to obtain a hit signal delayed in time.
The trombone is a very precise and repeatable programmable delay line, but with maximum 600ps delay.
The Keysight generator can overcome the limitations of the trombone and the silicon lab board. The Keysight generator is more flexible in generating delay and pulses, but with higher internal non-linearity and jitter.
16
Samuele Altruda
25/05/2020
Notes: Sweep test is performed to quantify the linearity of picoTDC.
To perform sweep test we use a dual channel pattern generator to provide the clock signal and the hit signal which is then delayed by the trombone.
The trombone is a very precise and repeatable, mechanical, programmable delay line.
We provide a hit signal synchronized and delayed with the clock source. We can provide than a measurement of the hit time.
16

## Slide 17
PicoTDC: Delay sweep test
The pulse is generated in sync with the clock and is going to the hit input of the picoTDC. There is a delay between the 2 signals, which goes with steps of 0.3 ps.
PicoTDC main DLL clock cycle is 1.28 Ghz, 780ps, the course time counter clock cycle is 40 Mhz, 25ns, so usually delay sweep measurements are provided in this 2 forms to check any non-linearity in both elements.
500 hits per delay step are acquired, calculating from them: Standard deviation (std in red), minimum (min in green), maximum (max in green)
INL or RMS_resolution or also called single shot RMS resolution, comprise:
  Jitter
  Quantization
  INL (standard INL calculated in code density test, not including jitter and quantization).
This value is calculated as RMS value through all the measurements taken without any averaging.
17
Samuele Altruda
25/05/2020
Notes: Sweep test is performed to quantify the linearity of picoTDC.
To perform sweep test we use a dual channel pattern generator to provide the clock signal and the hit signal which is then delayed by the trombone.
The trombone is a very precise and repeatable, mechanical, programmable delay line.
We provide a hit signal synchronized and delayed with the clock source. We can provide than a measurement of the hit time.
17

## Slide 18
PicoTDC: Code density test
18
Code density test is performed to measure the effective bin size. The code density test is in principle made with uncorrelated random hit generator (or systematic sweep with silicon labs card)
The system can provide the same amount of expected pulses in each bin.
The resulting histogram is the result of the measured amount of hits collected per each bin. This test could be done with bin size of 3 ps (fine time), 12 ps (coarse time).
The measurements don’t include jitter and quantization, INL and DNL don’t include jitter and quantization too.
Samuele Altruda
25/05/2020
Notes: Code density test is performed to measure the effective bin size. The measurement doesn’t include jitter and quantization.
To perform code density test we use an old RC based generator, which provides us random hits.
As a random source we expect the same amount of hits in each bin.
18

## Slide 19
PicoTDC delay sweep test
19
Samuele Altruda
25/05/2020

## Slide 20
PicoTDC: Delay sweep test
20
Ch 1, not adjusted, coarse mode, bin 12ps, RMS_resolution = 4,165ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/4Jc4wv68w8V4Ofi
* Link for interactive web plot
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/GvY7rTUIu6CgGBx
* Link for interactive web plot
25/05/2020
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
20

## Slide 21
PicoTDC: Delay sweep test
21
Ch 1, not adjusted, fine mode, bin 3ps, RMS_resolution = 3,256ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/zbxF6ikAMAB0BAu
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/fpqJKaSbkbPG0vX
* Link for interactive web plot
25/05/2020
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
21

## Slide 22
 |  |  | Sweep Test (PicoTDCV2) | Sweep Test (PicoTDCV1)
 | Ch | adjusted | RMS_resolution | RMS_resolution
Coarse time | 1 | NO | 4,165ps | X
 | 31 | NO | 4,066ps | 4,129ps
 | 35 | NO | 4,085ps | X
 | 48 | NO | 4,015ps | X
Fine time | 1 | NO | 3,256ps | X
 | 31 | NO | 2,904ps | 3,416ps
 | 35 | NO | 2,822ps | X
 | 48 | NO | 2,707ps | X
PicoTDC: performances summary
22
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
25/05/2020
Notes: Coarse time we expect to have the same result with common adjustment correction, but with new chip version with less mismatch
Fine time we expect to have a final result with common adjustment betwenn adjusted and not adjusted
22

## Slide 23
23
PicoTDC code density test comparison old/new PicoTDC
Samuele Altruda
25/05/2020

## Slide 24
PicoTDC: Code density test comparison
Ch 31, not adjusted, coarse mode, bin 12ps
24
Samuele Altruda
PicoTDC v1
PicoTDC v2
25/05/2020
Notes: As a random source we expect the same amount of hits in each bin, the deviation from a flat histogram means we need to adjust the bin size.
Mention DNL and INL
24

## Slide 25
Ch 31, not adjusted, fine mode, bin 3ps
PicoTDC: Code density test comparison
25
Samuele Altruda
PicoTDC v1
PicoTDC v2
25/05/2020
Notes: As a random source we expect the same amount of hits in each bin, the deviation from a flat histogram means we need to adjust the bin size.
Not visible but remind that we have mismatch which in fine mode create misdetection of trailing and leading edge.
Mention DNL and INL
25

## Slide 26
PicoTDC Focus on
channel 48
26
Samuele Altruda
25/05/2020

## Slide 27
PicoTDC: Code density test
Ch 48, not adjusted, coarse mode, bin 12ps, RMS_dnl= 2,478ps, RMS_inl= 2,192ps
27
Samuele Altruda
Ch 48, adjusted, coarse mode, bin 12ps, RMS_dnl= 0,356ps, RMS_inl= 0,454ps
Adjustment for Ch 48 is performed for this channel alone
Notes: As a random source we expect the same amount of hits in each bin, the deviation from a flat histogram means we need to adjust the bin size.
Mention DNL and INL
27

## Slide 28
PicoTDC: Delay sweep test
28
Ch 48, adjusted, coarse mode, bin 12ps, RMS_resolution = 3,662ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/wH7CpS2kFXkOHOs
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/yeXL80r9OfnLc9U
* Link for interactive web plot
25/05/2020
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
28

## Slide 29
Ch 48, not adjusted, fine mode, bin 3ps, RMS_dnl= 2,141ps, RMS_inl= 2,559ps
PicoTDC: Code density test
29
Samuele Altruda
Ch 48, adjusted, fine mode, bin 3ps, RMS_dnl= 0,323ps, RMS_inl= 0,322ps
Adjustment for Ch 35 is performed for this channel alone
Notes: As a random source we expect the same amount of hits in each bin, the deviation from a flat histogram means we need to adjust the bin size.
Not visible but remind that we have mismatch which in fine mode create misdetection of trailing and leading edge.
Mention DNL and INL
29

## Slide 30
PicoTDC: Delay sweep test
30
Ch 48, adjusted, fine mode, bin 3ps, RMS_resolution = 1,257ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/3OFMDr2Mz4brwMA
https://cernbox.cern.ch/index.php/s/bL1LSva2binJkb7
* Link for interactive web plot
25/05/2020
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
30

## Slide 31
PicoTDC Common adjustment
31
Samuele Altruda
25/05/2020

## Slide 32
Ch 1-31, coarse mode, bin 12ps, RMS_dnl_notadj= 2,074ps, RMS_dnl_adj= 1,485ps
32
Samuele Altruda
NOT ADJUSTED
COMMON ADJUSTMENT
PicoTDC: Common adjustments
25/05/2020
Notes: As a random source we expect the same amount of hits in each bin, the deviation from a flat histogram means we need to adjust the bin size.
Mention DNL and INL
32

## Slide 33
PicoTDC: Common adjustments
33
Ch 1-31, coarse mode, bin 12ps,
Samuele Altruda
NOT ADJUSTED
COMMON ADJUSTMENT
Notes: X axis delay introduced by the trombone, Y axis INL in LSB of measurement.
Specify min max values.
INL in rms mode, not averaged, but calculated through all the measurement points, around 16000 per delay step
33

## Slide 34
Ch 1-31, fine mode, bin 3ps, RMS_dnl_notadj= 2,141ps, RMS_dnl_adj= 1,607ps
34
Samuele Altruda
NOT ADJUSTED
COMMON ADJUSTMENT
PicoTDC: Common adjustments
Notes: As a random source we expect the same amount of hits in each bin, the deviation from a flat histogram means we need to adjust the bin size.
Mention DNL and INL
34

## Slide 35
35
Ch 1-31, fine mode, bin 3ps,
Samuele Altruda
NOT ADJUSTED
COMMON ADJUSTMENT
PicoTDC: Common adjustments
Notes: X axis delay introduced by the trombone, Y axis INL in LSB of measurement.
Specify min max values.
INL in rms mode, not averaged, but calculated through all the measurement points, around 16000 per delay step
35

## Slide 36
PicoTDC TOT measurements
36
Samuele Altruda
25/05/2020

## Slide 37
Ch 1, fine mode, bin 3ps, LSB TOT 24.4 ps Clock period 20Mhz
PicoTDC: TOT test
37
Samuele Altruda
In TOT mode, both leading edge and TOT have programmable resolution.
25/05/2020
Notes: Adjustment worked very well.
Mention DNL and INL
37

## Slide 38
PicoTDC Crosstalk test
38
Samuele Altruda
25/05/2020

## Slide 39
Common Hit signal
FMC1
picoTDC
FPGA
Ch 1 Hit signal
Ch 0
Ch 2-30
Ch 32-34
Ch 36-47
Ch 49-63
Test scheme simplified
Common Hit signal delay sweep from 0ps to 60000ps
Ch 1 Hit signal at fixed delay
PicoTDC: Crosstalk test
Cross talk test is performed to quantify the noise introduced by the others channels against one. Worst case in exam: one against all.
39
Samuele Altruda
25/05/2020
Notes: Cross talk test is performed to quantify the noise introduced by the others channels against one. Worst case in exam: one against all.
Common Hit signal delay sweep from 0ps to 60000ps. Ch 31 Hit signal at fixed delay
39

## Slide 40
Ch 1, coarse mode, bin 12ps, 1 channel vs all
PicoTDC: Crosstalk test
40
Samuele Altruda
Ch 1, fine mode, bin 3ps, 1 channel vs all
Ch 1, fine mode, bin 3ps, 1 channel vs 1 channel adjacent (ch 2)
Notes: As a random source we expect the same amount of hits in each bin, the deviation from a flat histogram means we need to adjust the bin size.
Not visible but remind that we have mismatch which in fine mode create misdetection of trailing and leading edge.
Mention DNL and INL
40

## Slide 41
PicoTDC Cable delay test
41
Samuele Altruda
25/05/2020

## Slide 42
picoTDC
Ch 31 Hit signal
Ch 0
Ch 2-30
Ch 32-34
Ch 36-47
Ch 49-63
Test scheme simplified
Hit generator outputs a delay sweep hits.
Hits go through different cables with different lengths to channel 31 and 48. Our TDC will measure the time of arrival of the two hits.
PicoTDC: Cable delay test
Cable delay test is performed to quantify the linearity of the system over multiple clock cycle, around 200ms
42
Samuele Altruda
Ch 48 Hit signal
Hit Generator
25/05/2020
Notes: 42

## Slide 43
PicoTDC: Cable delay test
43
Samuele Altruda
Ch 31, Ch 48, not adjusted , fine mode
Bin 9966 ~30 ns
Cable length difference
Ch 31, Ch 48, not adjusted , coarse mode
Bin 2492 ~30 ns
Cable length difference
Ch 31, Ch 48, adjusted each half , fine mode
Bin 9968 ~30 ns
Cable length difference
25/05/2020
Notes: 43

## Slide 44
PicoTDC Integrated Pulse Generator
44
Samuele Altruda
25/05/2020

## Slide 45
The PicoTDC contains a pulse generator for system testing purposes, able to generate a differential output signal.
The output can either be a 1.28GHz clock with selectable phase shift or a configurable pulse.This works internally with a selection of one of the 256 clock phases of the upper half of the timing macro.
In this example, the pulse generator was configured to use clock phase 192 (relation to the 40MHz clock), the time of the rising edge is set to 1, the falling edge to 4 and the repetition time to 7.
PicoTDC: Integrate Pulse Generator test
45
Samuele Altruda
25/05/2020
Notes: Adjustment worked very well.
Mention DNL and INL
45

## Slide 46
Ch 1, phase 3, rising 513(LSB), falling 1027(LSB), reload 1028(LSB)
PicoTDC: Integrate Pulse Generator test
46
Samuele Altruda
Period= reload(LSB)*0.78125ns=803 ns
Measurement done with limited bandwidth and sampling frequency scope.
More measurements to be done with high speed scope back to CERN.
25/05/2020
Notes: Adjustment worked very well.
Mention DNL and INL
46

## Slide 47
PicoTDC new version
47
Samuele Altruda
25/05/2020

## Slide 48
48
Samuele Altruda
PicoTDC v2
New features available on version 2:
Configurable pulse generator on chip
Trigger from channel 1
Production schedule:
PicoTDCv2 at CERN bonded on old test board available and tested
Functional and timing performance test of PicoTDCv2 bonded on old PCB ongoing
Functional and timing performance test of PicoTDCv2 on generic package in April - May 2020
Functional and timing performance test of PicoTDCv2 in final package in July - September 2020
Share point link for info and material: http://cern.ch/PicoTDC
For more info and request access to the material write to:
Jorgen Christiansen <Jorgen.Christiansen@cern.ch>
Samuele Altruda <samuele.altruda@cern.ch>
25/05/2020
Notes: 48

## Slide 50
PicoTDC: Delay sweep test
50
Ch 31, not adjusted, coarse mode, bin 12ps, RMS_resolution = 4,066ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/b2bdVf2JDZU1wPc
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/MVsmlaZT0ztRpNQ
* Link for interactive web plot
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
50

## Slide 51
PicoTDC: Delay sweep test
51
Ch 35, not adjusted, coarse mode, bin 12ps, RMS_resolution = 4,085ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/Od3gfXKGbDd4gGA
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/9bvaWodjuWe0puq
* Link for interactive web plot
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
51

## Slide 52
PicoTDC: Delay sweep test
52
Ch 48, not adjusted, coarse mode, bin 12ps, RMS_resolution = 4,015ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/u2E7sgOQ92txr6i
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/ft5glBUL7PSWP4k
* Link for interactive web plot
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
52

## Slide 53
PicoTDC: Delay sweep test
53
Ch 31, not adjusted, fine mode, bin 3ps, RMS_resolution = 2,904ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/fskC5r9l2V7ZMVq
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/QtcVm2R0UYcXNNa
* Link for interactive web plot
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
53

## Slide 54
PicoTDC: Delay sweep test
54
Ch 35, not adjusted, fine mode, bin 3ps, RMS_resolution = 2,822ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/kFmKMVg7X0JlZzC
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/ZdxoMEgEyYxAXc4
* Link for interactive web plot
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
54

## Slide 55
PicoTDC: Delay sweep test
55
Ch 48, not adjusted, fine mode, bin 3ps, RMS_resolution = 2,707ps
Samuele Altruda
*RMS_resolution includes jitter, quantization and INL
https://cernbox.cern.ch/index.php/s/j1Q6Kxats3H8L2L
* Link for interactive web plot
https://cernbox.cern.ch/index.php/s/g3sH0f4zMbw4KOf
* Link for interactive web plot
Notes: CH, adjustment, bin size
The measurement of the hit time includes also jitter from picotdc and pattern generator.
X axis delay introduced by the trombone, Y axis LSB of measurement.
Right zoom look bin size, min max and std increasing at bin intersection, while reducing in the middle of a bin.
Black line is the interpolation line, so the expected measurement line
Mention adjustment, mode, INL
55

## Slide 56: Electrical Interfaces
Hits: Differential (LVDS “compatible”, common mode from 0.2V to 1.2V)
  Highest speed (resolution) @ ~800mV common mode
Time reference: 40MHz differential
  Low jitter reference critical for high time resolution
Trigger/Event-Rst/BX-Rst/Reset: Sync Yes/No
Control/monitoring: I2C at CMOS 1.2V-levels
Readout: 4 readout ports of 8 differential signals
  Common mode 0.6V, programmable current 1-5mA
  Compatible with LpGBT and FPGAs
Packaging: 400 BGA (1mm pitch)

## Slide 57: Config / Control / Status Interface
I2C Interface, up to 1MBit/s
1.2V CMOS Levels
348 Bytes configuration / control
  Additional 322 bytes delay adjust
300 Bytes status

## Slide 58: Readout
1 or 4 differential readout ports with 8 bits
  40 - 320MHz
  Bandwidth:
    Min 320Mbits/s ( ~0.15 Mhits/s per channel)
    Max 10Gbits/s ( ~4 Mhits/s per channel)
Readout data: 32 bit words
  TDC data, headers, trailers etc.

## Slide 59: 32 Bit Frames
Field A (13)
Type (4)=100?
Field B (13)
00
Event headers (up to two)
Chan-Grp-ID (2)
Type (4)=1010
Event trailers
TDC data (31)
Type (1)=0
TDC measurement
59
Hit Count (13)
00
Possible fields: Event ID, Bx ID, Natural ID
Type (4)=1111
0x0000000 (26)
Channel group separator (for single readout port)
Event ID (13)
0x0D0D0D0 (28)
Idle frame
Type (4)=1101

## Slide 60: Absolute TDC data
Relative to Trigger
FULL TDC data, DEFAULT FORMAT
Channel (4)
Coarse cnt (13)
Med. cnt (5)
DLL int (6)
Res int (2)
Edge (1)
Type (1)
Triggered with relative time: Same as absolute
Channel (4)
Coarse cnt (13)
Med. cnt (5)
DLL int (6)
Res int (2)
Edge (1)
Type (1)
B: Triggered with relative leading and TOT: Same as absolute Lead. + TOT
Channel (4)
Leading (16)
TOT(11)
Type (1)
Channel (4)
Leading (19)
TOT(8)
Type (1)

## Slide 61: Leading + TOT
Packet Type: 1bit
Channel ID: 4 bits, for single port readout +2 bit group separator
Leading: 16/19 bits
  Large dynamic range
    16bit 3ps resolution: 200ns
    19bit 3ps resolution: 1600ns
  Programmable part of full 25bits leading TDC
  (Relative to trigger to be useable)
TOT (Relative to leading): 11/8 bits
  Short dynamic range:
    8bit 3ps resolution: 780ps
    11bit 3ps resolution: 6.1ns
  Programmable part of full 25bits TOT difference
    TOT assumed to be used for offline time-walk correction of leading.
Alternative: Readout of Individual Leading and Trailing edges with full range/resolution
  2x readout bandwidth
Channel (4)
Leading (16)
TOT(11)
Type (1)
Channel (4)
Leading (19)
TOT(8)
Type (1)

## Slide 62: Estimated Power Consumption
Highly dependent on hit rate, values based on 1 MHz per channel
High resolution, 64 channels: 1300mW
High resolution, 32 channels: 900mW
Low Resolution, 64 channels: 850mW
Low Resolution, 32 channels: 550mW