## Slide 1: Updated calibration for CLICpix2+planar assemblies
Vertex and tracking meeting
Morag Williams
29/5/2020

## Slide 2: Contents
Introduction
1) Threshold calibration for odd and even columns
2) Updated e/mV and capacitance calculations
3) Updated ToT calibration
4) Calibrated test-beam results

## Slide 3: introduction
Discussed last meeting (5/5/2020) the differences in capacitance seen between odd and even columns of CLICpix2+planar assemblies during calibration.
Surrogate function parameters of the ToT calibration curves had a clear systematic differences.

## Slide 4: introduction
Discussed last meeting (5/5/2020) the differences in capacitance seen between odd and even columns of CLICpix2+planar assemblies during calibration.
Surrogate function parameters of the ToT calibration curves had a clear systematic differences.
Difference due to the mirrored column design of the ASIC.
Suggestion: try pixel-by-pixel threshold calibration
Top metal layer of CLICpix2 ASIC

## Slide 5: 1) Threshold calibration for odd and even columns
Aim to convert threshold units from THL DAC to electrons.
Original threshold calibration procedure:
- X-ray fluorescence measurements using copper and iron targets are used
- threshold scan can locate the threshold that corresponds to the energy of the fluorescence peak
- graph of energy vs. threshold provides the e/THL DAC calibration constant for the assembly
Updates to this procedure:
Initial idea was pixel-by-pixel determination of the fluorescence peak position of each target
 unfortunately not enough statistics in the data set
Good compromise is to determine the threshold of the fluorescence peaks separately for odd and even columns

## Slide 6
 | Even | Odd
Copper | 1393 | 1399
 | ± 17 THL DAC | ± 16 THL DAC
Iron | 1368 | 1371
 | ± 14 THL DAC | ± 13 THL DAC
Top row: histogram of counts vs. threshold for copper and iron target.
Note that only single pixel clusters are taken into account in this analysis.
Gaussian fit to high-threshold range of peak used to find the corresponding threshold.

## Slide 7
 | Even | Odd
Copper | 1393 | 1399
 | ± 17 THL DAC | ± 16 THL DAC
Iron | 1368 | 1371
 | ± 14 THL DAC | ± 13 THL DAC
Gradient | 18.9 | 19.6
 | ± 0.8 e/THL DAC | ± 0.7 e/THL DAC
Gaussian fit to high-threshold range of peak used to find the corresponding threshold.
Bottom row: linear fit to the points corresponding to the target materials and the baseline.
Gradient = e/THL DAC calibration factor.

## Slide 8: 2) TOT calibration: e/vm and capacitance
For CLICpix2 ToT calibration, utilise results of the threshold calibration, therefore this also needs to be updated.
Threshold calibration used to calculate the capacitance of the test pulse capacitor and an electrons per millivolt conversion factor.
Procedure:
- for fixed test-pulse height, perform a threshold scan
- plot fraction of hit pixels vs. threshold
- fit Gauss error function to this s-curve
- mean of the fit provides the threshold 8corresponding to the test-pulse height used

## Slide 9: 2) TOT calibration: e/vm and capacitance
Update: use separate curves for odd and even columns and appropriate e/THL DAC factor in calculations

## Slide 10: 2) TOT calibration: e/vm and capacitance
 | Even | Odd
S-curve mean | 1518 | 1525
 | ± 0.078 THL DAC | ± 0.086 THL DAC
e/mV factor | 64.17 | 68.24
 | ± 2.31 e/mV | ± 2.87 e/mV
Capacitance | 10.28 | 10.93
 | ± 0.37 fF | ± 0.46 fF

## Slide 11: 3) TOT calibration: calibration curves
For the calibration curves of ToT vs. electrons, use the same procedure as before:
- for fixed threshold, perform a scan of test-pulse height
- convert test-pulse height in mV to electrons using conversion factor
- plot ToT vs. electrons
- fit surrogate function to parameterise curve
Update: now use appropriate e/mV conversion factor for odd and even columns.

## Slide 12
Results: ToT calibration curves still fit the data set well in the linear and low-ToT regions

## Slide 13
In surrogate function, A parameter describes gradient of the linear region
See significant difference in this parameter between odd and even columns
To be expected, since use a different e/mV factor for odd and even columns, increasing the difference in the gradient

## Slide 14
In surrogate function, B parameter describes y intercept of the linear region
Very similar distributions for odd and even columns

## Slide 15
In surrogate function, C parameter describes curvature of the function at low ToT. Hard parameter to constrain, due to low statistics in this range.
In 2D, see difference in parameter between top and bottom rows of the matrix.
Still see large spread of values in 1D, with a tail to higher C values.

## Slide 16
In surrogate function, T parameter describes x intercept. This is an interesting parameter to constrain as this is the threshold level in electrons.
See double peaked structure for both odd an even columns

## Slide 17
Using ‘old’ calibration method
Using separated even-odd column calibration

## Slide 18
In surrogate function, B parameter describes y intercept of the linear region
In 2D, see more apparent row difference, comparing top and bottom halves of the matrix.
For both odd and even columns, obtain a χ2/ndof close to ~1

## Slide 19: 4) Test-beam results
Re-ran test-beam analysis of DESY data using the new ToT calibration.
Data from July 2019 test-beam at DESY using assembly 20 (130um sensor) at a bias voltage of -25V and a threshold of ~885 electrons.

## Slide 20: associated cluster and pixel charge
electrons
Associated cluster charge has landau-gauss shape that is expected, peaking around ~10,200 electrons.
Is this what we expect?
for MIPs, expect ~80 electrons per um of silicon: 80 x 130um = 10,400 electrons

## Slide 21: associated cluster and pixel charge
electrons
Associated pixel charge has peaked structure, occurs due to the integer nature of ToT measurements. (see backup slides for more detail).

## Slide 22: in-pixel charge plots
In-pixel plot showing charge of the seed pixel of each cluster, where seed pixel has the highest charge in the cluster, follows the same pattern as the in-pixel cluster size, as expected.
In-pixel plot of cluster charge is ~uniform, this is expected as the charge of a particle should not depend on the position it hits within a pixel.

## Slide 23: conclusion
Updated threshold and ToT calibration for CLICpix2 assemblies to be calculated separately for odd and even columns separately.
Observe a ~0.65fF difference in capacitance between odd and even columns using this method.
Showed surrogate function parameters of ToT calibration curves for odd and even columns, where the T parameter has an interesting distribution.
This calibration has been applied to test-beam data, the results of which agree very well with theoretical expectations.
Next steps: apply an eta-correction to test-beam data, to correct for non-linear charge sharing between pixels.

## Slide 24: backup

## Slide 25: associated Pixel charge distribution

## Slide 26: associated Pixel charge distribution (rough calculation)
ToT of 1 = ~1.0 ke
ToT of 2 = ~1.0 ke
ToT of 3 = ~1.2 ke
ToT of 4 = ~1.6 ke
ToT of 5 = ~2.2 ke
ToT of 6 = ~3.0 ke
ToT of 7 = ~3.9 ke
ToT of 8 = ~4.8 ke
Steps of:
0
0.2
0.4
0.6
0.8
0.9
0.9

## Slide 27: associated Pixel charge distribution (rough calculation)
Peaks observed at:
1.0-1.4 ke
1.8 ke
2.5 ke
3.5 ke
4.5 ke
5.5-12 ke
0.7
1.0
1.0

## Slide 28: associated Pixel charge distribution (rough calculation)
Peaks observed at:
1.0-1.4 ke
1.8 ke
2.5 ke
3.5 ke
4.5 ke
5.5-12 ke
Steps of:
0
0.2
0.4
0.6
0.8
0.9
0.9
0.7
1.0
1.0