## Slide 2: Cryolab measurement capabilities and R&D – related to MSC
Torsten Koettig on behalf of the whole Cryolab team

## Slide 3: Outline
Introduction
Standard experimental infrastructure in the Cryolab
R&D projects performed together with MSC
Ongoing developments and new experimental systems
Future topics
Summary

## Slide 4: The Cryolab team TE/CRG-CI
2.5 scientists
1 technical engineer
1 mechanical engineer
2 Fellows
2 PhD students
2-3 technical students
2 FSU

## Slide 5: Cryolab projects 2019/20 35-40 short and long term projects
Technology development
RF cavity diagnostics BCCCA upgrade Cryo - Beam Loss Monitor RRR thin films, strips and bulk samples Thermal conductivity test stand Thermal dilatation test stand HFM λ-plate seal heat inleak Helios refill LKr purity monitor NA62 CAST Compass cooldown Rades 2 K - cavity validation Gas permeability LHC beam screen heaters RF surface impedance
Project driven
HL-LHC beam screen: non-standard conditions HL-LHC inner triplet current feeders Thermal cycling of 11 T coil He II heat transport in Nb3Sn cable geometries He II cooling simulations of Nb3Sn coils LAr calorimeters high-density feedthroughs BPM thermal link Near Tc wire characterization Nb3Sn FCC MLI test stand EASITrain - SC thin film on substrate DarkSide FCC transparent cryostat thermal performance LAr robot, CryFish Shape Memory Alloys ALPHA - Current leads test
Generic R&D
Thermal mapping of SRF cavities Cryocooler remote cooling study Cryogenic force calibrator Sub-Kelvin cryogenic cooling study
Training and Outreach
Cryogenic Safety courses: Fundamentals, LHe transfer, TSO Safety course: Self-Rescue mask Visits/Demonstration LN2 Open Days 2019 HSSIP
projects directly linked with MSC

## Slide 6: Standardized test stations and R&D projects
Material properties
Heat transfer
Simulations
Materials e.g. composites
Interfaces at low temperature
Transient conditions
& application
Remote cooling options
Standardized test stands
He II in confined geometries
Metal/epoxy to He
Novel concepts and R&D
HEX num. model to adapt to applic.
Custom design
& installation
New setups
Thermodynamics of Quench
MSC related topics

## Slide 7: Standardized test stations

## Slide 8: Cryolab test capabilities
Cryolab infrastructure for tests of:
  RRR for films, foils and bulk samples
10 x 35 mm2
2 x 100 mm2
6 bulk samples
2 x 2 x 120 mm3
He phase separator
Films on substrate
Foils e.g. Quench heaters
Bulk e.g. metals, SC like Nb

## Slide 9: Cryolab test capabilities
Cryolab infrastructure for tests of:
  RRR for films, foils and bulk samples
  SC splice resistance
Schematic of the setup Induced current in a SC loop with a resistive connection => R ~ 1-2 nΩ
Example of a SC cable with joint

## Slide 10: Cryolab test capabilities
Cryolab infrastructure for tests of:
  RRR for films, foils and bulk samples
  SC splice resistance
  Thermal conductivity test stand
  Thermal diffusivity test option
Courtesy: Patricia Borges de Sousa

## Slide 11: Cryolab test capabilities
Cryolab infrastructure for tests of:
  RRR for films, foils and bulk samples
  SC splice resistance
  Thermal conductivity test stand
  Thermal diffusivity test option
Courtesy: Patricia Borges de Sousa
Thermal conductivity of different 11 T dipole impregnation samples

## Slide 12: Cryolab test capabilities
Cryolab infrastructure for tests of:
  RRR for films, foils and bulk samples
  SC splice resistance
  Thermal conductivity test stand
  Thermal diffusivity option
  Thermal cycling (small and large scale)
next 2 slides

## Slide 13: Nb3Sn small scale samples thermal cycling to 2 K
Felix Wolf
4 samples to be cooled down
Immersion in He II
Warm up to 4.2 K => 293 K
80 K/h for T> 50 K
faster to reach 2 K
Sample box in G10
TT in sample space
Cover to protect from
He gas stream

## Slide 14: Thermal cycling 11 T coil thermal cycling 10 x to 80 K in GN2
Evaluation of the effect of thermal cycling on impregnated Quench heater insulation to the coil and the respective resistance (SP07)
  10 thermal cycles to 80 K in GN2 atmosphere
  Controlled ΔT during magnet cooldown
  Insulation resistance measurement after each cycle 660 V, 2 min
Setup of a gas distribution system, guaranteeing the max. thermal gradient across the magnet of 50 K (100 K) and cooldown in 24 h.
10 thermal cycles have been performed while staying within the maximum allowed temperature gradients.
No thermal degradation has been identified by MSC.
Humidity and temperature influence during the insulation tests => MSC
EDMS: 2150812 /2332230

## Slide 15: Cryolab test capabilities
Cryolab infrastructure for tests of:
  RRR for films, foils and bulk samples
  SC splice resistance
  Thermal conductivity test stand
  Thermal diffusivity option
  Thermal cycling (small and large scale)
R&D:
  A) Thermal dilatation, classic and laser based system
  B) Near Tc test stand for SC wires
  C) Tc of SC films on Cu substrate (cavities)

## Slide 16
Dilatation test stands in the Cryolab
The Laser dilatometer #3 (based on a cryocooler) is a completely new developed dry experimental station, and is under commissioning in the Cryolab right now.
The goal is to have a versatile measurement setup esp. in terms of sample dimension and composition.
Number | Setup | Sample dimensions | Temperature in K | Remarks
1 | Pyrex cryostat insert | 50 mm x 8 mm x 8 mm | 293 K reference, 78 K and 4.2 K | Requires LN2 or LHe
2 | Laser based cryostat insert | 80 mm x 15 mm x 15 mm | 5 K to 293 K continuous | Requires LHe 5 samples measured for MSC
3 | Laser based interferometer R&D | All dimensions up to: 200 mm x 70 mm x 50 mm Max. 18 mm dx from edge | 8 K to 293 K continuous | Sample in vacuum Cryocooler based Laser beam distance 21 mm

## Slide 17: Dilatation test stand – Pyrex insert
Reference material used for calibration of the glass insert (NIST)
Liquid level
0.8 m
Old system => cryostat insert
Sensor at RT
Inermet
50 mm sample length with special interfaces
Measurement points: 293 K reference, 78 K LN2 and 4.2 K LHe
Measurements: Inermet, Ultem 1000, G10 II and
Ͱ

## Slide 18: Dilatation test stand laser interferometer cryostat
Measurement with LN2 or LHe
Sample in vacuum
Cu shield + MLI around sample platform ensures
Reduction of the viewing factor at platform screen
Temperature uniformity by thermal screen
Experimental platform enables samples: 80 mm x 15 mm x 15 mm
Instrumentation feedthrough
Laser head
Thermal
screen

## Slide 19: Dilatation test stand – results Nb3Sn samples
Courtesy: Remy Kriboo, Flory ten Broeke
Matthias Michels
Axial direction

## Slide 20: Dilatation test stand – results Nb3Sn samples
Courtesy: Remy Kriboo, Flory ten Broeke
Matthias Michels
10 stack Radial
Axial direction
10 stack Azimuthal

## Slide 21: Dilatation test stand – results Nb3Sn samples
Courtesy: Remy Kriboo, Flory ten Broeke
Matthias Michels
10 stack Radial
Axial direction
10 stack Azimuthal
Additional test with azimuthal
straight 18 stack sample

## Slide 22: Upgrade dilatation test stand to cryocooler and laser
Vacuum enclosure
Thermal Shield
Cryocooler First Stage
Thermal shield of
sample environment
Laser interferometer
Courtesy: Remy Kriboo, Flory ten Broeke
Features:
LHe free system
Vacuum environment
Variable temperature 8 K to 293 K
Long term stable temperature
Controlled ramp rates
Easy to mount sample
Sample: 200 x 70 x 50 mm3
Open points:
Decoupling and alignment
Parasitic heat leaks

## Slide 23: Finished R&D topics: - Near Tc characterization of Nb3Sn wire (mechanical stress applied at room temperature)
Patrick Ebermann

## Slide 24: A) Near Tc test stand for Ic measurement of reacted Nb3Sn wire
PhD topic of Patrick Ebermann: “Analysis of electrical degradation by measuring Ic as a function of transversal stress applied at room temperature:
Cryogenic test 1 =>
=> Cryogenic test 2 =>
1st compression force applied at room temperature
2nd compression force applied at room temperature

## Slide 25: A) Near Tc test stand for Ic measurement of reacted Nb3Sn wire
Temp.
sensors
Nb3Sn wire
Transportable platform
Ttest=16-18 K
150 A current leads
PID of the Near Tc test stand, P. Ebermann
GHe LHe
Dewar
500 ltr
Phase Separator
Purge
45/2
45/3
180/2 Safety
outlet
4.3 K LHe
Bypass
Vacuum
pump
FT
001
FT
001
TT
060
TT
060
050
TT
050
TT
ET
022
ET
022
ET
021
ET
021
040
TT
040
TT
091
TT
091
TT
LT
001
LT
001
PI
001
PI
001
101
TT
101
TT
081
TT
081
TT
092
TT
092
TT
102
TT
102
TT
082
TT
082
TT
ET
031
ET
031
ET
032
ET
032
TT
007
TT
007
020
TT
020
ET TT
012
ET
012
FC
001
FC
001
Bypass
FT
002
FT
002
FC
002
FC
002
Cooling plate
with heat exchanger
(experimental platform)
Cooling plate
flow controller
Shield
flow controller
Vacuum gauge
< 10-6 mbar (HV)
EH 011 EH 012
Thermalisation of
current leads
on shield
molecular
pump
Turbo
030
TT
030
TT ET
011
ET
011
Dewar
pressure
controller
Bypass
PT
003
PT
003
PC
003
PC
003
155 mbarg
200 bar
Outlet
vapor
Outlet
liquid
Portable
sample holder
with sample
Cooling support
plate
Thermal shield
with heat
exchanger
He recovery
system
(30 mbarg)
Demountable sample holder
thermalized at 16 -18 K with dT=0.1 K

## Slide 26: A) Near Tc test stand for Ic measurement of reacted Nb3Sn wire
Courtesy: Patrick Ebermann, Irreversible degradation of superconducting Nb3Sn wires and cables caused by transverse stress at room temperature.
Keep Tsample constant

## Slide 27: A) Near Tc test stand for Ic measurement of reacted Nb3Sn wire
Courtesy: Patrick Ebermann, Irreversible degradation of superconducting Nb3Sn wires and cables caused by transverse stress at room temperature.
Keep Tsample constant
Controlled ramping Tsample

## Slide 28: A) Near Tc test stand for Ic measurement of reacted Nb3Sn wire
Feasibility shown:
  Test near Tc (16 K-18 K) possible
  Dismountable support allowing for room temperature stress application
  Heat loads are reasonable (CL 150 A) at 50 K and 15 K
  Possible to test with a cryocooler system

## Slide 29: Finished R&D topics: - Near Tc characterization of Nb3Sn wire - HIE ISOLDE thermalisation of RF cable and coupler

## Slide 30: HIE-ISOLDE RF cable temperature distribution
Established thermal mockup of HIE-ISOLDE semi rigid RF cable thermalisation from 293 K to 4.2 K incl. RF coupler and its antenna
Actively stabilized thermal shield
4.2 K coupler interface
13 Pt100 sensors at screen and RF cable
4 TVO at coupler and LHe reservoir
Heat meter at the cavity interface to calibrate after RF power tests
Results:
No thermal run-away observed in the tested set-up with RF power up to 95 W for 24 h => heat load to cavity in that case 0.8 W
Heat intercept of RF coax cable at thermal shield temperature effective, see TT8 at 1.38 m

## Slide 31: HIE-ISOLDE RF cable temperature distribution

## Slide 32: HIE-ISOLDE RF cable transverse thermal conductivity
[1] http://cryogenics.nist.gov/MPropsMAY/materialproperties.htm
[2] Perepechko, Low temperature Properties of Polymers

## Slide 33: Finished R&D topics: - Near Tc characterization of Nb3Sn wire - HIE ISOLDE thermalisation of RF cable and coupler - MLI thermal performance as fct. (Twarm and pvac)

## Slide 34: MLI performance test stand => host and support
7x10-7 mbar
Courtesy: Valentina Venturi
Vacuum and He degraded vacuum conditions
Different MLI configurations

## Slide 35: MLI performance test stand => host and support
2x10-5 mbar
Courtesy: Valentina Venturi
Vacuum and He degraded vacuum conditions
Different MLI configurations

## Slide 36: Finished R&D topics: - Near Tc characterization of Nb3Sn wire - HIE ISOLDE thermalisation of RF cable and coupler - MLI thermal performance as fct. (Twarm and pvac) - HFM Lambda plate seal – thermal heat leak

## Slide 37: HFM lambda plate seal test - heat load He I to He II
D-shape profile silicone based
choice tested
EDMS: 1845733 v.1

## Slide 38: HFM lambda plate seal test - heat load He I to He II
Arnaud van de Craen
Heat meter calibration with Indium seal
Silicone seal showed similar results of 2 W like In
EDMS: 1845733 v.1

## Slide 39: R&D topics - Thermal performance of He II cooled magnet coil samples

## Slide 40
Thermal performance of He II cooled NbTi & Nb3Sn coil samples
SC coil samples from prototype magnets are tested in pressurized He II
Step function heating
time constant response
Steady state temperature difference
Evaluation of the step heating response:
Steady state => Rth
Transient => Rth∙C = τ
Num. simulations
Courtesy: Rob van Weelderen, Tiemo Winkler, Mario Grosso, Kirtana Puthran, Lise Eder Murberg
Sample: LARP 007, tested in MQXFS3 conductor type RRP108/127 => MQXF coil
NbTi cable

## Slide 41
Thermal performance of He II cooled NbTi coil-pack samples
Courtesy: Tiemo Winkler, Jens Schundelmeier Rob van Weelderen
Original machine wrapped cable
Instrumentation in removed strand
Mechanical compression of 50 and 75 MPa

## Slide 42
Thermal performance of He II cooled Nb3Sn coil samples
Predicted He porosity 0.35-0.65% in the sample
Sample preparation for new D11T in progress (latest design)
To approach predicted operational heat loads on the magnets.
  Exploring thermal behaviour at higher heat loads by using an external heater and/or
  Variable frequency current source ν< 50 Hz
Steady state temp. D11 T sample
Steady state temp. MQXF sample
Volumetric heat input in mW/cm3
Volumetric heat input in mW/cm3
Time constant D11 T sample
Time constant MQXF sample
0.35%
0.6%
0.4%
0.65%
Courtesy: Rob van Weelderen, Tiemo Winkler, Mario Grosso, Kirtana Puthran, Lise Eder Murberg
He II

## Slide 43
Thermal performance of He II cooled Nb3Sn coil samples
Effect of magnetic field penetration and frequency dependency of the AC magnetic field.
Courtesy: Rob van Weelderen, Kirtana Puthran, Lise Eder Murberg

## Slide 44
Thermal performance of He II cooled Nb3Sn coil samples
Effect of magnetic field penetration and frequency dependency of the AC magnetic field.
Courtesy: Rob van Weelderen, Kirtana Puthran, Lise Eder Murberg

## Slide 45
Thermal performance of He II cooled Nb3Sn coil samples
Effect of magnetic field penetration and frequency dependency of the AC magnetic field.
Courtesy: Rob van Weelderen, Kirtana Puthran, Lise Eder Murberg
20 VAC peak, bipolar PS

## Slide 46: R&D topics - Thermal performance of He II cooled magnet coil samples - Thermal performance of HL-LHC Inner Triplet Current Feeders

## Slide 47: Thermal validation: HL-LHC Inner Triplet Current Feeders
Mock-up measurement setup
K-mod configuration
2 x 10mm2 Cu cable (Kapton insulated in stainless steel tubes)
Equivalent thermal path like in HL-LHC
Thermal performance validation
  Temperature feedthroughs to outside
  Heat load to 1.9 K pressurized bath
  Cooldown behavior and thermal stability

## Slide 48
Thermal validation: HL-LHC Inner Triplet Current Feeders
Mockup of electric current feeders 35 AAC
Equivalent thermal path like in the application => K-mod
Feeder connection from room temperature flange directly to He II bath.
Goals:
  Feedthrough temperature
  Temperature profile
  Thermal stability
  Heat inleak into 1.9 K LHe
  Electrical tests with CERN colleagues
Courtesy: Joanna Liberadzka

## Slide 49
Thermal validation: HL-LHC Inner Triplet Current Feeders
Heat input to the He II system (2 leads):
Simulation expected Q=1.9 W
Results
Warm end temperatures with and w/o HX and electrical current 25 A DC:
dT is important for dew point in the LHC tunnel air
Courtesy: Joanna Liberadzka

## Slide 50
Thermal validation: HL-LHC Inner Triplet Current Feeders
TE/MPE-EE: Electrical tests with pulses up to 7 kA were successful in the Cryolab setup, see EDMS: 2332911 v.1
Courtesy: Joanna Liberadzka
Simulations cable
Simulations tube

## Slide 51: R&D topics - Thermal performance of He II cooled magnet coil samples - Thermal performance of HL-LHC Inner Triplet Current Feeders - HTS small coil test

## Slide 52: Variable temperature cryostat setup for HTS small test coils
Test stand to determine the performance of novel HTS solenoids
Non-metallic cryostat insert
Operating temperatures at 4.2 K, 20 K … 80 K => LN2 or LHe or He vapor
Long ramp duration (stable for up to 6 h)
2 kA current leads => in-house build to fit the gas cooling insulated loop
Vehicle for MSC to test new technologies (like partial or non-insulated coils)
Drawings courtesy: D. Schoerling

## Slide 53: Vapor cooled non-magnetic cryostat setup, 2 kA current leads
Test of small HTS coils
Controlled inter-turn resistance
Temperature steps 4.2 K < T< 80 K
Adapted current leads for max. 2 kA
Fully non magnetic insert (Quench)
Glass epoxy chamber with temperature stabilization for several hours
Cryo system is commissioned
Courtesy: Joanna Liberadzka

## Slide 54: R&D - possibly interesting for MSC
- Remote cooling with He fluid circuits

## Slide 55: LHe zero boil off cryostat for the Antimatter Factory (BCCCA)
Detector with SC shield and SQUID sensor measuring the antiproton beam current.
Inner vessel
Detector
LHe
Courtesy: Miguel Abreu Fernandez, Jocelyn Tan, Andrew Lees

## Slide 56: LHe zero boil off cryostat for the Antimatter Factory (BCCCA)
Warm bore cryostat for Antiproton beam pipe
Stand alone cryo system
Complex design with ceramic breaks
Q<1.2 W available cooling power
Low values of: vibration, dT and magnetic field variations are crucial
=> paved the way to remote cooling circuit R&D
Courtesy: Miguel Abreu Fernandez, Jocelyn Tan, Andrew Lees

## Slide 57: Helium remote cooling circuit R&D
Research goal: Protect the experiment from cryocooler influences (vibration/mag. disturbance) or
separate the cryocooler from harsh environment (radiation/mag. field).
Novel HEX design
Enhanced tube in tube geometry
Highly compact geometry
High effectiveness of the HEXs is key for system performance
Design code
Adaptable to various applications
High pressure
hot flow
Low press.
cold flow
Low press.
cold flow
2 stage cryocooler
Results:
Performance is tested in between 293 K to 50 K,
CFHEX 1 achieved 94.6 % overall effectiveness, which will enable an increased cooling performance of the circulation loop compared to the cryocooler alone.
Courtesy: Aleksandra Onufrena

## Slide 58: New / further topics
Thermal conductivity of coil impregnation and a double Rutherford cable with one impregnation layer axial direction:
Emissivity of materials at T ≤ 4 K
Vacuum break at non-standard geometries and higher than 4 K temperatures
MQXF impregnation and a double cable structure

## Slide 59: Summary
Capabilities of the Cryolab and its personnel for tests and projects,
R&D in the range between 25 mK to 1.5 K to room temperature,
Staff is guiding the young generation of scientist and engineers in their projects
Our work is free of charge
We are interested to participate in the R&D of MSC
Goal is to further improve our testing capabilities and perform R&D together with you
Cryolab capabilities document: https://edms.cern.ch/document/2384399/1