## Slide 1
Radiation Protection at CERN
Helmut Vincke
CERN, HSE-RP
ESIPAP
26/1/2021
Notes: 1

## Slide 2: Content
CERN – Accelerators, our main source of radiation
Most important quantities used in Radiation Protection
Radiation Fields occurring at High Energy Accelerators
A few words about cosmic radiation
Radiation Protection carried out at CERN
Radiation detectors used in mixed radiation fields

## Slide 3: CERN - CONSEIL EUROPÉEN POUR LA RECHERCHE NUCLÉAIRE
1954:
Created by 12 European States
First European Organisation
Focus on nuclear physics (“nucleaire”)
Today: Particle Physics

## Slide 4
Intergovernmental Organization
SPS
CNGS
Meyrin
Prévessin
PS
LHC
~ 2400 staff
~ 800 other paid personnel
> 11000 registered users from ~ 100 countries
~ 3500 registered contractors
23 Member States: Austria, Belgium, Bulgaria, the Czech Republic, Denmark, Finland, France, Germany, Greece, Hungary, Israel, Italy, Netherlands, Norway, Poland, Portugal, Rumania, Serbia Slovakia, Spain, Sweden, Switzerland and the United Kingdom.
7 Observers to Council: Japan, the Russian Federation, the United States of America, the European Union, JINR and UNESCO
Notes: Exploration of a new energy frontier - Proton-proton collisions at ECM = 14 TeV
Start-up of the Large Hadron Collider (LHC), one of the largest and truly global scientific projects ever, is the most exciting turning point in particle physics.
Full membership to be confirmed very soon : Romania
Group of candidates for membership: Israel, Cyprus, Serbia, Slovenia, Turkey.
CERN was founded 1954: 12 European States
Geographical enlargement going on at present
Non-Member States with co-operation agreements
Algeria - Argentina - Armenia - Australia - Azerbaijan - Belarus - Bolivia - Brazil - Canada - Chile - China - Colombia - Croatia - Cyprus - Ecuador - Egypt - Estonia - Former Yugoslav Republic of Macedonia (FYROM) - Georgia - Iceland - Iran - Jordan - Korea - Lithuania - Malta - Mexico - Montenegro - Morocco - New Zealand - Pakistan - Peru - Saudi Arabia - Serbia - Slovenia - South Africa - Ukraine - United Arab Emirates - Vietnam
Other Non-Member States with Scientific Contacts at CERN
China (Taipei) - Cuba - Ghana - Ireland - Latvia - Lebanon - Madagascar - Malaysia - Mozambique - Palestinian Authority - Philippines - Qatar - Rwanda - Singapore - Sri Lanka - Thailand - Tunisia - Uzbekistan - Venezuela

## Slide 5: CERN’s main accelerator chain

## Slide 6
Main accelerator chain at CERN (low-energy part)
LINAC (160 MeV)
200 m
Booster (2.0 GeV)
Proton Synchrotron (26 GeV/c)

## Slide 7
1 km
Super Proton Synchrotron (450 GeV/c)
LHC (2 x 7 TeV)
Accelerators at CERN
(high-energy part)
SPS Circumference: 7 km
Intensity per filling of the ring: 6E13 protons
SPS beam
Fixed target experiments
AWAKE
LHC
FT
CNGS

## Slide 8
Particle type: Protons
Beam energy: 7 TeV
Number of stored particles: 2×4·1014
Stored energy: ~ 2×450 MJ
Mass at rest: ~ 1 ng
Mass in laboratory system: ~10 mg
Large Hadron Collider (LHC)

## Slide 9
The same amount of energy as in one LHC beam is stored in:
An F16 airplane at supersonic speed (mach 1)
In an aircraft carrier at a speed of 12 km/h
Notes: 11 GJ stored in magnets

## Slide 10
CERN operates powerfull (and dangerous) beams: full impact of a typical high-energy beam on tungsten (~ 0.16 % of the LHC beam energy)
Beam energy 440 GeV/c (SPS beam momentum)
1.08×1013 protons on tungsten alloy (Inermet 180)
Beam impact within ~ 8 us
Question: How will the three 3 cm long tungsten alloy blocks digest the beam impact?
By courtesy of A. Bertarelli

## Slide 11: High speed camera catching the beam impact
Beam
~10.5 mm
Beam
Details: Bertarelli et al., An experiment to test advanced materials impacted by intense proton pulses at CERN HiRadMat facility, Nucl. Instr. Meth. B (2013) http://dx.doi.org/10.1016/j.nimb.2013.05.007

## Slide 12: Most important quantities used in Radiation Protection

## Slide 13: Absorbed dose
Unit: J/kg = Gray (Gy)
refers to the energy deposited (not released) in matter.
It reflects the sum of the energies dEdep deposited by incident particles in a sample of matter, divided by the mass dm of the sample.
dm

## Slide 14: Equivalent dose in an organ or tissue, HT
ICRP publication 103
The radiation weighting factor (especially for neutrons) has been revised over time and remains controversial
Unit: Sievert (Sv)

## Slide 15: Effective dose, E
Organ | Tissue weighting factor |  | Organ | Tissue weighting factor
Gonads | 0.08 |  | Esophagus | 0.04
Red Bone Marrow | 0.12 |  | Thyroid | 0.04
Colon | 0.12 |  | Skin | 0.01
Lung | 0.12 |  | Bone surface | 0.01
Stomach | 0.12 |  | Salivary glands | 0.01
Breasts | 0.12 |  | Brain | 0.01
Bladder | 0.04 |  | Remainder of body | 0.12
Liver | 0.04 |  |  | 
Different organs show different sensitivity to equivalent dose deposited
Unit: Sievert (Sv)

## Slide 16: Ambient-dose-equivalent, H*(10)
denotes the operational dose quantity used for area monitoring of penetrating radiation. Such a quantity is required since the effective dose is not directly measurable (different weighting factors for organs and particles). The H*(10) quantity is measured via the ICRU sphere:
ICRU sphere: A sphere of 30 cm diameter made of tissue equivalent material with a density of 1 g/cm3 and a mass composition of 76.2% oxygen, 11.1% carbon, 10.1% hydrogen and 2.6% nitrogen
Dose equivalent in this point (10 mm inside the sphere)
30 cm
Unit: Sievert (Sv)
Ambient dose equivalent is a conservative measure for effective dose.

## Slide 17: Radiation Fields around High Energy Accelerators
Introduction
Impact of ionizing radiation in accelerators
Dose to people, shielding
  Radiation fields lateral to beam impact points
  Radiation fields downstream of beam impact points
Contents

## Slide 18: Radiation Fields around High Energy Accelerators
Prompt ionizing radiation – beam on
REM counter
High pressure ionization chamber
Ionizing Radiation due to induced radioactivity – beam off
Air filled ionizat-ion chamber
a, b-, g-radiation,
Main g energies: < 2.76 MeV
Whole particle zoo with E up to initial beam energy
Handheld devices

## Slide 19: Prompt Ionising Radiation
Cosmos
Hadron accelerator
Particle impact creates high-energy mixed radiation fields
Highest particle energy measured: 3.2×1020 eV
up to 7 TeV

## Slide 20
Prompt Ionizing Radiation in Accelerators
Next slides show the production of ionizing radiation by ONE hadron (120 GeV/c ) on copper Target
Ionizing radiation in accelerators is produced by any beam impacts of high energy particles  secondary radiation
Impact of very energetic particles produce particle showers

## Slide 21: What happens when a 120 GeV proton impacts? Hadronic cascade
Proton
Neutron
Pion+
Pion-
Electron
Positron
Photon

## Slide 22: What happens when a 120 GeV proton impacts? Full cascade
Proton
Neutron
Pion+
Pion-
Electron
Positron
Photon

## Slide 23: Ionizing particles on matter: Impact and consequences
  Activation of material
  Radiation triggered failure of electronics
Focused energy deposition in material  heat development, shockwaves destruction of materials
  Aging of organic materials like insulations
  Dose to people

## Slide 24
Relation between absorbed dose and damage caused by radiation
Gy
LD 50/30
Annual dose limit (Sv) for radiation workers in Switzerland
Radiation impact
Problems with electronics equipment occur
Problems with cable insulation occur
Problems with magnet coil insulation occur

## Slide 25: Dose to people

## Slide 26
Example of full beam loss (7 TeV) in the LHC
Dose caused by a full beam loss (4E14 protons@7TeV) reaches levels of tens of thousands of Sv close to the loss point.
 strong shielding required to significantly reduce dose
mSv/full beam loss

## Slide 27
Main aspects of radiation field attenuation in accelerator environments
Lateral to beam impact point
Downstream to beam impact point

## Slide 28
Radiation lateral to the beam impact point

## Slide 29
Lateral Shielding Configuration: Simulation to calculate radiation propagation through a lateral shielding wall
A 450 GeV proton beam is sent onto a 5 m long target with a diameter of 5 cm.
Target is surrounded by particle detectors to evaluate the particles’ spectra
One quarter clipped for viewing the inner structure of geometry

## Slide 30
Buildup
Attenuation
~e-d/l
Dose analysis
Shielding
Point source/line of sight model
Hp/2 … source term (at 1 m) for quantity of interest (e.g. effective dose)
r= r0+d+r1 in m
l… hadronic interaction length
Source
r0
d
r1
Point of interest

## Slide 31
Typical particle fluence spectra at areas located behind lateral thick concrete shielding
Dose contributors
86% n 7 % p 4% p 3% g

## Slide 32
Radiation downstream to the beam impact point

## Slide 33
Radiaton fields occurring downstream of an impact point of an high-energy proton beam
450 GeV proton beam on 50 cm long copper target
2m

## Slide 34
Dose of all particles
Dose by muons only
Dose simulation results
Muons strongly dominate the dose seen downstream the heavy shielding

## Slide 35
Spectral analysis of the fluence seen in detector 1 and detector 3.
Strong domination of muons downstream the shielding
Particles at this locations can be traced back to muon interactions
High-energy muons lose ~1 GeV when traveling through 1.8 m of concrete or 70 cm of iron.
 To shield all muons a shielding of hundreds of m of iron is required.
450 GeV
Detector 1
Detector 3

## Slide 36
Iron shielding and dump
4 m
2.8 m
2 m
Concrete shielding
Demonstrating the muon problem by Conceptual design of an irradiation facility
Courtesy of E. Feldbaumer

## Slide 37
15 µSv/h
Contours:
15 µSv/h
450 GeV/c
24 GeV/c
Dose rate mapping at beam height
Comparison: 450 GeV/c vs 24 GeV/c
µSv/h
1010 particles/spill
1011 particles/spill
Example to show the muon dominance downstream of thick shielding construction in high-energy facilities
Courtesy of E. Feldbaumer

## Slide 38
Chart: Energy [GeV]
| 2.2250312357358189E-2 | 2.8674129611899357E-2 | 3.695198706429749E-2 | 4.7619586138478824E-2 | 6.1367341232939486E-2 | 7.9084225386609378E-2 | 0.10191760839030641 | 0.13134256735727379 | 0.16925741933516619 | 0.21811874151480068 | 0.28109135703539473 | 0.36224342920196634 | 0.46682108649888582 | 0.6015955648107788 | 0.7752776883156125
Neutrons | 7.2316e-09 | 4.2767e-09 | 1.4405e-08 | 1.1105e-08 | 1.3413e-08 | 1.3572e-08 | 1.24e-08 | 1.1577e-08 | 4.597e-09 | 5.2429e-09 | 1.1648e-09 | 2.4219e-09 | 4.5893e-10 | 5.3441e-11 | 9.6375e-12
Charged Hadrons | 2.5143e-10 | 1.2955e-11 | 6.5219e-11 | 2.2451e-11 | 3.4928e-11 | 6.0308e-10 | 1.9063e-10 | 5.7232e-11 | 2.3732e-09 | 3.7599e-10 | 1.005e-10 | 5.1933e-11
Muons | 1.1504e-08 | 1.118e-08 | 1.7444e-08 | 4.3252e-08 | 7.3337e-08 | 5.6665e-08 | 1.0815e-07 | 1.4694e-07 | 1.6816e-07 | 2.2152e-07 | 2.8235e-07 | 3.7844e-07 | 3.6518e-07 | 4.349e-07 | 4.3976e-07 | 3.8736e-07 | 2.8053e-07 | 2.2079e-07 | 1.3704e-07 | 1.1133e-07 | 9.9137e-08 | 4.0963e-08 | 3.032e-08 | 4.1247e-08
Neutrons | 9.2276e-10 | 7.9645e-10 | 1.7438e-09 | 1.5425e-09 | 1.0796e-09 | 1.2069e-09 | 1.7085e-09 | 9.0437e-10 | 9.9713e-10 | 5.5428e-10 | 2.2357e-10 | 5.6485e-11 | 4.4135e-11 | 1.1021e-11 | 1.4891e-11
Charged Hadrons | 2.6526e-11 | 2.1282e-12 | 3.9051e-11 | 8.177e-12 | 4.5268e-12 | 1.6635e-11 | 6.8622e-11 | 9.752e-12 | 2.2616e-11 | 2.647e-11 | 2.572e-11 | 2.1117e-13
Muons | 1.2577e-10 | 3.6022e-10 | 3.3307e-10 | 5.0678e-10 | 7.9553e-10 | 1.1014e-09 | 1.2229e-09 | 1.6925e-09 | 9.8701e-10 | 1.4929e-09 | 1.8362e-09 | 2.5402e-09 | 3.6125e-09 | 1.4601e-09 | 2.2468e-10
Spectra behind experimental setup consisting of a 50 cm copper target + 2 m of air + 4 m of iron dump
Courtesy of E. Feldbaumer

## Slide 39
Rough estimates for the muon dose as a function of shielding strength (collimated beam assumed)
Concrete shielding strength = Iron strength*2.6
The given shielding strength does not include the 4 m of iron between the target and the reference point.
Assuming a beam intensity of 1E11 p/s we would need the following amount of shielding to remain below 2.5 uSv/h (non designated area) a locations downstream the dump
~ 100 m of additional iron for 450 GeV/c
~ 5 m of additional iron for 24 GeV/c

## Slide 40
Cosmic radiation environment

## Slide 41
Contents:
Basic introduction to cosmic radiation fields
Interesting phenomena at high energies
Radiation Protection aspects in space

## Slide 42
Classification of cosmic (ionizing) radiation
Galactic Cosmic Radiation (GCR)
Solar Cosmic Radiation (SCR)
high-energy particles coming from the sun
consist of protons, electrons and ions with energies ranging from a few tens of keV to GeV
Two main processes of their production:
  solar-flares
  shock waves caused by coronal mass ejections.
Particles which originate from sources outside of the solar system, distributed throughout our Milky Way galaxy and beyond.

## Slide 43
Properties of Galactic Cosmic Radiation
Spectrum:
Main part of the GCR particles have an energy below 10 GeV
Interaction with solar magnetic field modulate the particle’s energy
 GCR fluence up to 10 GeV shows a dependence on the solar activity
There is a high-energy component of the GCR spectrum, reaching energies higher than 1020 eV.
Composition: 2% electrons and 98% nuclei
Fisk (1979): Mechanisms for energetic particle acceleration in the solar wind
Proton spectra observed during various years (1965: solar minimum, 1969: solar maximum)
Composition of nuclei:
87% protons 12 % a-particles 1% heavy nuclei
Almost no anti-matter detected

## Slide 44: A BIRD'S EYE VIEW OF THE ALL-PARTICLE CR SPECTRUM
In the high energy region the particle “flux” can be described by:
Knee
Foot (?)
Modulated by solar activity
1 particle per m2×second
1 particle per m2×year
2nd knee
Ankle
Fingers (?)
1 particle per km2×year
All nuclei
Expected GZK cutoff
| EAS experiments
ballons & satellites |
Picture taken from http://hep.fi.infn.it/PAMELA/naumov/Eng/UHECR/UHECR.html
F ~ E-a
with
a= 2.7 for E < 3×1015 eV
and
a= 3 for E > 3×1015 eV
Highest energy measured: 3.2×1020 eV (Fly’s Eye detector)

## Slide 45: Interesting effects at such high energies (assuming particle was a proton)
Energy of particle : 3.2×1020 eV = 51 J
Kinetic energy of a golf ball (170 km/h)
From
we calculate Lorentz factor g:
= 3.4×1011
1 second for the proton are 10773 years for us!!!
Length contraction:
Time dilatation: t= t0×g
For the proton the earth appears as a disk with a
maximum thickness of 39 micro m.

## Slide 46
Relativistic Doppler Effect
High energy particles are subject to Doppler Effect, changing the frequency of photons.
Considering a head-on collision between a proton and a photon, the frequency of the photon is seen as:
With: b=v/c … b (3.2×1020 eV)= (1 - 4.8×10-24)
Typical energy of microwave photon: 1.1×10-3 eV.
Proton should interact with microwave photons producing pions.
This calculation coincides with the Greisen–Zatsepin–Kuzmin limit (GZK limit), defining a theoretical upper limit (6E19 eV) for the energy of cosmic rays coming from sources beyond ~ 50 MParsec.
However, we observe such particles, although no sources for such high-energy particles can be pinpointed within this radius. These observations remain an unsolved riddle.

## Slide 47
Dose to space- and aircraft crews
Ambient-dose-equivalent rates as a function of standard barometric altitude
(at 2 GV vertical geomagnetic cut-off rigidity and mid solar cycle, calculated by S. Rollet, taken from Oxford University Press et al. Journal of the ICRU 2010;10:17-21)

## Slide 48
Dose exposure during a space trip to Mars
In space crafts only limited shielding power for the protection of the crew can be provided.
During a one way trip to Mars, the dose received by the crew is estimated to (330 ± 6) mSv, resulting in a dose of 660 mSv for a round trip. Exposure during Mars (no magnetic field protection) is not included in this calculation.
These results are based on measurements carried out in the Mars Science Laboratory spacecraft*.
Real dose during flight depends also strongly on the sun activity
*Science 31 May 2013: Vol. 340 no. 6136 pp. 1080-1084, DOI: 10.1126/science.1235989

## Slide 49: Aging of organic materials like insulations

## Slide 50: Examples for radiation damage Resin used for magnet coil insulation
0 5x106 107 2.5x107 5x107
D (Gy)
Notes: La dosimétrie à haute dose
Les points de mesures
Le long de accélérateur
La fréquence des mesures
Les ancienne répartition de dose des années passer

## Slide 51
Radiation damage on cable insulations

## Slide 52
Radiation damage is mainly caused by braking hydrogen bridge bounds in molecules
In radiation facilities insulation material shall be chosen according to the radiation level in the area.
In the last millennium many radiation hardness tests were carried out and documented at CERN.
Material catalogues from the past are available and should be used
Radiation resistance of cable materials

## Slide 53
Induced radioactivity in and around (high-energy) particle accelerators

## Slide 54
Introduction to radioactivity and isotope decay
Activation in accelerators
Calculation procedures to forecast activation in accelerators
Examples of activation at high-energy accelerators
Contents:

## Slide 55: Radioactivity

## Slide 56: What is radioactivity?
Spontaneous emission of radiation from unstable nuclei. The consequence of most of the radioactivity reactions are combined with the change of the emitting nucleus into another kind of nucleus.
Modes of radioactive decay
a-decay: Emission of an alpha particle, a part of the nucleus consiting of 2 protons and 2 neutrons.  Anew= Aold-4 and Znew=Zold-2
b--decay: A neutron in the nucleus is transformed into a proton via the emission of an electron and an anti-electron neutrino.  Anew= Aold and Znew=Zold+1
Example:
Example:
U-235
Th-231
a

## Slide 57
Modes of radioactive decay
Electron capture: An electron from the atomic orbit is captured by a proton resulting to the transformation into a neutron.  Anew= Aold and Znew=Zold-1
b+-decay: A proton in the nucleus is transformed into a neutron via the emission of a positron and an electron-neutrino  Anew= Aold and Znew=Zold-1
Example:
Alternative decay of C-11
Example:
g-decay: In a gamma decay a nucleus changes from a higher energetic state to a lower energetic state by emitting a high-energy photon (gamma particle). The composition of the nucleus remains unchanged.
All decay pictures taken from Jefferrson lab web page

## Slide 58: Radiation types emitted by radioactivity
How can we shield these particles?
How dangerours are such particles when being incorporated?
Radioactivity results in the emission of a, b+, b- and g radiation.
Sheet of paper
Several mm of aluminium
Heavy shielding
The risk caused by a-radiation is highest when being incorporated and decaying inside the body.

## Slide 59: Radioactive decay
Decay of a radioactive material per time unit (activity) as a function of time:
Unit: Bq: 1 Bq = 1 decay per second
N... Number of radioactive isotopes
A ... Activity, decay per time unit
l… Decay constant, defining the speed of decay
Proportionality factor is called decay constant (l) and it defines the probability of the deacy per time unit
l
With N(t=0) = N0

## Slide 60
General production-decay chains
A decay of an isotope can result in a chain (even several chains) of decays
Beside production via decay some or even all isotopes can be produced by external production processes. E.g.: activation of materials in accelerators.
Uranium-Radium decay chain
Each possible chain in this tree can be described by coupled differential equations
PTh-234
Pra-226
PRn-222
PPb-214
PPo-210
PPo-218
PU-238

## Slide 61
…
…
…
…
…
…
Nn … Number of isotope n
Pn … Production rate of isotope n
ln … Decay constant of isotope n
bn … Branching ratio from isotope n-1 into n
Mathematical expression of one production-decay chain (Bateman equation)
System of coupled differential equations
Solving by Laplace transformation of system of differential equations
)

## Slide 62
Laplace transformation of system of differential equations  system of linear algebraic equations
Laplace transformation to find solutions for complicated
radioactive decay problems
It transforms a function being dependent from t into a new function F being dependent from s

## Slide 63
Nn … number of isotope n
Pn … production rate of isotope n
ln … decay constant of isotope n
bn … branching ratio from isotope n-1 into n
Fn (s) … Laplace transformed of
Nn(t)
Mathematical expression of one production-decay chain (Bateman equation)
Laplace transformed equations = system of linear equations, to be solved in the Laplace domain as a function of s.
To obtain the final result for a given isotope the contributions of the various chains have to be summed up.
…
…
…
…
Inverse Laplace transformation of Fn (s) Nn(t)
)
Decay term
Production term

## Slide 64
Activation: Radioactivity production in accelerators

## Slide 65
First questions about activation
What is activation?
Activation can be described as the imposed change of nuclear composition of given isotopes resulting in the production of radioactivity.
Impact of activation:
Accelerators: due to beam operation’s losses, accelerator and environment will become radioactive  dose to personnel and environment.
What can be done to reduce activation?
Reduce beam losses
Reduce activation prone material

## Slide 66
Which production mechanisms of activation occur at high-energy accelerators?
At high-energy accelerators primary particles interact with matter. The primary particle itself or secondary particles interacting with nuclei can produce radioactive isotopes. Main production channels of activation at high-energy accelerators are:
  Spallation processes, (n,2n), (n,p), (n,alpha), ..
  Particle capture (mainly neutrons)
  (g,n)-reactions (important for electron accelerators)

## Slide 67
Why is activation important for high-energy accelerators?
Required information to classify activation:
  Specific activity: classification of material between radioactive and non-radioactive material
  Dose rate around the activated components
Air activation
Water activation
Soil activation
Environment
Material activation of beam line, tunnel and shielding components
Questions about activation III/III
Tunnel and accelerator

## Slide 68
Parameters responsible for production of radioactivity in a given material sample at high-energy accelerators
Composition of material sample
Beam type and energy
Irradiation history
Cooling time
Radionuclide production in sample per lost proton
Evolution of radioactivity as a function of time
Beam intensity
Location of sample in given radiation environment
Material around loss point

## Slide 69
Total production of radio-nuclide i per lost proton
Physics principles of radio nuclide production per lost proton
Beam on target situation:
Sample positions
Target
Shielding
Energy (E) and particle type dependent production cross section to produce nuclide i from isotope j
Atomic density of isotope j in the given volume
Total track length of particle type k through volume of interest as a function of energy (E)
Chemical composition | 
Material | Wt-%
Fe | 65.0
Cr | 15.0
Ni | 14.0
Mo | 5.0
Co | 1.0

## Slide 70
Activation as a function of beam operation time and cooling time
Ni… Number of isotopes i
l… decay constant of nuclide i
Pi… production rate of isotope i per proton
I … proton intensity
Ai… Activity of isotope i after tirr and tcool
tirr… irradiation time
tcool… cooling time
Ai,max ≤ Pi ∙ I
Build-up and decay
90% of Ai,max are obtained after ~ 3.32∙t1/2
99% of Ai,max are obtained after ~ 6.64∙t1/2

## Slide 71: Calculation procedures to forecast activation and dose rates in high-energy accelerators
Analytic method
Full Monte Carlo calculation

## Slide 72
Analytic methods
Required input parameters:
  Track length spectra for all relevant particle types, Lk(E)
  Cross sections for radio nuclide production si,j,k
  Irradiation and cooling history
Analytical estimate of activation
}
Pro:
Fast activation result if input parameters are available
Cons:
Track length of various particles fields are required (very often Monte Carlo results)
Only rough dose rate estimate without self shielding effects

## Slide 73
Full Monte Carlo calculation of activation
Procedure inside code:
1) Simulation of particle cascade and isotope production around beam impact point
2) Radiation emerging from radio-isotopes are further transported to calculate dose rate in the surroundings of activated material
Input required:
3D geometry description
Beam energy and intensity
Irradiation history and cooling time(s)

## Slide 74
Applications for FLUKA activation simulations

## Slide 75
Particle type: Protons
Beam energy: 7 TeV
Number of stored particles: 2×4·1014
Applications at the Large Hadron Collider (LHC)

## Slide 76
CMS detector
Parameters of CMS detector:
Length: 21 m
Diameter: 15 m,
Mass: 12500 t
Operational parameters:
Two counter rotating beams @ 7 TeV
Luminosity: 1034 cm-2 s-1
109 proton-proton collisions/s
Collision point
21 m
7.5 m

## Slide 77
Residual dose rates to be expected after beam operation
Residual dose rate expected after 1st year of operation:
180 days of operation +
6 cooling times: 1 h, 1d, 1w, 1m, 4m

## Slide 78
180 days of irradiation, 109 pp/s, 1h of cooling

## Slide 79
180 days of irradiation, 109 pp/s, 1d of cooling

## Slide 80
180 days of irradiation, 109 pp/s, 1w of cooling

## Slide 81
180 days of irradiation, 109 pp/s, 1m of cooling

## Slide 82
180 days of irradiation, 109 pp/s, 4m of cooling

## Slide 83
Radiation Protection at CERN

## Slide 84
Contents
Mandate
Radiation Protection Regulations
Dose limits and objectives
Dosimetry, Operational Radiation Protection and Radiation Monitoring
Radioactive waste: treatment and elimination
ALARA at CERN

## Slide 85: Radiation Protection at CERN: Mandate
Monitoring
Riskanalysis
Controls
Authorization
Safety Code F
Legislation
Radioactive Waste

## Slide 86: Some Key Figures…
Radiation Areas and Radioactive Laboratories:
~ 45 km accelerator tunnel
Class A,B and C laboratories
RIB facility ISOLDE
Spallation source n-TOF
~ 50 - 60 access points
~ 160 experiments
~ 11000 radiation workers
Many new projects

## Slide 87
General Principles of Radiation Protection
Justification
any exposure of persons to ionizing radiation has to be justified
2) Limitation
the personal doses have to be kept below the legal limits
3) Optimization
the personal doses and collective doses have to be kept as low as reasonable achievable (ALARA)
Radiation Protection Regulation

## Slide 88: CERN’s Radiation Protection Regulation
Guideline 96/29 Euratom laying down the basic standards for protecting public and workers against the risk of ionising radiation
IAEA Basic Safety Standards
CERN is an intergovernmental organization and not bound to any national law* - but
CERN Safety Code F (Radiation Protection Ordinance) and underlying safety instructions, guidelines, etc.
*) CERN’s relation with its two Host States is defined in conventions between the parties
Taken from B. Lorenz, WKK Symposium April 2008 and modified

## Slide 89
Tripartite Agreement
RATIFICATION OF THE TRIPARTITE AGREEMENT ON SAFETY AND RADIATION PROTECTION
(September 2011)

## Slide 90: History of Radiation Protection
Chart: Year
M. Curie
ICRP 103
Source: Los Alamos Science
Nr. 23, 1995, p. 116

## Slide 91: Dose Limits
 | Dose limits for 12 months consecutive (mSv) |  | 
 | Non-occupationally exposed persons | Occupationally exposed persons | 
 |  | B | A
EURATOM | < 1 | < 6 | < 20
Germany/France | < 1 | < 6 | < 20
CERN | < 1 | < 6 | < 20
Switzerland | < 1 | < 20 | 

## Slide 92: CERN’s Dose Objectives
Category | Dose/Year
Reference group of Public | < 10 uSv
Non-professionally exposed personnel | < 100 uSv
Professionally exposed personnel | < 6 mSv

## Slide 93: Dosimetry, Operational Radiation Protection and Radiation Monitoring

## Slide 94: Individual Dosimetry
Up to 11000 monitored persons per year
98 - 99% of individual doses < 1 mSv
Dose interval (mSv) | Persons concerned | Persons concerned | Persons concerned | Persons concerned | Persons concerned | Persons concerned | Persons concerned | Persons concerned | Persons concerned | Persons concerned | Persons concerned
years | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | Jan- Nov 2019]
0.0 | 5042 | 5418 | 5315 | 6002 | 6273 | 7616 | 8704 | 8788 | 9034 | 9824 | 7246
0.1-0.9 | 1219 | 1514 | 1984 | 2030 | 2188 | 1816 | 1108 | 1003 | 1110 | 1260 | 1400
1.0-1.9 | 39 | 31 | 31 | 29 | 82 | 133 | 2 | 11 | 3 | 10 | 10
2.0-2.9 | 13 | 6 | 7 | 0 | 3 | 14 | 0 | 0 | 0 | 0 | 0
3.0-3.9 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
4.0-4.9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5.0-5.9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
> 6.0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SUM PERS | 6315 | 6969 | 7337 | 8061 | 8546 | 9580 | 9814 | 9802 | 10147 | 11094 | 8656
0.0 – 0.1
0.1 – 1.0
1.0 – 2.0
2.0 – 3.0
3.0 – 4.0
4.0 – 5.0
5.0 – 6.0
> 6.0

## Slide 95: Collective Personal Dose Equivalent Summary over 25 years
From 2015 onwards: low dose could be measured more precisely

## Slide 96: Operational Radiation Protection
1 hour
1 week
1 month
From design
to reality
Notes: HiRADMat
Ramses materials
-yellow inside facility
- Red ventilation outlet

## Slide 97: Radiation Monitoring
with
location ID
Radiation
Display
Display
Control box
Radiation
Monitor
Monitor
Controller
Direct
hardware
connection
Tap box
Basic Area
Controller
with
ID + local
database

## Slide 98
Start up of LHC – Seen by neutron monitors in USA15
30 nSv/h
Hourly values
10 nSv/h
Instrumentation

## Slide 99: Area classification: LHC - 2010

## Slide 100: LHC - 2011

## Slide 101: LHC - 2012
Point 1
Point 8
Point 7
Point 6
Point 5
Point 4
Point 3
Point 2
Expectation the LS1
LHC tunnel mostly classified as Supervised Radiation Area
Few areas might be classified as Simple Controlled Radiation Area
Collimation and dump areas will be classified as Limited Stay Area

## Slide 102: LHC during LS1

## Slide 103: Radioactive waste: treatment and elimination
Analyzed
Eliminated to final repository
Conditioned
Waste received
Chart: Incoming waste in RWTC premises
| 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 44168
Masse (T) | 193 | 126 | 358 | 427 | 295 | 263 | 132 | 523.7 | 323.1 | 638.6 | 191
Vol (m3) | 445 | 399 | 521 | 1030 | 792 | 495 | 315 | 684 | 479.7 | 980 | 561

## Slide 104: ALARA at CERN
ALARA procedure – 3 Levels:
  I Optimisation
  II Optimisation, documentation
  III Optimisation, documentation, ALARA Committee
Includes risk analysis

## Slide 105: ALARA
Mock-up training
Procedures
Approval by “stakeholders” – radiation protection incl.
Lessons learned

## Slide 106: ALARA example 1: SPS-LSS1 cabling and dump shielding wall design (ALARA III)

## Slide 107
The LSS1 area is the most radioactive zone in the SPS
Any work has to be fully optimized allowing to reduce dose to personnel to a bare minimum
Removal of highly radioactive equipment prior the 22 weeks lasting cable exchange campaign
 average dose rate in the LSS1+ area was reduced by a factor of 3.2  dose reduction of several tens of mSv
Remote handling with robots
Special equipment allowing for dose optimization
Cable exchange campaign in SPS-LSS1
Optimized cable removal and cutting
Remote cable pulling machine
Cables combs replacing cable ties

## Slide 108
FLUKA
geometry
In the SPS-CSAP a working group (chaired by RP) was founded to replace this shielding wall by a new design.
Ten options were studied in order to optimize the new wall with respect to:
  Dose-to-cable reduction
  Residual dose rate during beam-off periods
  Reduction of radioactive waste
  Transport handling issues
  Compatibility with other equipment
  Allowing cable maintenance or exchange without removing the wall
  Compatibility with cable routing plans
  …
The former shielding located beside the TIDV dump caused significant problems due to its high activation and contamination levels.
SPS beam dump shielding
Final shielding design, being the optimized choice in material, dimensions and installation position
Dose rate after 1 week of cooling

## Slide 109
ALARA example 2: Dismantling of former SPS target area

## Slide 110
Dose rates based on measurements and simulations
Job and Dose Planning
For all working steps i, the timei required and the given dose ratei need to be assessed beforehand.
Example: Work in a former SPS target area

## Slide 111: Removal of highly radioactive blocks being located in a former SPS target area

## Slide 112
ALARA example 3: Repair work of CNGS horn and reflector

## Slide 113: CNGS Horn und Reflector Repair
level II
optimisation and documentation
1.6 mSv collective dose
Shielded cabin
Mobile lead shield
Notes:     Rupture of ceramic insulator brazed to Ti flanges
  (no brazing, only under compression)
  all water outlets and inlets for both horn and reflector (improved geometry)

## Slide 114: Radiation detector systems used to measure mixed radiation fields at workplaces

## Slide 115: Goal of this chapter
We will introduce the RAMSES system using three different radiation detector types which are mainly used as fixed detector systems at CERN
We will learn about their main fields of application in terms of radiation fields.

## Slide 116
RP Data Base
long term storage
Monitor
Controller
Termination
Termination
Network-based
Area Controller
RP User
Interface
Remote User
Interface
Internet
LHC point
(local Network)
CERN Networks
(Ethernet)
RP User
Interface
with
location ID
Radiation
Display
Display
Control box
Radiation
Monitor
Monitor
Controller
Direct
hardware
connection
Tap box
Basic Area
Controller
with
ID + local
database
Conceptual proposal
Configuration and
Supervision
Manager
(RAdiation Monitoring System for Environment and Safety)
Monitoring system RAMSES

## Slide 117
REM counter
High pressure ionization chamber
Air filled ionization chamber
Typical radiation detectors used at CERN for high-energy particle fields

## Slide 118
IG5 monitor (high pressure ionization chamber)
Used mainly in accessible areas during machine operation
Comes in two flavours
H-filled type: Calibrated to neutron sources
Ar-filled type: Calibrated to photon sources

## Slide 119: IG5 - Design
Properties
2 types (Ar or H filled)
5,2 l active volume
pressurized at 20 bar
1200 V high-voltage
Dimensions
diameter – 18.33 cm
height – 45.6 cm
electrodes
active medium
electronics
board
Notes: Electronics = 60% epoxy resin, 35% silicon, 2.5% copper, 2.5% carbon

## Slide 120
Characteristics of IG5 chamber responses (Ar versus H)
IG5 hydrogen type
Strong sensitivity to neutrons at energies relevant behind thick shielding.
Reduced sensitivity to charged particles compared to the Ar type total detector response in a neutron dominated field is dominated by neutron triggered counts
IG5 argon type
10 times stronger sensitivity to charged particles and photons when being compared to H-type
Detector response in mixed fields is strongly triggered by charged particles (e.g. muons) and photons

## Slide 121
Locations of IG5 chambers in our accelerator environment
To be used at locations where the dose is dominated by neutrons:
Behind lateral shielding (e.g.: LHC experiments)
upstream of beam loss points
at the end of mazes
To be used at locations where the dose is dominated by photons or charged particles:
Close to gamma sources (e.g.: ion exchanges of cooling water)
Behind massive beam dumps or shielding downstream high-energy beam loss points (muons dominating dose)
beam
Target or loss point
Shielding
IG5-H
beam
dump
IG5-Ar
Target or loss point
IG5 Argon type
IG5 hydrogen type

## Slide 122
REM counters adapted to high-energy neutron fields (WENDI2 detector)
To be used in neutron dominated areas (beam-on)

## Slide 123
Reminder to our typical radiation fields behind lateral shielding in accelerators
beam
Target or loss point
Shielding
Radiation Detector
Majority of dose originates from neutrons in the energy range of a few hundred of keV to a few hundred of MeV
Detector is sensitive to neutrons only!!!

## Slide 124
Neutron detector used at CERN: FHT 762 Wendi-2
Wide Energy Neutron Detector (Thermo)
PE cylinder
W-layer
He-3 counter
Borated rubber
Reference: Olsher et al, Health Physics, 79(2): 170ff, 2000
Equivalent instruments exist also from other companies
This instrument is sensitive up to high neutron-energies but blind to all other particles.

## Slide 125
PMI detector (ionization chamber)
Anode: PE / graphite coated
Active volume
Filling gas: air atmospheric pressure
Active volume: 3l
PE wall (4mm) / inside graphite coated
Wall composition: C-H2
Working voltage: ~460 V
Various connectors
28.5 cm
21.5 cm
15.8 cm
Main installation place: radioactive equipment inside accelerator tunnel.

## Slide 126
Use of PMI monitors in RAMSES
Main use: remote readout of monitors prior machine intervention when beam is off.  Monitor measures residual dose rate (ambient dose equivalent) coming from activated material in Sv/h.
Secondary use: remote readout of monitors during machine operation.  Monitor measures prompt dose rate (in Gy/h!!! not Sv/h) of the cascade originating from lost particles. The measured value can be interpreted as dose rate in air within an accuracy of 20%.

## Slide 127
The IG5 chambers are used mainly to monitor accessible areas during beam-on There are two types available:
  The hydrogen type is used mainly behind lateral shielding during beam-on
  The argon type is used mainly to monitor dose behind dumps or close to gamma sources (e.g.: ion exchangers)
REM counters provide a clear discrimination to other particles than neutrons  measuring only the neutron component of a radiation field
The PMI (sometimes called IAM) detector is used to measure residual ambient dose equivalent rate (Sv) during “beam-off”. During beam-on it provides a very good indication for the dose in air (Gy) caused by the given mixed radiation fields
Summary of “Radiation detector systems used to measure mixed radiation fields at workplaces”

## Slide 128
Merci Beaucoup !
Notes: 128

## Slide 129
Monte-Carlo simulations as a tool for detector response evaluation

## Slide 130: Motivation
High-energy mixed radiation fields
Radiation detectors must be characterized for these fields
Measured detector counts desired quantity (Sv, Gy)
High-energy hadron accelerator – LHC, SPS ...

## Slide 131
Detector response evaluation with Monte Carlo simulation tools
Simulation of irradiation situation + simulation response of detector
Simulation provides knowledge of particle fields and the response of the chamber to this radiation field
The following question remains: Does the simulation also reflect the reality?
Comparison between simulation and measurement = Benchmarking of simulation
Field calibration factor

## Slide 132
How?  A mixed hadron beam (in this setup 120 GeV/c) is intercepted by a copper target  high-energy mixed radiation field produced by EM and hadronic cascades
Hadron beam
Concrete shielding
Copper target
1 m
PMI: Air filled ionization chamber
IG5: High pressure ionization chamber
Benchmark experiments in the CERF radiation field to test reliability of Monte Carlo program FLUKA
What is CERF? An irradiation facility at CERN providing high-energy mixed radiation fields
Two ionization chamber types were irradiated with secondary particles produced in high-energy hadronic interactions (like beam loss in accelerator)
PMI chambers: exposed to high-energy particles occurring close to a target
IG5 chambers: exposed to the same radiation field, however attenuated by 80 cm of concrete.

## Slide 133
Detector benchmarking at the CERF facility: Copper target irradiated with a mixed hadron beam at 120 GeV/c
Hadron beam
~ 2 Gy/h
~10 Gy/h
~0.02 Gy/h
Dose rate at 1E8 protons/spill

## Slide 134
Simulation picture of the PMI chamber
Anode: PE / graphite coated
Active volume
Filling gas: air atmospheric pressure
Active volume: 3l
PE wall (4mm) / inside graphite coated
Wall composition: C-H2
Working voltage: ~460 V
Connector to cathode
Connector to anode
Connector plug for power supply and signal outlet
28.5 cm
21.5 cm
15.8 cm

## Slide 135: IG5 - Geometry
Properties
2 types (Ar or H filled)
5,2 l active volume
pressurized at 20 bar
1200 V high-voltage
Dimensions
diameter – 18.33 cm
height – 45.6 cm
electrodes
active medium
electronics
board
Notes: Electronics = 60% epoxy resin, 35% silicon, 2.5% copper, 2.5% carbon

## Slide 136: Way 1 (indirect approach)
B) Calculate particle fluence spectra [cm-2]
C) Combination of A and B
Simulated counting rate of detector
A) Calculate detector fluence response [C cm2]
Used to calculate IG5 response
2 ways to calculate detector response

## Slide 137: A) Simulation of Response Fuctions
Energy deposition in active volume
Calculate number of ion pairs created
Charge created within active volume
Response RF [C cm2]
Response RKa [C/Gy], RH*10 [C/Sv]
Conversion factors
Circular parallel beam
Notes: Parallel beam from circular disk, covers the whole chamber

## Slide 138: Response to var. particles (RF for Ar)

## Slide 139
Hadron beam
B) Calculation of fluence in the range of the measurement positions
CT6/T10
CT4
CS2
CS-50U

## Slide 140: B) Particle fluence at detector position (CT6/T10)

## Slide 141
C) Combination of A and B
Charge leads to counting rate of detector
x

## Slide 142: Convolution between fluence and response functions (CT6/T10)
Total contribution to response (Ar)
Neutron | Proton | p | g
(30 ± 1)% | (24 ± 3)% | (11 ± 1)% | (35 ± 4)%
Total contribution to response (H)
Neutron | Proton | p | g
(59 ± 3)% | (17 ± 2)% | (4 ± 1)% | (20 ± 2)%

## Slide 143
Ratio between simulation and experiment
Notes: CS2 – No additional shielding!! -> less response from calculation!!
CT4: The ratio is expected to be above 1 because the spectra calculated for the simple geometry do not match position T4 but a position directly above the target (T8/T12). The we'd expect more low energy neutrons and thus the comparision has to be done keeping this in mind!

## Slide 144
Experimental set-up in the CERF target area
Pos 1
Pos 2
Pos 3
Pos 4
Pos 5
Pos 6
Hadron beam
Cu target
Beam composition:
60.7 % p+
34.8 % protons
4.5 % K+
Momentum:
120 GeV/c
Intensity:
9*107 hadrons/ SPS extraction (16.8 s with 4.8 s continuous beam)
Beam parameters:

## Slide 145
Within this simulation calculation of energy deposition in active volume of chamber
“Energy to ion+/e-” conversion factor leads to number of produced ion+/e- pairs.
Conversion of number of ion+/e- pairs into pC.
One pC corresponds with one PMI counts.
FLUKA calculation of the whole particle cascade in the experimental setup
Way 2 (direct approach)
Used to calculate PMI response

## Slide 146
Pos 6
Pos 5
Pos 4
Pos 3
Pos 2
Pos 1
beam
Particle fluence at detector position 2
Analysis of the fluence reaching the various detector positions

## Slide 147
Pos 6
Pos 5
Pos 4
Pos 3
Pos 2
Pos 1
beam
Particle fluence at detector position 4

## Slide 148
Pos 6
Pos 5
Pos 4
Pos 3
Pos 2
Pos 1
beam
Particle fluence at detector position 6

## Slide 149
Simulation result of the counting rate
Pos 6
Pos 5
Pos 4
Pos 3
Pos 2
Pos 1
beam

## Slide 150
Influence of the different particle types (%) to the final counting rate of the detectors at the various positions

## Slide 151
Comparison between simulation and measurement results
 | Simulation Counts/ prim. part. *10-6 | Simulation error *10-6 | Measurement Counts/ prim. part. *10-6 | Measurement error *10-6 | Simulation/ Measurement | Error | 
Pos 1 | 5,63 | ± 0,12 | 5,64 | ± 0,56 | 0.998 | ± 0.102 | 
Pos 2 | 16,06 | ± 0,44 | 15,58 | ± 1,56 | 1.031 | ± 0.107 | 
Pos 3 | 67,46 | ± 0,73 | 67,25 | ± 6,93 | 1.003 | ± 0.104 | 
Pos 4 | 85,33 | ± 0,64 | 79,00 | ± 8,67 | 1.080 | ± 0.119 | 
Pos 5 | 96,20 | ± 1,26 | 89,39 | ± 9,47 | 1.076 | ± 0.115 | 
Pos 6 | 108,31 | ± 0,82 | 115,74 | ± 17,99 | 0.936 | ± 0.146 | 

## Slide 152: Summary of MC based calibration
FLUKA benchmarking experiments were performed at CERF
Very good agreement between simulation and measurement results in the radiation field occuring at the CERF facility.
The results prove that:
  FLUKA calculates mixed high-energy radiation fields correctly.
  FLUKA calculates detector response of ionisation chambers correctly.
FLUKA can be used to calculate a suitable field calibration factors for high-energy radiation fields occuring at CERN

## Slide 153
Backup slides

## Slide 154: ActiWiz Nuclide inventory optimization in accelerators

## Slide 155: Motivation for optimization of nuclear inventories of materials placed in accelerators
Safety benefit
Lower dose rates and committed doses
Operational benefit
Reduced downtime due to faster access
Less restrictions for manipulation & access
End of life-cycle benefit
Smaller amount and less critical radioactive waste
Smaller financial burden
Beside other material properties also the radiological consequences of the implementation of a material have to be considered
Level of activation depends on the type of the material
Choosing materials with low radiological impact results in several benefits

## Slide 156: Strategy to develop a tool allowing an optimization of nuclear inventories
Categorization of radiation environments
Development of ActiWiz – code assessing radiation risks, dominant nuclides etc., for arbitrary materials
Radiological hazard catalogue for materials

## Slide 157: Categorization of the radiation environments
FLUKA calculations of typical hadronic particle spectra (p, n, p+, p-) in CERN’s accelerators
LHC
SPS
PS
Linac 4 + Booster
160 MeV (Linac4), 1.4 GeV (Booster), 14 GeV/c (PS), 400 GeV/c (SPS), 7 TeV (LHC)
2400 single Monte Carlo simulations
 157.000 nuclide inventories (10 GB of data)
Used as base for the ActiWiz program

## Slide 158: ActiWiz – program interface
1.) Select energy / location / irradiation times
2.) Define material composition based on 69 chemical elements
* Many thanks to R. Froeschl for providing activation data on Zinc
Evaluate radiological hazard for arbitrary materials with a few mouse clicks

## Slide 159: Main output of ActiWiz: Material categorization
Radiological hazard assessment of material allowing for radiological comparison of materials
Aluminum 5083
Copper CUZnO5
Steel 316Ti
Operational
Waste

## Slide 160: Secondary output of ActiWiz: RP quantities Example
Information about ambient dose equivalent rate for various materials as a function of cooling time
For a given irradiation scenario we obtain:

## Slide 161: Some relevant units for radiation protection and radiation physics

## Slide 162: Quantities discussed
Flux
Fluence
Fluence rate or Flux density
Differential fluence
Current
Kerma
Absorbed dose
Equivalent dose
Effective dose
Ambient dose equivalent
Cross section
Surface density
Activity
Particle momentum versus particle energy
The description of these quantities are taken from the relevant ICRU and ICRP reports

## Slide 163: FLUX (g: Fluß)
N … number of particles
t … time
No surface through which particles traverse is considered

## Slide 164: Fluence (Flußdichte)
N …. Number of particles incident on a sphere of cross-sectional area da
a ... Cross section of an infinitesimal sphere surrounding point of interest
l … track length of particles traversing the infinitesimal sphere of volume dV
More general for macroscopic bodies: average fluence in a given body
For a sphere:
l … average cord length of a sphere
Fluence is a quantity that is proportional to effects such as induced activity, dose, radiation damage. The longer the integrated track length of particles through matter the higher the number of interactions inside the body

## Slide 165: Average fluence on a surface
For a given surface with a infinitesimal thickness of dt the following can be concluded:
dt

= =
Considering a constant Q
Quantity called current

## Slide 166: Differential fluence
Fluence per energy occurring in the energy interval [E, E+dE]

## Slide 167: Different ways to display differential fluence
=
Lethargy plots
Reflects better the real amount of particles around a given energy

## Slide 168
From
to
Derivation

## Slide 169: Current
Particles (N) crossing a given surface (A)
No weighting with cos(Q)
Pure counting of particles through a surface

## Slide 170: Fluence rate or Flux density
F … Fluence
N … number of particles
t … time

## Slide 171: Kerma (K)
is the abbreviation of kinetic energy released in matter. It reflects the sum of the initial kinetic energies dEtr of charged particles that are liberated by uncharged particles in a sample of matter, divided by the mass dm.
Kerma must not be mixed up with Absorbed dose, having the same unit (Gy).
Unit: J/kg = Gray (Gy)

## Slide 172: Absorbed dose
Unit: J/kg = Gray (Gy)
refers to the energy deposited (not released) in matter.
It reflects the sum of the energies dEdep deposited by incident particles in a sample of matter, divided by the mass dm of the sample.

## Slide 173: Equivalent dose in an organ or tissue, HT
ICRP publication 103
The radiation weighting factor (especially for neutrons) has been revised over time and remains controversial
Unit: Sievert (Sv)

## Slide 174: Effective dose, E
Organ | Tissue weigthing factor |  | Organ | Tissue weigthing factor
Gonads | 0.08 |  | Oesophagus | 0.04
Red Bone Marrow | 0.12 |  | Thyroid | 0.04
Colon | 0.12 |  | Skin | 0.01
Lung | 0.12 |  | Bone surface | 0.01
Stomach | 0.12 |  | Salivary glands | 0.01
Breasts | 0.12 |  | Brain | 0.01
Bladder | 0.04 |  | Remainder of body | 0.12
Liver | 0.04 |  |  | 
Different organs show different sensitivity to equivalent dose deposited
Unit: Sievert (Sv)

## Slide 175: Ambient-dose-equivalent, H*(10)
denotes the operational dose quantity used for area monitoring of penetrating radiation. Such a quantity is required since the effective dose is not directly measurable (different weighting factors for organs and particles). The H*(10) quantity is measured via the ICRU sphere:
ICRU sphere: A sphere of 30 cm diameter made of tissue equivalent material with a density of 1 g/cm3 and a mass composition of 76.2% oxygen, 11.1% carbon, 10.1% hydrogen and 2.6% nitrogen
Dose equivalent in this point (10 mm inside the sphere)
30 cm
Unit: Sievert (Sv)
Ambient dose equivalent is a conservative measure for effective dose.

## Slide 176: Cross section
The cross section (s) of a target entity, for a particular interaction produced by incident particles is defined as:
P … probability of interaction occurring in a given volume
F fluence through the given volume
Unit of the cross section is m2. A special unit for the cross section used is barn, which is defined by:
1 barn= 10-28 m2= 10-24 cm2
Using the definition of the fluence applying the track length allows to use each target shape and size

## Slide 177: Surface density (Flaechendichte)
= =
Unit: kg/m2
Mass per unit area:
Explanation: Mass along a straight line starting at a given surface normalized to the size of the surface
A
l
rA … surface density
r … density
If r = const.

## Slide 178: Activity
Decays of a radioactive material per time unit
Unit: Bq:  1 Bq = 1 decay per second

## Slide 179: Particle momentum versus particle energy
E in GeV
p in GeV/c
m in GeV
Ekin= E – Energy equivalent of particle
Don’t forget considering mass:
E … total energy
p … momentum
m… mass of particle
m0… mass of particle at rest
v… velocity of particle
c… speed of light in vacuum
+
+