## Slide 1: FCC-ee Twin quadrupole prototype magnetic axis shift
J. Bauche (TE-MSC)
FCC-ee beam optics meeting, 27th November 2020

## Slide 2: FCC-ee quadrupole design
Magnetically coupled design
SPACER
Design key feature:
non-magnetic spacer balances the flux between the poles
Courtesy A. Milanese
CDR parameters FCC-ee quadrupoles

## Slide 3: FCC-ee quadrupole – prototype
Short model, ~1 m long (900 mm yoke)
Yoke assembled from machined solid iron (ARMCO)
Assembly tolerances ±20 µm (aperture diameters)
Coils 2 x 64 turns (converter matching)

## Slide 4: Prototype measurements – transfer function
Very good symmetry between the two apertures (diff. < 1.6‰)
Pessimistic BH curve on FEM (µr), diff ≈1%
FCC-ee range

## Slide 5: Prototype measurements - magnetic axis shift
~0.4 mm shift outwards for each aperture between low and high fields
Mismatch MM vs. FEM at low fields, not completely explained. Refined simulations needed as a next step.
Chart: Current [A]
| 25 | 50 | 100 | 150 | 200 | 250 | 200 | 150 | 100 | 50 | 25
ΔXAP1 MM | 0.75 | 0.22 | -0.07 | -0.17 | -0.22 | -0.29 | -0.23 | -0.18 | -0.1 | 0.15 | 0.59
ΔXAP2 MM | -0.75 | -0.23 | 0.07 | 0.16 | 0.22 | 0.27 | 0.22 | 0.17 | 0.09 | -0.16 | -0.59
ΔXAP2 FEM | 0.17 | 0.17 | 0.17 | 0.17 | 0.18 | 0.22 | 0.18 | 0.17 | 0.17 | 0.17 | 0.17
FCC-ee range

## Slide 6: Mitigations of magnetic centre shift
Active compensation with trim coil (busbars) in spacer  to be simulated
Use dipole trim coils for tapering (E saw tooth mitigation) to add dipolar component
Modify passive shimming at extremities at each change of energy  to be measured at manufacture for all energies