Slide_1
MPP Meeting
BLMINJ Software
David Medina Godoy, BE-BI-SW

---
Slide_2
Outline
System Overview
Interlocks
Crate Configuration & FESA instantiation
Deployed System Configurations

---
Slide_3
System Overview
- Provides several type of losses per cycle:
  - Running sums (2us, 600us and 1ms)
  - Losses over cycle
  - Losses over beam presence
  - Ambient losses
  - Evolution of the losses over the cycle
    - Ring configuration (PSB, PS) : 1ms/sample
    - Transfer line configuration (LN4, TT10, ...) : 2us/sample
  - Capture of the losses with adjustable ...
    - Start of recording event :
      - Basic Period, Beam Start, Beam End or Start of Cycle
    - Delay after start event
    - Sample Resolution [2us - 1.2s]
  - Provides diagnostics data for the experts

---
Slide_4
Interlocks
  - Software thresholds defined per loss type, monitor and cycle
    - Evaluation of possible SW interlock at the end of the cycle
    - Bad shots allowed before interlock counter setting provided
    - Possibility to mask software interlock per loss type, monitor and cycle
  - Hardware thresholds defined per loss type and monitor (non multiplexed)
    - Non maskable
    - Evaluation of possible HW interlock is immediate
  - RBAC/MCS protected settings:
    - HW/SW thresholds, bad shot counter and masks
  - Provided alarms include:
    - SW Interlock (ERROR)
    - HW interlock (ERROR)
    - Beam Presence loss type masked (WARNING)

---
Slide_5
Crate configuration & FESA Instantiation
Per crate:
1 combiner card with BP, BIN, BOUT, SCY timing inputs
(Max of) 8 acq. & proc. cards (64 monitors)
Problem: Crates with monitors that belong to different acc. zones (InCA settings should be split).
Solution: Family of monitors
1 FESA device per family in the crate
Max of 8 families/devices allowed (1 per card)
(Our standard configuration is 1 device per crate)
In cfv-361-blmbt we instantiate 4 devices:
BT.BLM (BE_BT acc. zone)
BTM.BLM (BTM_DUMP acc. zone)
BTP.BLM (BTP acc. zone)
BTY.BLM (BTM_BTY acc. zone)
cfv-361-blmbt

---
Slide_6
Deployed Configurations
PS RING
LN4 & PSB
TT10 (SPS)

---
Slide_7
Deployed Configurations
PS RING
LN4 & PSB
TT10 (SPS)

---
Slide_8
Deployed Configurations
PS RING
LN4 & PSB
TT10 (SPS)

---
Slide_9
Deployed Configurations
PS RING
LN4 & PSB
TT10 (SPS)

---
Slide_10
Deployed Configurations
PS RING
LN4 & PSB
TT10 (SPS)

---
Slide_11
Deployed Configurations
PS RING
LN4 & PSB
TT10 (SPS)

---
Slide_12
Deployed Configurations: LN4, PSB and PS Ring
Combiner card connected to BIS
BLMINJ_VIRT
1 Monitor = 1 Virtual Device
BLMINJ Expert GUI
BIS select configuration
BIBLMApp (ComRAD)
BLMINJ Virtual Settings
Courtesy: Yu Wu

---
Slide_13
Deployed Configurations : PS Ring
  - BLMSYNC:
    - Concentrate losses in PS ring (i.e. 2 vme crates)
    - Sum of the losses of all monitors interlock
      - Threshold, mask and counter settings provided
    - Monitors hw/sw interlock status from BLMINJ
    - Interlock capabilities acting on tail clipper LTIM
    - (Extra) Alarms added:
      - sumOfLossesMasked (WARNING)
      - sumOfLossesThresholdBreached (ERROR)
      - syncAlarm (ERROR)
PR.BLM
(cfv-359-blmpra)
BIX.TCLPS-BLM

---
Slide_14
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_15
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_16
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_17
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_18
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_19
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_20
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_21
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_22
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_23
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_24
Deployed Configurations : PS Ring
BLMSYNC Evolution Vistar:
To be used in PS BLM Vistar fixed displays
Calculates the sum (so far in the cycle) of the losses in the Evolution data
Publishes every 100ms ( 10Hz )
PS BLM Vistar
(currently in development)
Courtesy: Denis Gerard Cotte

---
Slide_25
Deployed Configurations : TT10 (SPS)
  - Only interested in injection losses
  - Summary Acquisition added:
    - Provides a summary of the injections in the cycle
  - Special timing configuration:
    - BP = SIX.W20-CT + 22ms (2ms after injection)
    - SCY = SIX.F1KLO-CT + 1002 ms (2ms after last inj.)
publish data
inj1 inj2

---
Slide_26
Deployed Configurations : TT10 (SPS)
Interlock:
BLMINJ publishes swInterlock flag after every injection
If swInterlock = true for injection
Further ones in the cycle will be inhibited by SIS
swInterlock will be cleared for the next cycle
If threshold breached for 3 cycles in a row:
SIS inhibits the beam

---
Slide_27
Expert Application & Diagnostics

---
Slide_28
Expert Application & Diagnostics

---
Slide_29
Expert Application & Diagnostics

---
Slide_30
Expert Application & Diagnostics

---
Slide_31
Expert Application & Diagnostics

---
Slide_32
Expert Application & Diagnostics

---
Slide_33
Expert Application & Diagnostics

---
Slide_34
Expert Application & Diagnostics

---
Slide_35
Expert Application & Diagnostics

---
Slide_36
Expert Application & Diagnostics

---
Slide_37
Expert Application & Diagnostics

---
Slide_38
Expert Application & Diagnostics

---
Slide_39
Expert Application & Diagnostics

---
Slide_40
Expert Application & Diagnostics

---
Slide_41
Expert Application & Diagnostics

---
Slide_42
Expert Application & Diagnostics

---
Slide_43
Conclusion
Highly configurable system
Adapted to the needs of HW experts and operators
System validated in LN4 and PSB dry runs

---
Slide_44
THANKS!!
Questions?