## Slide 2
Vacuum Systems
Lecture 1
V. Baglin
CERN TE-VSC, Geneva

## Slide 3: Vacuum? A transverse field across several engineering and scientific disciplines
Complementary information

## Slide 4
Outline
Introduction
Gas kinetic theory & gas flow
Measurement & production of vacuum

## Slide 5
1. Introduction

## Slide 6
Vacuum
“Perfect vacuum” does not exist on earth nor in space!
Interstellar medium in a galaxy such as the Milky Way:
  Composed of molecules, ions atoms, cosmic rays & dust (size 0.1 micron)
  In molecular clouds, which are cold (>10 k) and dense region n ~ 104 molecules/cm3
  Atoms density:
    50 H/cm3 at 100 K (~10-13 Pa)
    1 H/cm3 at 10 000 K (~10-13 Pa)
Herschel: molecules in the Orion nebula
Pilar of the creation
Nasa-ESA Hubble

## Slide 7
Vacuum Systems on Earth
Vacuum technology is present in many devices/ systems
Light bulb
0.7 atm of Ar
Vacuum pump
For wine
Electron beam welding machine
10-5 – 10-2 mbar
Semiconductor industry
10-6 – 10-2 mbar depending on process
Under vacuum brazing
10-7-10-6 mbar at 200-1300°C
Accelerator
The objective of Vacuum is to reduce the collision of molecules on the surrounding environment to preserve the quality of the process
Design: 1015 H2/m3 i.e. 2 10-9 mbar at 15 K
(~10-8 mbar at 300 K)

## Slide 8
Vacuum gauges & (capture) pumps
P. A. Redhead, Vacuum 53 (1999) 137-149
1964: 7 10-15 Torr
obtained with a multi-baked Aluminosilicate glass finger immersed in liquid helium
Modulated BA gauge
P. A. Redhead et al, Can. J. Phys. 40,1814 (1962)
J.P Hobson
P.A Redhead
C. Benvenuti
1993: ~ 10-14 Torr
Passively activated NEG at 350 °C
C. Benvenuti et al, Vacuum 44,511-13 (1993)
ST707 NEG mounted on a frame & Helmer gauge

## Slide 9
Lowest (measured) pressure
Recently obtained at the BASE experiment (the Baryon-Antibaryon Symmetry Experiment)
By measuring the decay of anti-protons due to beam-gas scattering, the estimated pressure in the cryogenic Penning trap is ~ 10-19 mbar i.e. 2 400 molecules/m3
BASE experiment implemented into the AD facility
BASE, Eur. Phys. J. Special Topics 224, 3055-3108 (2015)

## Slide 10
Measure of beam-gas scattering, using beam loss monitors
Pressure in the LHC arcs during RUN2
 | Design | 2015 | 2018
Life time | 100 h | 185 days | 4,5 years !!
2 10-11 mbar
2 10-10 mbar

## Slide 11
2. Gas kinetic theory
& gas flow

## Slide 12
2.1 Gas kinetic theory

## Slide 13: Introduction
Assume a large number of molecules, always moving in a disordered manner
The size of molecules is very small as compared to the intermolecular distance
Molecules moves in a rectilinear uniform manner between successive collisions
Collisions (elastic) are against molecules themselves or against the wall
The molecular trajectories are broken lines
http://www.matierevolution.fr

## Slide 14: Maxwell Boltzmann Distribution
Assume a pure gas, in thermal equilibrium and enclosed in an isothermal volume
In this case:
  The molecule density is constant in the volume and do not vary in time
  The direction of the molecule’s speed is uniform
  The speed distribution is stationary
The speed of the molecules follows a Maxwell-Boltzmann distribution
Most of the molecules have a speed around the maximum, vp
Less than 1/1000 of molecules have speed:
  v<0.1 vp or v>3 vp
vp
Maxwell 1860
Boltzmann 1870

## Slide 15: Maxwell Boltzmann Distribution
The most probable speed is given at the maximum of the distribution: d(dn/n)/dv =0. It equals
The mean thermal speed equals:
The average quadratic speed equals:
The corresponding kinetic energy is:
All the molecular speeds scales like ~  (T/m)

## Slide 16: Maxwell Boltzmann Distribution
The gas velocity:
  increase with increasing temperature
  is the largest for light molecules
~  (T/m)

## Slide 17: Average velocity
~  (T/m)
T (K) | He | Air | Ar
4.2 | 150 | 55 | 50
300 | 1300 | 470 | 400
600 | 1800 | 660 | 560
The mean thermal speed range from ~ 100 till 1000 m/s:
  The traveled distance in a second is much larger than the vacuum chamber dimensions!

## Slide 18: Collision on the wall
The molecular collision rate on the wall (the incidence rate), , can be derived from the Maxwell Boltzmann distribution
Monolayer formation time
Ice formed at atm pressure
At room temperature, a monolayer is formed in 1 s at 10-6 Torr : Langmuir formation time
1 L = 1 monolayer.s
Very low pressure is preferred to minimize surface contamination,
A monolayer is formed in:
1 ns at 1 atm
1 s at 10-6 mbar
1 day at 10-11 mbar !!!
1015 molecules/cm2
LHC beam pipe is under vacuum

## Slide 19: Pressure & Ideal gas law
Molecules which collide a wall of surface A, produce a force, F, onto it.
The pressure P, exerted on the wall by the molecules is defined by the ratio of the force to the surface:
It can be shown that, for any particle velocity distribution, the pressure is given by the mass density and the quadratic speed:
In the special case the particle velocity distribution, follows a Maxwell Boltzman distribution, the pressure is given by the IDEAL GAS LAW:
For such a gas, the pressure, P [Pa], is defined by the gas density, n [molecules.m-3] , the temperature of the gas, T [K] and the Boltzman constant k , (1.38 10-23 J/K)
The pressure increases linearly with the gas temperature

## Slide 20: Ideal gas law: illustration
Ultra-High Vacuum
  take LHC design
  n= 1015 molecules/m3
    k = 1.38 10-23 J/K
    T = 300 K i.e. room temperature
    P = 1015 x 1.38 10-23 x 300 = 4 10-6 Pa
The ideal gas law relation is another expression on the Avogadro’s law:
  Molar volume
  The occupied volume by one mole in standard condition equals 22.4 ℓ (i.e. 22.4 10-3 m3)
     n = N / V = 6.02 1023 / 0.0224 m3 = 2.7 1025 molecules/m3
  Atmospheric pressure at 0°C (standards conditions)
    P = 101 300 Pa,
    k = 1.38 10-23 J/K
    T = 273 K i.e. 0°C
    n = 101 300 / (1.38 10-23 x 273) = 2.7 1025 molecules/m3

## Slide 21
Units
The pressure is the force exerted on the wall by the molecules per unit of surface :
1 Pa = 1 N/m2
 | Pa | kg/cm2 | Torr | mbar | bar | atm
1 Pa | 1 | 10.2 10-6 | 7.5 10-3 | 10-2 | 10-5 | 9.81 10-6
1 kg/cm2 | 98.1 103 | 1 | 735.5 | 980 | 0.98 | 0.96
1 Torr | 133 | 1.35 10-3 | 1 | 1.33 | 1.33 10-3 | 1.31 10-3
1 mbar | 101 | 1.02 10-3 | 0.75 | 1 | 10-3 | 0.98 10-3
1 bar | 1.01 105 | 1.02 | 750 | 103 | 1 | 0.98
1 atm | 101 300 | 1.03 | 760 | 1 013 | 1.01 | 1
Examples:
  10-6 Torr is converted to 1.33 x 10-6 = 1.3 10-6 mbar
  10-6 Torr is converted to 133 x 10-6 = 1.3 10-4 Pa
  4 10-6 Pa is converted to 4 10-6 / 133 = 3 10-8 Torr
  4 10-6 Pa is converted to 4 10-6 / 100 = 4 10-8 mbar

## Slide 22
Force applied on a vacuum vessel
When the vacuum vessel is evacuated, a force is applied onto it
It amounts to 1 kg/cm2
Example: force applied on a blank flange
As a consequence of the « vacuum force » …
D (mm) | 16 | 35 | 63 | 80 | 100 | 130 | 150 | 212
kg | 2 | 10 | 32 | 52 | 81 | 137 | 182 | 363
Picture Kurt J. Lesker

## Slide 23
“Vacuum force”: Magdeburg hemisphere
1654, Regensburg
55 cm diameter  25 Tons

## Slide 24
Work with the Mechanical Design Office !
0.5 ton
6.6 tons
Otherwise…
60

## Slide 25: Typical accidents with UHV!
Case of the CERN ISR in the 70’s :
Imploded “flat fish” at an ISR intersection
“spontaneous” breaking of a bellow (either due to a bad design or to a fixed point not well attached)

## Slide 26: Even in modern times …
Accident still possible nowadays even with modern computing tools …
Case of the QRL’s bellows in the LHC deformed during pump down,
Origin attributed to a non-conform bellow with a too small corrugation high

## Slide 27
Total Pressure and Partial Pressure
The gas is usually composed of several types of molecules (ex : air, residual gas in vacuum systems)
The total pressure, PTot, is the sum of all the partial pressure, Pi (Dalton law)
John Dalton, 1801

## Slide 28
2.2 Gas flow

## Slide 29
Mean Free Path
It is the path length that a molecules traverse between two successive impacts with other molecules. It was derived by Clausius.
It is a function of the pressure, P, of the temperature, T, and of the molecular diameter, σ.
Molecule | Diameter (Å)
H2 | 2.8
H2O | 2.9
O2 | 2.9
N2 | 3.7
CO2 | 3.2
R. Clausius

## Slide 30
Mean Free Path: air at room temperature
At atmospheric pressure, λ = 70 nm
At 1 Torr, λ = 50 μm
At 10-3 Torr, λ = 5 cm
At 10-7 Torr, λ = 500 m
At 10-10 Torr, λ = 500 km
Increasing mean free path
when decreasing pressure

## Slide 31
Classification of vacuum
From atmospheric pressure to very low pressure, the mean free path varies over more than 10 orders of magnitude!
Low or Rough vacuum
  105 to 102 Pa ; 103 to 1 mbar
Medium vacuum
  102 to 10-1 Pa ; 1 to 10-3 mbar
High vacuum (HV)
  10-1 to 10-5 Pa ; 10-3 to 10-7 mbar
Ultra-high vacuum (UHV)
  10-5 to 10-10 Pa ; 10-7 to 10-12 mbar
Extreme-high vacuum (XHV) (below actual limit of “standard” instrumentation)
  <10-12 Pa ; <10-12 mbar

## Slide 32
Flows
The turbulent flow is established around the atmospheric pressure
In the low vacuum (103-1 mbar), the flow is viscous and laminar.
In the high vacuum (10-3 – 10-7 mbar) and ultra-high vacuum (10-7–10-12 mbar), the flow is molecular. The mean free path is much larger than the vacuum chamber diameter. Molecules interact only with the vacuum chamber walls
Molecular flow is the main regime of flow to be used in vacuum technology
In this regime, the vacuum vessel has been evacuated from its volume.
The pressure inside the vessel is dominated by the nature of the surface.

## Slide 33
Molecular Flow: domain of application
Assume an accelerator ring operating under UHV:
  vacuum chamber diameter ~ 10 cm
  average pressure 10-8 mbar (= 7.5 10-9 Torr)
     PxD = 7.5 10-9 x 10 = 7.5 10-8 Torr.cm  molecular regime
  mean free path:
     λ = 5 10-5 / 7.5 10-8 = 667 m
Assume a large vacuum vessel, e.g. a large cryostat, operating under UV:
  vacuum chamber diameter ~ 10 m (=1000 cm)
  average pressure 10-5 mbar (= 7.5 10-6 Torr)
     PxD = 7.5 10-6 x 1000 = 7.5 10-3 Torr.cm  molecular regime
  mean free path:
     λ = 5 10-5 / 7.5 10-6 = 7 m

## Slide 34: Molecule interaction with the wall
In the molecular regime, Knudsen observed that the speed of the molecules is reduced due to interactions with the pipe
Following the collision on the wall, owing to the roughness of the surface, the molecule is re-emitted into the vacuum system according to a Beer-Lambert law.
This observation introduced the concept of conductance (Dushman)
Knudsen number
Saul Dushman (1883–1954)

## Slide 35
Conductance
It is defined by the ratio of the molecular flux, Q, to the pressure drop along a vacuum vessel. It is a function of the shape of the vessel, the nature of the gas and its temperature.
Adding conductances in parallel
Adding conductances in series
P1
P2
Q
P1
P2
Q
Q
C1
C2
Q
P1
P2
C1
C2

## Slide 36
Conductance Calculus in Molecular Regime
For a (thin) orifice :
The conductance of an orifice of 10 cm diameter is 900 l/s
For a tube :
The specific conductance of a tube of 10 cm diameter is 120 l/s.m
To increase the conductance of a vacuum system, it is better to have a vacuum chamber with large diameter and short lenght

## Slide 37
Scaling of conductances
Gas | M | C_orifice [l/s] | C_tube [l/s.m]
Air | 29 | 900 | 120
H2 | 2 | 3427 | 457
CH4 | 16 | 1212 | 162
CO | 28 | 916 | 122
CO2 | 44 | 731 | 97
The conductance scales like:
at room temperature

## Slide 38
Pumping Speed
The pumping speed, S, is the ratio of the flux of molecules pumped to the pressure
S range from 10 to 20 000 l/s
Q range from 10-14 mbar.l/s/cm2 for metalic tubes to 10-5 – 10-4 mbar.l/s/cm2 for plastics
l/s
mbar.l/s
mbar
3 orders of magnitude for pumping
vs
10 orders of magnitude for outgassing
Outgassing MUST be optimised to achieve UHV

## Slide 39: Effective pumping speed
P
P’
Q, C
S
It is the pumping speed seen from P’ trough the pipe of conductance, C
This is the result of adding in series the conductance C with the pumping speed S
If:
  C=S then Seff = S/2
  C>> S then Seff = S
  C<< S then Seff = C, the system is “conductance limited”
Large conductances preserve the efficiency
of the pumping system
Pumping speed
Conductance

## Slide 40: Effective pumping speed
Assume a 10 m long 10 cm diameter vacuum pipe evacuated by a turbomolecular pump of 60 l/S
Seff, P
C
S
C = 120 / 10 = 12 l/s
S = 60 l/S
Seff = (12x60)/(12+60) = 10 l/s
The system is conductance limited and the (effective) pumping speed, hence pressure, at the upstream part is defined by the conductance of the pipe!

## Slide 41: Differential pumping system
This system allows to decouple a vacuum system from another one (e.g. in Linac source, between baked and unbaked vacuum system)
P3
Q1, C
Q2, S
P2
Q, C
Q4, C
Q3, S
P1
P4
Commercial differential pumps
Pictures Edwards
10-1
UHV

## Slide 42
3. Measurement &
production of vacuum

## Slide 43
3.1 Vacuum gauges

## Slide 44: Vacuum gauges pressure range
Pa
104
102
10
10-2
10-4
10-6
10-8
10-10
10-12
mbar
102
10
10-2
10-4
10-6
10-8
10-10
10-12
10-14
Capacitance
Spinning rotor
Piezo
Pirani
Penning
Bayard-Alpert
Helmer, extractor
16 orders of magnitude !

## Slide 45
Pirani Gauge
Pirani gauges are used in the range 1 atm -10-4 mbar to monitor the roughing of a vacuum system
Robust, gas dependent, accuracy 10-100 %
Operating principle based on the variation of the heat conductivity trough a gas
   pressure reading above 1 mbar are wrong
Inficon Pirani gauge
Pfeiffer Pirani gauge
Edwards Pirani gauge
wikipedia
M. Pirani 1906
Tungsten wire

## Slide 46
Penning Gauge
Penning gauges are commonly used in the range 10-5 -10-10 mbar. They are use for interlocking purposes
Robust, gas dependent, accuracy ~ 20-50 %
It is a cold cathode ionisation gauge i.e. there are no hot filament: electron are produced by field emission
The operating principle is based on the measurement of a discharge current in a Penning cell which is a function of pressure : I+ = Pn, n is close to 1
F. Penning 1937
B
Pfeiffer Penning gauge

## Slide 47: A discharge in a Penning gauge
Penning gauge ON behind a window
In the dark, the discharge is seen
Pictures courtesy B. Henrist, TE-VSC
Courtesy Pfeiffer
a Penning gauge:

## Slide 48
Bayard-Alpert Gauge
Bayard-Alpert gauges are used for vacuum measurement purposes in the range 10-6 -10-12 mbar.
It is a hot filament ionisation gauge. Electrons emitted by the filament perform a few oscillations inside the grid and ionise the molecules of the residual gas. Ions are then collected by an electrode.
Ion collector = 0 V
Filament = + 50 V
Grid = + 150 V
Modulator = + 150 V
L.G. Pittaway. J. of Phys. D: Appl. Phys. 3, 1113-1121, 1970
Courtesy B. Jenninger

## Slide 49: A burned filament
Pictures courtesy B. Henrist, TE-VSC
Obviously, even is there are 2 filaments, the gauge is polluted therefore the pressure measurement will not be correct !
It is wise to exchange the vacuum gauge

## Slide 50
Bayard-Alpert Gauge: Sensitivity
The ion current collection can be described by:
The gauge needs to be calibrated for several gases
SN2 ~ 40 mbar-1
The pressure reading is expressed in nitrogen equivalent
In UHV, typical collected current are in the pA range
Where :
I+ is the ion current
I- is the filament current
σ is the ionisation cross section
n the gas density
L the electron path length
The vacuum gauge sensitivity, S in mbar-1, is defined by:
Ie (mA) | P (mbar) | I+ (pA)
4 | 10-10 | 16
4 | 10-12 | 0.16
0.1 | 10-10 | 0.4
0.1 | 10-12 | 0.004

## Slide 51
Bayard-Alpert Gauge: Sensitivity
The sensitivity can be measured and also computed from simulations
B. Jenninger et al. Vacuum 138 (2017) 173-177
 | H2 | He | CH4 | Ne | N2 | CO | C2H6 | Ar | CO2 | Xe
Si (mbar-1) | 19.06 | 7.46 | 60.62 | 10.48 | 41.84 | 42.30 | 114.71 | 53.19 | 54.48 | 7.50
Srel,i (mbar,i/mbar,N2) | 2.20 | 5.61 | 0.69 | 3.41 | 1.00 | 0.99 | 0.36 | 0.79 | 0.77 | 4.83
The sensitivity relative error equals ~ 10 %
A pressure reading in N2 equivalent is: Pread =2 10-10 mbar, in the case the main molecular species is H2, the real pressure is: P = Srel,H2 x Pread = 4.4 10-10 mbar
Relative sensitivity:
Space charge
effect
4 mA

## Slide 52
3.2 Gas analysis

## Slide 53
Residual Gas Analysers
Residual Gas Analysers are used in the range 10-4 -10-13 mbar. Their purpose is to do gas analysis
A filament produces electrons which ionise the residual gas inside a grid.
A mass filter is introduced between the grid and the ion collector.
The ion current can be measured in Faraday mode or in secondary electron multiplier mode.
G.J. Peter, N. Müller. CAS Vacuum in accelerators CERN 2007-003
Range: 10-14 till 10-5 A
Gain ~ 103 -104 electrons/ion
ΔM at FHWM = 0.5 AMU
1 mA
Q ~ 2 10-9 mbar.l/s
Picture Pfeiffer

## Slide 54
RGA: Cracking pattern
Ions produced inside the grid can be fragmented into sub-species by the collisions with electrons
Example Mass 28
Nitrogen is traced by mass:
  28 (N2+)
  14 (N+): I14 = 0.14 I28
Carbon monoxide is traced by mass:
  28 (CO+)
  12 (C+): I12 = 0.06 I28
Carbon dioxide is traced by mass:
  44 (CO2+)
  28 (CO+): I28 = 0.13 I44
  16 (O+): I16 = 0.16 I44
  12 (C+): I12 = 0.10 I44
The table gives the percentage of the fragments with respect to the main peak

## Slide 55
RGA: Spectrum
A typical spectrum of a baked vacuum system: P = 10-10 mbar
28
44
40
2
15
12
16
19
M/e | Ion | Molecule
2 | H2+ | H2
15 | CH3+ | CH4
28 | CO+ | CO
40 | Ar+ | Ar
44 | CO2+ | CO2
12 | C+ | CO+CO2
16 | O+ | Filament artefacts
19 | F+ | 
H2
Di-Hydrogen
M/e=2;1
CH4
Methane
M/e=15;(16)
H2O
Water
M/e=18;17;16
CO
Carbon monoxide
M/e=28;12
CO2
Carbon dioxide
M/e=44;16;28;12
N2
Nitrogen
M/e=28;14
Ar
Argon
M/e=40;20
O2
Oxygen
M/e=32;16

## Slide 56
RGA: Partial Pressure
The RGA needs to be calibrated against a total pressure gauge for standard gases.
To consider the RGA ageing, relatives sensitives are used and a normalization factor, K, is introduced
 | H2 | CH4 | N2 | CO | Ar | CO2
Srel,i,RGA (Torr N2/Torr) | 1.09 | 0.93 | 0.99 | 1 | 1.18 | 1.96
According to the Dalton’s law, the reading given by the total pressure gauge shall equal the sum of the partial pressure, expressed in nitrogen equivalent:
So:
Therefore:

## Slide 57
Typical spectrum of a baked vacuum system
Baked system with P= 10-10 mbar N2 eq
 | H2 | CH4 | CO | Ar | CO2
I (A) | 1 10-9 | 2 10-12 | 7 10-11 | 5 10-13 | 2 10-11
P (mbar) | 2 10-10 | 1 10-13 | 6 10-12 | 4 10-14 | 3 10-12
% Pi | 96 | 0 | 3 | 0 | 1
28
44
2
Note: a simple estimation from the total pressure measurement would give P= 2.2 10-10 mbar!

## Slide 58
Typical spectrum in the presence of an air leak
RGA are also useful to identify, trace and estimate air leaks
Air leak
M/e | Ion | Molecule
14 | N+ | N2
28 | N2+ | N2
32 | O2+ | O2
40 | Ar+ | Ar
Assume a 8 cm diam. tube with specific conductance ~ 60 ℓ.s.m:
  a leak detected 10 m away from the gauge (at P= 8 10-7 mbar)
  Hence, a pumping speed of 6 ℓ/s at the level of the leak gives a leak rate of:
    6 x 8 10-7 = 5 10-6 mbar.ℓ/s (i.e.)
Beware, oxygen being highly chemically reactive, its mass is not always present in the spectrum!

## Slide 59
3.3 Vacuum pumps

## Slide 60: Vacuum pumps pressure range
Pa
104
102
10
10-2
10-4
10-6
10-8
10-10
10-12
mbar
102
10
10-2
10-4
10-6
10-8
10-10
10-12
10-14
Oil sealed mech. pump
Turbo molecular pump
Sputter ion pump
Cryopumps
Ti sublimation pump
NEG & cryogenics pumps
16 orders of magnitude !

## Slide 61
Turbomolecular pumping group
Used to pump down from atmosphere and commission vacuum sectors down to 10-11 mbar
Mobile system based on rotary vane primary pump and turbomolecular pump
Valve
Turbomolecular pump
Primary pump
Primary pump
Turbomolecular pump
Primary pump

## Slide 62
Sputter Ion Pump
This pump operate in the range 10-5 -10-11 mbar. It is used to maintain the pressure in the vacuum chamber of an accelerator.
Pumping speed range from 1 to 500 l/s.
Titanium sputtered from the cathode bombarded by accelerated ions provides pumping.
The ion current is proportional to pressure, hence these pumps are used for interlocks
Picture Agilent, Varian

## Slide 63: Lecture 1 summary
The kinetics of gas molecules is described by a Maxwell-Boltzmann distribution.
The pressure is defined by the ideal gas law.
As a function of mean free path, several regime of flow exits.
The molecular flow is the regime of vacuum systems: the molecules interact only with the vacuum chamber wall.
The conductance characterise the pressure drop along a vacuum component
The pressure in a vessel is defined by the ratio of the gas flow to the pumping speed
Main instruments for vacuum measurement are: Pirani, Penning, Bayart-Alpert gauges and residual gas analysers
Main devices for vacuum pumping are: primary, turbomolecular and ion pumps.

## Slide 64
Some References
Cern Accelerator School, Vacuum technology, CERN 99-05
Cern Accelerator School, Vacuum in accelerators, CERN 2007-03
Cern Accelerator School, Vacuum for particle accelerators, CERN-ACC-2020-0009
The physical basis of ultra-high vacuum, P.A. Redhead, J.P. Hobson, E.V. Kornelsen. AVS.
Scientific foundations of vacuum technique, S. Dushman, J.M Lafferty. J. Wiley & sons.
Les calculs de la technique du vide, J. Delafosse, G. Mongodin, G.A. Boutry. Le vide.
Vacuum Technology, A. Roth. Elsevier Science
Foundations of vacuum science and technology, Ed by J.M. Lafferty. J. Wiley & sons.
Some Journals Related to Vacuum Technology
Journal of vacuum science and technology
Vacuum

## Slide 65
Thank you for your attention !!!

## Slide 67
Complementary
informations

## Slide 68
2.1 Gas kinetic theory

## Slide 69: Collision on the wall
The molecular collision rate on the wall (the incidence rate), , can be derived from the Maxwell Boltzmann distribution
Monolayer formation time
H2 at 300 K
At room temperature, a monolayer is formed in 1 s at 10-6 Torr : Langmuir formation time
1 L = 1 monolayer.s
Very low pressure is preferred to minimize surface contamination:
1 ns at 1 atm
1 s at 10-6 mbar
1 day at 10-11 mbar !!!
1015 molecules/cm2
Complementary information

## Slide 70: Collision on the wall: cryogenic temperature
The molecular collision rate on the wall (the incidence rate), , can be derived from the Maxwell Boltzmann distribution
H2 at 4.2 K
n (H2/m3) | V (H2/cm2/s) | Time
1010 | 5 107 | 220 d
1011 | 5 108 | 22 d
1012 | 5 109 | 2
1013 | 5 1010 | 5 h
1014 | 51011 | 30 min
1015 | 5 1012 | 3 min
Complementary information

## Slide 71
2.2 Gas flow

## Slide 72
Turbulent and Viscous Flows
When pumping down from atmospheric pressure, the physics is characterised by different flow regimes. It is a function of the pressure, of the mean free path and of the components dimensions.
Reynolds number, Re :
  if Re > 2000 the flow is turbulent
  it is viscous if Re < 1000
The turbulent flow is established around the atmospheric pressure
In the low vacuum (103-1 mbar), the flow is viscous. The flow is determined by the interaction between the molecules themselves. The flow is laminar. The mean free path of the molecules is small compared to the diameter of the vacuum chamber
Complementary information

## Slide 73
Transition and Molecular Flows
In the medium vacuum (1-10-3 mbar), the flow is transitional. In every day work, this range is quickly crossed when pumping down vacuum chambers. In this regime, the calculation of the conductance is complex. A simple estimation is obtained by adding laminar and molecular conductances.
In the high vacuum (10-3 – 10-7 mbar) and ultra-high vacuum (10-7–10-12 mbar), the flow is molecular. The mean free path is much larger than the vacuum chamber diameter. The molecular interactions do not longer occurs. Molecules interact only with the vacuum chamber walls
Molecular flow is the main regime of flow to be used in vacuum technology
In this regime, the vacuum vessel has been evacuated from its volume.
The pressure inside the vessel is dominated by the nature of the surface.
Complementary information

## Slide 74: As a function of pressure
P (mbar) | Force per unit surface | Density (molec.cm-3) | Mean free path | Incidence rate (cm-2.s-1) | Monolayer formation time
1013 | 1 kg/cm2 | 2.5 1019 | 0.1 μm | 2.9 1023 | 3.4 ns
1 | 1 g/cm2 | 2.5 1016 | 0.1 mm | 2.9 1020 | 3.4 μs
10-3 | 1 mg/cm2 | 2.5 1013 | 10 cm | 2.9 1017 | 3.4 ms
10-6 | 1 μg/cm2 | 2.5 1010 | 100 m | 2.9 1014 | 3.4
10-9 | 10 μg/m2 | 2.5 107 | 100 km | 2.9 1011 | 1 h
10-12 | 10 ng/m2 | 2.5 104 | 105 km | 2.9 109 | 40 days
Complementary information

## Slide 75
The flux, Q, correspond to the current, I
The pressure, P, corresponds to the voltage, V
The inverse of the conductance, C, corresponds to the resistance, R
Adding resistances in parallel
Adding resistances in series
Conductance: Analogy to electricity
V2
V1
R, I
Ohm’s law:
R1
R2
R1
R2
Complementary information
Vacuum systems can be computed using software developed for electrical engineering!

## Slide 76: More conductances
Complementary information
Short tube: the conductance of an orifice multiplied by the Clausing factor i.e. the transmission probability trough the “short tube”
For long circular tube, the transmission probability equals:
Table & plot courtesy of Y. Li and X. Liu
http://uspas.fnal.gov/materials/15ODU/Session1_Fundamentals.pdf

## Slide 77: Calculating conductances with Molflow+
Complementary information
A test particle Monte-Carlo code for molecular flow
http://molflow.web.cern.ch/
R. Kersevan – M. Ady
courtesy of Y. Li and X. Liu

## Slide 78
3.1 Vacuum gauges

## Slide 79
Pirani Gauge
Pirani gauges are commonly used in the range 1 atm -10-4 mbar.
The operating principle is based on the variation of the thermal conductivity of the gases as a function of pressure. A resistor under vacuum is heated at a constant temperature (~ 120°C). The heating current required to keep the temperature constant is a measure of the pressure.
In the viscous regime, the thermal conductivity is independent of the pressure. Therefore pressure readings given above 1 mbar are wrong !
True vs indicated pressure
K. Jousten. J.Vac.Sci. Technol. 26(3), May/Jun 2008, 352-359
Complementary information

## Slide 80
P. Redhead. J.Vac.Sci. 21(5), Sept/Oct 2003, S1-S5
(cathode)
Penning Gauge
Penning gauges are commonly used in the range 10-5 -10-10 mbar. They are use for interlocking purposes
It is a cold cathode ionisation gauge i.e. there are no hot filament
The operating principle is based on the measurement of a discharge current in a Penning cell which is a function of pressure : I+ = Pn, n is close to 1
At high pressure the discharge is unstable due to arcing.
At low pressure, the discharge extinguishes which means zero pressure reading …
Electrons are produced by field emission and perform oscillations due to the magnetic field
Along the path length, molecules are ionised and ions are collected onto the cathode
WARNING : leakage current on the HV cables simulates a higher pressure
Complementary information

## Slide 81: Spinning rotor gauge
It is a reference standard which allows an absolute measurement of the pressure
It is fully bake able up to 450 deg (so low outgassing rate)
SRG gauges are used in the range 5 10-7 -10-2 mbar, accuracy ~ 2 %, uncertainty ~0.5 %.
A stainless steel ball is magnetically levitated and rotated by a drive assembly
Collisions of the gas molecules with the surface of the ball decelerates the ball: the principle of measurement is mechanical
The viscosity of the gas shall be entered into the power supply for pressure measurement
Viscosity, sgaz | H2 | He | CH4 | H2O | N2 | CO | Ar | CO2 | Xe
10-6 Pa.s | 8.8 | 19.6 | 10.8 | 9 | 17.5 | 17.6 | 22.1 | 14.6 | 21.9
Complementary information

## Slide 82: Capacitance diaphragm gauge
It allows an absolute direct measurement of the pressure: the reading is
independent of the gas species!
It operates in the range 10-5 – 104 mbar
Electrodes are placed in a vacuum of ~ 10-7 mbar maintained by a getter.
The capacitance varies under the deflection of a 25 micron thick diaphragm due to the vacuum force (the pressure in the vacuum system)
The sensitivity of the diaphragm deflection is 0.4 nm (10-4 pF) !
It is sensitive to thermal fluctuations: a thermostated housing is used (45 +/0.02 degres C)
Uncertainty : 0.3 to 10 % - Accuracy: 0.15 %
K. Jousten, S. Nael, J.Vac. Sci.Technol. A 29(1), Jan/Feb 2011, 011011-1
N2
Courtesy Vacom
Complementary information

## Slide 83: Strain gauge
It allows a direct measurement of the pressure: the reading is independent of the gas species!
A strain sensor is placed in a sealed volume
The change of the pressure causes a diaphragm to move against the piezo sensitive plate of the strain sensor.
This induce the change of the resistance of the piezoelectric transducer part of a bridge network
Complementary information

## Slide 84
Bayard-Alpert Gauge: CERN construction
Large electron path length are needed to increase the vacuum gauge sensitivity: L ~ 7 cm
The ionization volume (grid volume) must be large
The gauge is mounted on a DN63 flange
It is bake able to 400°C
Low outgassing rate (~ 5 10-10 mbar.l/s)
Large sensitivity ~ 40 mbar-1 for N2
a. W Filament: Ø0.18, height 30
b. Closed grid: Ø35 x 45, pitch 2
Platinium-iridium wire: Ø0.13
c. W Collector: Ø0.05, length 42
d. Modulator: Ø0.7, length 42
Vgrid = 50 cm3
Complementary information

## Slide 85
Bayard-Alpert Gauge: Filament
Electron are produced at the W filament by thermo-electronic emission
The emission is a function of the work function, Φ, of the material
Filaments are made of Iridium/Ytrium, W/Thorium coated to reduce the work function, thereby reducing the operating temperature, T, (thus gas load into the system)
Some electrons can bombard the grid with ~ 150 eV: they can desorb gas molecules and produce an artificial increase of the pressure.
The vacuum gauges are therefore “degassed” by electron bombardment at large current (~ 10 mA) and energy () of the grid and electrodes
Filaments have pumping speed: 0.1 – 1 l/s
D. Alpert, Le Vide 17 (1962),19
Typical chemical reactions at the filament
Richardson-Dushman equation
A = (1-r) Ao
A0 = 120 A/(cm2K2)
Material | A | Φ (eV)
W | 60 | 4.54
W-Th | 3 | 2.63
Ir | 60 | 5.3
Ir-Y | 2 | 2.6
Complementary information

## Slide 86
Bayard-Alpert Gauge: X-rays
Electrons which are bombarding the grid, creates photons by bremsstrahlung
A fraction of those photons irradiated the collector producing a photoelectron
The photoelectron production is interpreted by the electronic system as a positive charge i.e. as a constant pressure, Px-ray
X-ray limit of Pxray ~ 2 10-12 mbar
Complementary information

## Slide 87
Bayard-Alpert Gauge: Modulation
A correction of the residual pressure due to the x-ray can by applied by the “modulation” technique
X-ray limit of Pxray ~ 2 10-12 mbar
 | H2 | He | CH4 | N2 | CO | C2H6 | Ar | CO2
k | 0.86 | 0.91 | 0.89 | 0.89 | 0.89 | 0.89 | 0.89 | 0.89
Modulator | High Pres. | Low Pres.
0 V | I’1 = I’+ | I1 = I+ + Ix
+ grid | I’2 = k I’+ | I2 = k I+ + Ix
=
Complementary information

## Slide 88: Helmer gauge
Electrostatically-deflected ion beam with an electron suppressor grid
Very low x-ray value ~ 5 10-14 mbar
10-13 to 10-6 mbar with Ie ~ 10 mA
10-6 to 10-4 mbar with Ie ~ 0.18 mA
Sensitivity: 20 mbar-1 for N2
Bakeable to 450 deg C, degassing current at 100 mA with + 500 V on the grid
Complementary information

## Slide 89: Helmer gauge and XHV
Recently measured pressure is ~ 5 10-14 mbar in a dedicated set up
Courtesy B. Jenninger
Complementary information

## Slide 90: Extractor gauge
Anode: grid at 220 V
Cathode: filament at 100 V
Reflector at 205 V
Ion collector at 0 V
The ion collection (collector) is located outside the ionization region (grid)
The reflector used to reflect the ions on the collector tip for sensitivy enhancement
Low x-ray value ~ 10-12 mbar
10-12 to 10-4 mbar
Operating current: 1.6 mA, filament in Iridium with yttric oxid coating
Sensitivity: 5 mbar-1 for N2 (due to the small grid volume Ø12 x 25)
Bakeable up to 350°C
Vgrid = 10 cm3
Complementary information

## Slide 91
3.2 Gas analysis

## Slide 92
RGA: Unbaked system
A typical spectrum of an unbaked vacuum system
M/e | Ion | Molecule
2 | H2+ | H2
18 | H2O+ | H2O
 | H2 | CH4 | H2O | CO | CO2
I (A) | 1 10-9 | 1 10-10 | 2 10-8 | 7 10-10 | 2 10-10
P (mbar) | 1 10-9 | 3 10-11 | 1 10-8 | 4 10-10 | 2 10-10
% Pi | 10 | 0 | 85 | 3 | 1
Complementary information

## Slide 93
RGA: Contamination
RGA spectra are used to qualify the cleanliness of a vacuum system
A typical spectrum of a hydrocarbons contaminated system
Peaks are separated by 14 units corresponding to one or more CH2 groups (CnH2n+2)
Acceptance limits are defined for validation tests
Complementary information

## Slide 94
3.3 Vacuum pumps

## Slide 95
Primary Pumps
Are used to pump down from atmosphere down to 10-2 mbar with a speed of a few m3/h
They are usually used as a backing pump of turbomolecular pumps
Two categories : dry and wet pumps.
Dry pumps are expensive and need additional cooling (water)
Wet pumps are operating with oil which acts as a sealing, a lubricant, a heat exchanger and protects parts from rust and corrosion
A.D. Chew. CAS Vacuum in accelerators CERN 2007-003
Oil Sealed Rotary Vane Pump
Complementary information

## Slide 96
Turbomolecular Pump
This pump operates in the molecular regime and is used to pump down an accelerator vacuum system. Usually, it is installed with its primary pump on a mobile trolley : it can be removed after valving off
Its ultimate pressure can be very low : 10-11 mbar
Its pumping speed range from 10 to 3 000 l/s
The compression ratio (Pinlet/Poutlet) increase exponentially with √M : “clean” vacuum without hydrocarbons. So, the oil contamination from the primary pump is avoided
The pumping mechanism is based on the transfer of impulse. When a molecule collide a blade, it is adsorbed for a certain lenght of time. After re-emission, the blade speed is added to the thermal speed of the molecules. To be significant, the blade speed must be comparable to the thermal speed hence it requires fast moving surfaces (~ 40 000 turns/min)
Picture Pfeiffer
Complementary information

## Slide 97
Sputter Ion Pump
This pump operate in the range 10-5 -10-11 mbar. It is used to maintain the pressure in the vacuum chamber of an accelerator.
Their pumping speed range from 1 to 500 l/s.
When electrons spiral in the Penning cell, they ionised molecules. Ions are accelerated towards the cathode (few kV) and sputter Ti. Ti, which is deposited onto the surfaces, forms a chemical bounding with molecules of the residual gas. Noble gases and hydrocarbons ,which does not react with Ti, are buried or implanted onto the cathode.
Advantage : like for a Penning gauge, the collected current is proportional to the pressure. It is also used for interlock.
Picture Agilent, Varian
Triode
(noble gas)
Complementary information

## Slide 98: Sublimation pump
This pumps operates in the UHV regime in a range << 10-9 mbar
It is used as a complementary pumping system to e.g. ion pumps.
Their pumping speed is very large : ~5 000 l/s, typical surface ~ 1000 cm2
A Ti filament is sublimated by Joule effect. Ti is adsorbed on the vacuum chamber wall and provides molecular pumping by a “gettering” effect.
Getterable gases are H2, H2O, O2, CO, O2 , CO2 and,
When assisted with liquid N2 cooling H2 pumping speed is boosted
Sublimation at regular intervals is needed.
l/s cm-2 | H2 | D2 | H2O | CO | N2 | O2 | CO2
300 K | 2.6 | 3.1 | 7.3 | 8.2 | 3.5 | 8.7 | 4.7
77 K | 17.0 | 6.2 | 14.6 | 11.0 | 8.2 | 11.0 | 9.3
K. Welch, Capture pumping technology, Pergamon press, 1991
Complementary information

## Slide 99: NEG cartridge
This pumps operates in the UHV regime in a range < 10-8 mbar
It is used as a complementary pumping system to e.g. ion pumps.
Zr based getter materials are sintered into porous disks.
After activation at ~ 500ºC for 1h, the pumping speed can be large: 100 – 2 000 l/s
Large sorption capacity: >0.1 – 10 mbar.l for N2, CO and > 10 mbar.l for H2!
Reminder: 1 mbar.l = 4.3 1019 molecules, ~ 1 monolayers of a 10m long,Ø10 tube
H2 diffuse into the bulk, CO is adsorbed in 1 active site whereas N2 requires 6.
Gas mixture: CO adsorption inhibit H2 and N2 pumping
Courtesy G. Bregliozzi, J. Callagher, 2012
Complementary information

## Slide 100: NEG coated chamber
Complementary information
A ~ 1 μm getter getter film made of TiZrVa is coated inside a vacuum chamber
Ti-Zr-V is coated by magnetron sputtering with Kr gas
Courtesy P. Chiggiato
Pumping Speed
Very large pumping speed : ~ 250 l/s/m for H2, 20 000 l/s.m for CO
Very low outgassing rate
But : limited capacity and fragile coating sensitive to pollutant (hydrocarbons, Fluor …)
P. Costa Pinto, P. Chiggiato / Thin Solid Films 515 (2006) 382-388